## READ THIS
Olink® Explore HT Validation Data | Last updated October 4, 2023
The performance characteristics of Olink® Explore HT were carefully validated before market introduction. Results for the analytical measurement range, precision, measured plasma protein levels and information regarding observed cross-reactivity are presented in this document. Analytical measurement range: For each biomarker, recombinant antigens in 16 concentrations (1:4 dilutions) were used to define the calibration curve and hook. Hook: Highest NPX value in the calibration curve Hooked: TRUE: Biomarker reached hook, antigen excess relative to the antibody probes at given NPX. FALSE: Biomarker did not reach hook, antigen not excess relative to the antibody probes at given NPX. NA: Data is not available Limit of detection (LOD) was calculated based on the negative controls (NC) in validation runs. LOD: Median (NCs) + 3SD(NCs) (or Median (NCs) + 0.2, whichever is highest) If the biomarker’s maximum counts in NCs is ≤ 150 counts, the LOD is set at counts level using a read count threshold. The count threshold is equal to 2*maximum counts in NCs (or 150, whichever is highest). SD: Standard Deviation Precision: Coefficient of variation (CV) is calculated from 6 individual control samples within each of 3 replicate runs. CV is used to measure technical variation within a plate (intra-CV) or between plates (inter-CV). The variability of data below the background cannot be accurately defined. Therefore, only data above LOD is used to estimate CV. Biomarkers with no data above LOD are reported as NA in the table. Plasma protein level: Commercially available plasma samples from healthy controls (n=21) and patients (n=80) with a range of diseases including cardiovascular, autoimmune, cancer and neurological diseases were analyzed in three replicate runs. Most of the disease samples had multiple diagnosed diseases. All replicates were treated as individual samples to calculate the distribution of protein levels in each group. Biomarkers with no data above LOD are reported as NA in the table. Specificity: In practice, it is not possible to test all Explore HT biomarkers against each other to evaluate potential cross-reactivity between assays. Here an alternative approach was used, where 179 antigens with high homology (over 80%) within the Explore HT product were selected to create a "specificity test sample pool" in order to investigate potential signal contribution from closely related proteins. All Explore HT biomarkers were tested against the "specificity test sample pool". Note: The in-vivo concentrations of the protein biomarkers may be much higher or lower than the concentration tested (5 ng/mL). Therefore, the signal contribution at endogenous level may differ between proteins and individuals.

## Olink Explore HT
|  | Last updated October 4, 2023 |  |  | Analytical range |  |  | Precision |  | Plasma levels (NPX) in Normal samples |  |  | Plasma levels(NPX) in Disease samples
UniProt | Protein name | Gene name | Dilution factor | Block | LOD (NPX) | Hook (NPX) | Hooked* | Intra-CV (%) | Inter-CV (%) | 10th %tile | Median | 90th %tile | 10th %tile | Median | 90th %tile
P04217 | Alpha-1B-glycoprotein | A1BG | 1:100000 | Block_8 | -3.75 | NA | NA | 6.9722 | 6.0114 | -0.37826 | 0.1328 | 0.7473 | -0.31329 | 0.26925 | 0.73782
P43652 | Afamin | AFM | 1:100000 | Block_8 | -0.28785 | NA | NA | 8.0596 | 13.1925 | -0.074933 | 0.32457 | 0.88878 | NA | 0.28105 | 0.66519
P02743 | Serum amyloid P-component | APCS | 1:100000 | Block_8 | -2.83 | 2.9106 | FALSE | 6.3494 | 5.3103 | -0.68225 | 0.0086072 | 1.0416 | -0.50118 | 0.2464 | 0.82491
P02647 | Apolipoprotein A-I | APOA1 | 1:100000 | Block_8 | -5.15 | 1.3286 | FALSE | 8.6856 | 12.2089 | -0.34804 | 0.43001 | 1.6291 | -0.43937 | 0.51924 | 1.5479
P06727 | Apolipoprotein A-IV | APOA4 | 1:100000 | Block_8 | -1.81 | 4.6702 | TRUE | 4.9073 | 21.6701 | -0.56324 | -0.040948 | 0.87746 | -0.63639 | 0.22464 | 1.5363
P05090 | Apolipoprotein D | APOD | 1:100000 | Block_8 | -1.25 | 5.7826 | TRUE | 40.4575 | 33.7875 | -0.44387 | 0.14728 | 0.63723 | -0.61003 | 0.015699 | 0.77385
P02649 | Apolipoprotein E | APOE | 1:100000 | Block_8 | -0.04 | 4.8726 | TRUE | NA | 6.7899 | NA | 0.0078236 | 0.61416 | NA | 0.035464 | 0.59728
O75882 | Attractin | ATRN | 1:100000 | Block_8 | -0.94233 | 6.7823 | FALSE | 10.7214 | 8.0406 | -0.24164 | 0.20445 | 0.77965 | -0.15165 | 0.28191 | 0.78675
P61769 | Beta-2-microglobulin | B2M | 1:100000 | Block_8 | -1.1 | 9.598 | TRUE | 16.0875 | 17.1979 | -0.29706 | 0.21331 | 0.78638 | -0.024236 | 0.66736 | 2.0341
P43251 | Biotinidase | BTD | 1:100000 | Block_8 | -3.23 | 6.9218 | TRUE | 6.6945 | 5.7392 | -0.58777 | -0.074765 | 0.51911 | -0.52742 | 0.04965 | 0.56968
Q9NZP8 | Complement C1r subcomponent-like protein | C1RL | 1:100000 | Block_8 | -1.683 | 7.3033 | TRUE | 7.7068 | 8.492 | -0.38391 | 0.10876 | 0.73265 | -0.17681 | 0.3493 | 0.78975
P09871 | Complement C1s subcomponent | C1S | 1:100000 | Block_8 | -1.8457 | 8.5355 | TRUE | 6.6068 | 8.46 | -0.40868 | 0.20779 | 0.77617 | -0.22915 | 0.34527 | 0.84134
P01024 | Complement C3 | C3 | 1:100000 | Block_8 | -4.2 | 4.0877 | FALSE | 6.4349 | 14.4634 | -0.99864 | -0.29186 | 0.88228 | -0.62513 | 0.20735 | 0.86335
P0C0L5 | Complement C4-B | C4B | 1:100000 | Block_8 | -6 | NA | NA | 7.283 | 13.6761 | -0.94791 | 0.15258 | 1.1358 | -0.93451 | 0.3492 | 1.181
P01031 | Complement C5 | C5 | 1:100000 | Block_8 | -0.80947 | 10.2382 | FALSE | 9.9842 | 9.3921 | NA | -0.07226 | 0.86154 | NA | -0.13312 | 1.5299
P10643 | Complement component C7 | C7 | 1:100000 | Block_8 | -3.95 | 6.8422 | TRUE | 8.9845 | 6.2581 | -0.97324 | 0.14084 | 0.66556 | -0.26745 | 0.40409 | 1.4217
P07358 | Complement component C8 beta chain | C8B | 1:100000 | Block_8 | -0.35105 | 4.7544 | FALSE | 6.8175 | 8.2531 | -0.22986 | 0.088982 | 0.49467 | -0.24723 | 0.11781 | 0.53
O43866 | CD5 antigen-like | CD5L | 1:100000 | Block_8 | -3.3 | 8.9726 | TRUE | 7.336 | 6.6819 | -0.36855 | 0.63098 | 1.4648 | -0.16446 | 0.80469 | 2.3366
P00751 | Complement factor B | CFB | 1:100000 | Block_8 | -4.61 | 3.5362 | FALSE | 6.674 | 6.861 | -0.45022 | -0.0064433 | 1.0636 | -0.29579 | 0.23029 | 0.95031
P00746 | Complement factor D | CFD | 1:100000 | Block_8 | -1.16 | 8.647 | TRUE | 5.2319 | 5.7494 | -0.40469 | 0.063198 | 0.53356 | -0.12337 | 0.41352 | 1.252
P08603 | Complement factor H | CFH | 1:100000 | Block_8 | -6.86 | -1.8728 | FALSE | 6.3126 | 4.2567 | -0.36514 | 0.092209 | 0.82963 | -0.050944 | 0.45923 | 0.94219
P36980 | Complement factor H-related protein 2 | CFHR2 | 1:100000 | Block_8 | -4.99 | 7.0351 | TRUE | 5.4819 | 8.3559 | -0.45052 | 0.19957 | 0.6839 | -0.5659 | 0.55209 | 1.3789
Q92496 | Complement factor H-related protein 4 | CFHR4 | 1:100000 | Block_8 | -1.349 | 6.6763 | TRUE | 6.7044 | 8.8751 | -1.1997 | 0.15634 | 1.0272 | -0.90793 | 0.24456 | 1.0921
P05156 | Complement factor I | CFI | 1:100000 | Block_8 | -3 | 6.7781 | TRUE | 6.7211 | 5.9182 | -0.3072 | 0.060466 | 0.81602 | 0.034218 | 0.45188 | 0.93811
P27918 | Properdin | CFP | 1:100000 | Block_8 | -3.03 | 1.949 | FALSE | 6.5532 | 6.7733 | -0.42743 | 0.13011 | 0.84096 | -0.32963 | 0.23674 | 0.82287
P10909 | Clusterin | CLU | 1:100000 | Block_8 | -1.11 | 4.8644 | TRUE | 12.2346 | 13.2462 | -0.53483 | 0.071971 | 0.74164 | -0.3794 | 0.21193 | 0.70925
Q96IY4 | Carboxypeptidase B2 | CPB2 | 1:100000 | Block_8 | -1.1677 | 5.1913 | TRUE | 6.7744 | 6.8348 | -0.28702 | 0.14772 | 0.7593 | -0.25755 | 0.23531 | 0.66876
Q16610 | Extracellular matrix protein 1 | ECM1 | 1:100000 | Block_8 | -3.47 | 8.8384 | FALSE | 5.5367 | 5.1926 | -0.61654 | 0.0027832 | 0.60554 | -0.384 | 0.077491 | 0.67738
P00748 | Coagulation factor XII | F12 | 1:100000 | Block_8 | -2.0545 | 6.9585 | TRUE | 6.4971 | 4.3741 | -0.33546 | 0.23625 | 0.63644 | -0.41773 | 0.31934 | 0.86998
P05160 | Coagulation factor XIII B chain | F13B | 1:100000 | Block_8 | -1.1007 | 6.9707 | TRUE | 6.889 | 5.9944 | 0.0065093 | 0.49419 | 1.3658 | 0.02194 | 0.57816 | 1.0982
P00734 | Prothrombin | F2 | 1:100000 | Block_8 | 0.11761 | 7.7658 | FALSE | NA | 1.4582 | NA | 0.22727 | 1.4912 | NA | NA | 0.71533
P02671 | Fibrinogen alpha chain | FGA | 1:100000 | Block_8 | -1.1245 | 2.4449 | FALSE | 6.6141 | 8.7175 | -0.29717 | 0.34142 | 1.4756 | -0.23789 | 0.70236 | 1.6741
P02679 | Fibrinogen gamma chain | FGG | 1:100000 | Block_8 | -1.1101 | NA | NA | 9.7725 | 11.6708 | -0.57989 | 0.52974 | 2.086 | -0.52992 | 0.94266 | 2.1976
P02751 | Fibronectin | FN1 | 1:100000 | Block_8 | -0.59747 | 2.2313 | TRUE | 9.1914 | 5.4262 | NA | 0.043172 | 1.2359 | -0.48523 | 0.21262 | 0.9019
P02774 | Vitamin D-binding protein | GC | 1:100000 | Block_8 | -4.5 | 3.0837 | TRUE | 4.9968 | 4.987 | -0.45909 | 0.19754 | 0.75378 | -0.34656 | 0.14375 | 0.64774
Q14520 | Hyaluronan-binding protein 2 | HABP2 | 1:100000 | Block_8 | -0.58706 | NA | NA | 10.0631 | 12.0645 | -0.42786 | 0.088588 | 1.2009 | -0.33212 | 0.3616 | 1.1336
P04196 | Histidine-rich glycoprotein | HRG | 1:100000 | Block_8 | -3.75 | 5.1434 | TRUE | 5.9893 | 10.2813 | -0.31726 | 0.28176 | 1.0625 | -0.17032 | 0.43334 | 1.098
P0DOY2 | Immunoglobulin lambda constant 2 | IGLC2 | 1:100000 | Block_8 | -0.12448 | 9.1852 | TRUE | 13.3496 | 10.7553 | NA | 0.053843 | 0.63224 | NA | 0.20716 | 1.3692
P19827 | Inter-alpha-trypsin inhibitor heavy chain H1 | ITIH1 | 1:100000 | Block_8 | -0.9053 | 7.6813 | TRUE | 4.6864 | 7.1536 | -0.26279 | 0.051321 | 0.66686 | -0.32438 | 0.14719 | 0.84925
Q14624 | Inter-alpha-trypsin inhibitor heavy chain H4 | ITIH4 | 1:100000 | Block_8 | -4 | 4.2261 | TRUE | 10.3992 | 9.3894 | -0.57959 | 0.091033 | 0.68012 | -0.33265 | 0.48883 | 1.2159
P03952 | Plasma kallikrein | KLKB1 | 1:100000 | Block_8 | -1.332 | 6.6889 | TRUE | 5.5865 | 6.8308 | -0.36386 | 0.55075 | 1.076 | -0.21338 | 0.46904 | 0.97384
Q08380 | Galectin-3-binding protein | LGALS3BP | 1:100000 | Block_8 | -1.6702 | 5.553 | FALSE | 6.1794 | 5.4438 | -0.905 | -0.20135 | 0.9634 | -0.64718 | 0.094641 | 1.3369
P08519 | Apolipoprotein(a) | LPA | 1:100000 | Block_8 | -3.0267 | 4.5752 | TRUE | 52.7584 | 49.7501 | -2.9876 | -0.51854 | 1.0341 | NA | -2.279 | 0.35206
P02750 | Leucine-rich alpha-2-glycoprotein | LRG1 | 1:100000 | Block_8 | -1.775 | 4.6859 | TRUE | 6.8124 | 7.5244 | -0.74928 | 0.1028 | 0.79563 | -0.42634 | 0.50634 | 1.5329
P51884 | Lumican | LUM | 1:100000 | Block_8 | -0.27784 | NA | NA | 13.0977 | 6.4484 | NA | 0.37537 | 1.0081 | NA | 0.40845 | 1.0073
Q9Y5Y7 | Lymphatic vessel endothelial hyaluronic acid receptor 1 | LYVE1 | 1:100000 | Block_8 | -1.2671 | 11.1214 | TRUE | 12.9985 | 6.6037 | -0.83594 | 0.21095 | 0.66051 | -0.50486 | 0.41992 | 1.3512
P11226 | Mannose-binding protein C | MBL2 | 1:100000 | Block_8 | -3.5 | 0.91947 | FALSE | 6.287 | 5.1133 | -1.0262 | 0.57987 | 1.8776 | -1.4717 | 0.27551 | 1.7144
P02763 | Alpha-1-acid glycoprotein 1 | ORM1 | 1:100000 | Block_8 | -1.6748 | 3.3473 | FALSE | 8.8889 | 9.1431 | -0.63219 | 0.0034902 | 0.8006 | -0.36155 | 0.3484 | 1.1767
Q96PD5 | N-acetylmuramoyl-L-alanine amidase | PGLYRP2 | 1:100000 | Block_8 | -3 | NA | NA | 5.6136 | 7.8373 | -0.72622 | 0.32067 | 0.88262 | -0.39239 | 0.29589 | 0.94299
P00747 | Plasminogen | PLG | 1:100000 | Block_8 | -4 | 4.4372 | TRUE | 5.967 | 6.2862 | -0.38569 | 0.16134 | 0.704 | -0.21227 | 0.20584 | 0.74021
P02775 | Platelet basic protein | PPBP | 1:100000 | Block_8 | -1.55 | 9.2582 | TRUE | 4.3803 | 4.8163 | -0.19532 | 1.3367 | 2.6075 | 0.56014 | 2.6756 | 4.1058
P20742 | Pregnancy zone protein | PZP | 1:100000 | Block_8 | -0.64605 | 7.8199 | FALSE | 4.3209 | 9.0531 | NA | 0.11084 | 0.98331 | NA | -0.21547 | 1.2508
P02753 | Retinol-binding protein 4 | RBP4 | 1:100000 | Block_8 | -5 | NA | NA | 8.8789 | 9.3902 | -0.49528 | 0.57511 | 1.6956 | -0.40401 | 0.90122 | 2.2772
P49908 | Selenoprotein P | SELENOP | 1:100000 | Block_8 | -1.88 | 6.5275 | TRUE | 5.3542 | 5.5904 | -0.49691 | 0.074669 | 0.66147 | -0.32949 | 0.27384 | 0.84673
P01009 | Alpha-1-antitrypsin | SERPINA1 | 1:100000 | Block_8 | -1.0686 | 2.8916 | FALSE | 10.4905 | 9.1913 | -1.0643 | -0.60242 | -0.13653 | NA | -0.5505 | 0.59211
P01011 | Alpha-1-antichymotrypsin | SERPINA3 | 1:100000 | Block_8 | -2.29 | 3.6865 | TRUE | 5.5643 | 7.8684 | -0.57333 | 0.048681 | 0.53457 | -0.012933 | 0.5449 | 1.1906
P29622 | Kallistatin | SERPINA4 | 1:100000 | Block_8 | -4.36 | NA | NA | 5.8425 | 5.179 | -0.29878 | 0.1103 | 0.77939 | -0.3373 | 0.28319 | 0.74726
P05154 | Plasma serine protease inhibitor | SERPINA5 | 1:100000 | Block_8 | -2.98 | 2.7609 | FALSE | 11.9099 | 32.9327 | -2.5909 | -0.65051 | 0.079097 | -2.7774 | -1.0431 | 1.0123
P08185 | Corticosteroid-binding globulin | SERPINA6 | 1:100000 | Block_8 | -1.2802 | 7.5219 | FALSE | 7.8135 | 10.2227 | -0.26588 | 0.17718 | 0.80422 | -0.24089 | 0.27642 | 0.70818
P05543 | Thyroxine-binding globulin | SERPINA7 | 1:100000 | Block_8 | -4.09 | 7.1534 | TRUE | 7.0564 | 5.2685 | -0.36736 | 0.11544 | 0.93355 | -0.52404 | 0.20161 | 0.7254
P01008 | Antithrombin-III | SERPINC1 | 1:100000 | Block_8 | -1.9744 | 3.9843 | TRUE | 7.01 | 5.1967 | -0.13854 | 0.32241 | 0.78839 | -0.0074151 | 0.48523 | 0.94548
P05546 | Heparin cofactor 2 | SERPIND1 | 1:100000 | Block_8 | -1.9102 | 2.02 | FALSE | 7.4976 | 5.3651 | -0.41282 | -0.01604 | 0.62145 | -0.43677 | 0.1802 | 0.84444
P36955 | Pigment epithelium-derived factor | SERPINF1 | 1:100000 | Block_8 | -0.47332 | 5.239 | FALSE | 7.3177 | 5.0742 | -0.35333 | 0.021936 | 0.49738 | -0.16512 | 0.25472 | 0.69549
P08697 | Alpha-2-antiplasmin | SERPINF2 | 1:100000 | Block_8 | -4.06 | 4.4514 | FALSE | 8.4893 | 9.1678 | -0.23619 | 0.24136 | 0.88083 | -0.24462 | 0.32445 | 0.86647
P05155 | Plasma protease C1 inhibitor | SERPING1 | 1:100000 | Block_8 | -5 | 2.8997 | TRUE | 5.3235 | 7.9694 | -0.61774 | -0.18271 | 0.25485 | -0.31825 | 0.081704 | 0.58865
P02787 | Serotransferrin | TF | 1:100000 | Block_8 | -4.67 | 1.7271 | TRUE | 8.1264 | 9.5206 | -0.51475 | 0.17029 | 0.92947 | -0.72483 | 0.00084839 | 0.63007
P02766 | Transthyretin | TTR | 1:100000 | Block_8 | -5.87 | 4.0273 | TRUE | 7.0594 | 8.1092 | -0.43602 | 0.2581 | 0.74959 | -0.44899 | 0.31316 | 0.97366
P04004 | Vitronectin | VTN | 1:100000 | Block_8 | -5 | NA | NA | 6.8318 | 13.1719 | -1.1357 | -0.22273 | 1.6139 | -1.35 | -0.083377 | 0.90114
P16442 | Histo-blood group ABO system transferase | ABO | 1:1000 | Block_7 | -0.30095 | 13.6471 | TRUE | 8.2275 | 7.5973 | NA | NA | 1.1953 | NA | -0.0038337 | 1.1307
P12821 | Angiotensin-converting enzyme | ACE | 1:1000 | Block_7 | -1.4657 | 9.1026 | FALSE | 10.5642 | 4.6259 | -0.42313 | 0.4504 | 0.87575 | -0.19167 | 0.60062 | 1.3521
Q9NZK5 | Adenosine deaminase 2 | ADA2 | 1:1000 | Block_7 | -5.3126 | 7.4977 | TRUE | 5.5228 | 5.0037 | -0.54373 | 0.36552 | 1.038 | -0.11495 | 0.85717 | 2.132
P01019 | Angiotensinogen | AGT | 1:1000 | Block_7 | -4.4 | 1.299 | FALSE | 12.3094 | 6.8834 | -1.1466 | -0.19403 | 1.254 | -1.0498 | 0.12002 | 1.947
P02765 | Alpha-2-HS-glycoprotein | AHSG | 1:1000 | Block_7 | -5.7792 | -0.85436 | FALSE | 7.6211 | 15.2122 | -0.69816 | -0.11734 | 0.43231 | -0.75304 | -0.0998 | 0.48242
P04746_P0DTE7_P0DTE8_P0DUB6_P19961 | Alpha-amylase 1/2 | AMY1_AMY2 | 1:1000 | Block_7 | -1 | 6.8583 | TRUE | 6.4248 | 3.0782 | -0.7364 | -0.19962 | 0.35258 | -0.67591 | -0.10275 | 0.43742
P03950 | Angiogenin | ANG | 1:1000 | Block_7 | -4.8979 | 2.7709 | TRUE | 6.4745 | 4.3899 | -0.51718 | -0.0082703 | 0.66056 | -0.38827 | 0.36051 | 1.0741
Q9Y5C1 | Angiopoietin-related protein 3 | ANGPTL3 | 1:1000 | Block_7 | -5.4525 | 6.4969 | TRUE | 6.9657 | 4.6012 | -0.31479 | 0.34748 | 1.0238 | -0.23279 | 0.72049 | 1.3891
P15144 | Aminopeptidase N | ANPEP | 1:1000 | Block_7 | -2.7425 | 5.4568 | FALSE | 5.4973 | 8.9082 | -0.26339 | 0.16445 | 0.59955 | -0.35529 | 0.40093 | 1.1229
Q16853 | Membrane primary amine oxidase | AOC3 | 1:1000 | Block_7 | -2.0029 | 4.4531 | TRUE | 5.1203 | 8.2947 | -0.26237 | 0.59657 | 1.1596 | -0.15398 | 0.63751 | 1.3221
P02652 | Apolipoprotein A-II | APOA2 | 1:1000 | Block_7 | -3.194 | NA | NA | 12.8096 | 13.3796 | -0.7884 | 0.69661 | 1.5314 | -1.2344 | 0.29583 | 1.5694
P04114 | Apolipoprotein B-100 | APOB | 1:1000 | Block_7 | 0.4629 | 5.5853 | FALSE | NA | NA | NA | 0.75172 | 1.7646 | NA | NA | 1.3968
Q13790 | Apolipoprotein F | APOF | 1:1000 | Block_7 | -6.6574 | -1.2783 | FALSE | 4.0255 | 3.7022 | -0.94124 | 0.25988 | 1.0483 | -1.142 | 0.89684 | 1.4107
O14791 | Apolipoprotein L1 | APOL1 | 1:1000 | Block_7 | -6.92 | 2.102 | FALSE | 4.1002 | 5.8528 | -3.4146 | -2.5966 | -0.52538 | -3.6956 | -2.7619 | -0.079271
O95445 | Apolipoprotein M | APOM | 1:1000 | Block_7 | -5.4779 | 4.6169 | TRUE | 7.249 | 6.1126 | -0.67715 | 0.16681 | 0.78345 | -0.84193 | -0.061031 | 0.54288
Q13508 | Ecto-ADP-ribosyltransferase 3 | ART3 | 1:1000 | Block_7 | -3.6027 | 6.6976 | TRUE | 3.7351 | 8.1171 | -0.65957 | -0.18212 | 0.58817 | -0.38516 | 0.27746 | 1.2676
P07307 | Asialoglycoprotein receptor 2 | ASGR2 | 1:1000 | Block_7 | -0.4938 | 5.1052 | TRUE | 8.3028 | 9.388 | -0.47044 | 0.19907 | 1.0827 | 0.056934 | 0.6895 | 1.6437
P30530 | Tyrosine-protein kinase receptor UFO | AXL | 1:1000 | Block_7 | -2.5814 | 6.4735 | TRUE | 5.6694 | 6.3029 | -0.19937 | 0.18736 | 0.82363 | 0.081927 | 0.75662 | 1.327
P06276 | Cholinesterase | BCHE | 1:1000 | Block_7 | -5.3905 | 3.1818 | TRUE | 5.3408 | 2.9532 | -0.52178 | 0.12711 | 0.57206 | -0.40467 | 0.35473 | 0.78492
Q9BXJ1 | Complement C1q tumor necrosis factor-related protein 1 | C1QTNF1 | 1:1000 | Block_7 | -3.0501 | 6.578 | TRUE | 14.1739 | 9.8439 | -0.36742 | 0.70683 | 2.3627 | 0.09233 | 0.95561 | 2.1095
P00736 | Complement C1r subcomponent | C1R | 1:1000 | Block_7 | -4 | NA | NA | 6.8154 | 5.444 | -0.008791 | 0.498 | 0.93476 | 0.081982 | 0.52198 | 0.92589
P06681 | Complement C2 | C2 | 1:1000 | Block_7 | -5.448 | 2.1149 | FALSE | 5.9324 | 4.5733 | -1.187 | -0.67416 | -0.017987 | -0.9737 | -0.41551 | 0.67932
P02748 | Complement component C9 | C9 | 1:1000 | Block_7 | -3.7 | NA | NA | 16.8459 | 16.3013 | -0.78413 | 0.15995 | 0.89524 | -0.4658 | 0.42516 | 1.3
P00915 | Carbonic anhydrase 1 | CA1 | 1:1000 | Block_7 | -2.8029 | 5.8054 | FALSE | 3.6703 | 5.5364 | 0.60572 | 2.4719 | 3.2093 | 1.2289 | 2.5047 | 3.8259
P04040 | Catalase | CAT | 1:1000 | Block_7 | -0.72 | 6.3625 | TRUE | 9.7841 | 14.7584 | 0.15322 | 1.4636 | 2.0355 | 0.80796 | 1.7574 | 2.8495
Q16627 | C-C motif chemokine 14 | CCL14 | 1:1000 | Block_7 | -4.8472 | 6.4858 | TRUE | 7.583 | 7.5234 | -0.61146 | -0.17334 | 0.76633 | -0.039596 | 0.82187 | 1.8401
P55774 | C-C motif chemokine 18 | CCL18 | 1:1000 | Block_7 | -4.0472 | 6.112 | TRUE | 6.6649 | 4.4665 | -0.94649 | -0.34251 | 1.0364 | -0.41899 | 0.61031 | 1.5631
P13501 | C-C motif chemokine 5 | CCL5 | 1:1000 | Block_7 | -1.2245 | 9.0814 | TRUE | 9.1669 | 14.3597 | -0.11066 | 3.0504 | 4.2245 | 1.8951 | 4.1018 | 5.9533
P08571 | Monocyte differentiation antigen CD14 | CD14 | 1:1000 | Block_7 | -7.3669 | 8.1531 | FALSE | 7.7303 | 9.326 | -0.71396 | 0.44535 | 1.4052 | 0.068621 | 1.0756 | 2.1696
Q86VB7 | Scavenger receptor cysteine-rich type 1 protein M130 | CD163 | 1:1000 | Block_7 | -3.9932 | 5.951 | TRUE | 5.0354 | 4.5566 | -0.58367 | 0.16131 | 0.63419 | -0.19801 | 0.67857 | 1.5407
P15529 | Membrane cofactor protein | CD46 | 1:1000 | Block_7 | -3.6287 | 8.0558 | TRUE | 6.2932 | 6.0721 | -0.1199 | 0.39037 | 0.99649 | 0.61987 | 1.0907 | 1.7496
P08174 | Complement decay-accelerating factor | CD55 | 1:1000 | Block_7 | -2.2566 | 6.614 | TRUE | 3.9308 | 5.168 | -0.239 | 0.20595 | 0.63324 | 0.10736 | 0.673 | 1.3125
Q9NPY3 | Complement component C1q receptor | CD93 | 1:1000 | Block_7 | -5.3439 | 6.8274 | TRUE | 4.2783 | 3.8291 | -0.57992 | 0.27542 | 0.87135 | -0.069355 | 0.64653 | 1.5987
P12830 | Cadherin-1 | CDH1 | 1:1000 | Block_7 | -2.3689 | 6.4078 | TRUE | 7.3569 | 8.5612 | -0.54175 | 0.24985 | 0.84708 | -0.10852 | 0.55415 | 1.48
Q9BXR6 | Complement factor H-related protein 5 | CFHR5 | 1:1000 | Block_7 | -7.5 | 2.0628 | TRUE | 5.9128 | 3.6842 | -0.45226 | 0.035372 | 0.79096 | -0.39879 | 0.32442 | 0.9146
P0DN86 | Choriogonadotropin subunit beta 3 | CGB3_CGB5_CGB8 | 1:1000 | Block_7 | -3.023 | 6.5841 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P36222 | Chitinase-3-like protein 1 | CHI3L1 | 1:1000 | Block_7 | -2.85 | 6.5048 | TRUE | 3.9747 | 5.3 | -1.8642 | -1.0459 | 0.69259 | -0.91525 | 0.72303 | 2.3394
O00533 | Neural cell adhesion molecule L1-like protein | CHL1 | 1:1000 | Block_7 | -1.9521 | 5.3833 | TRUE | 6.7147 | 7.7369 | -0.32342 | 0.3705 | 0.85084 | -0.2511 | 0.39591 | 0.97991
Q96KN2 | Beta-Ala-His dipeptidase | CNDP1 | 1:1000 | Block_7 | -3.817 | 6.4841 | FALSE | 3.5103 | 11.0553 | -0.13458 | 0.91951 | 2.1074 | -0.83874 | 0.24063 | 1.6628
P49747 | Cartilage oligomeric matrix protein | COMP | 1:1000 | Block_7 | -4.2995 | 5.7599 | TRUE | 3.6314 | 4.9597 | -0.2106 | 0.50351 | 1.0229 | -0.17007 | 0.70948 | 1.5957
P22792 | Carboxypeptidase N subunit 2 | CPN2 | 1:1000 | Block_7 | -3.1195 | NA | NA | 4.4835 | 6.9362 | -0.45965 | -0.031515 | 0.33456 | -0.4179 | -0.02416 | 0.3465
P17927 | Complement receptor type 1 | CR1 | 1:1000 | Block_7 | -2.8604 | 8.9268 | TRUE | 11.5138 | 6.5857 | -0.61128 | -0.21068 | 0.69669 | -0.24593 | 0.62571 | 1.2389
P20023 | Complement receptor type 2 | CR2 | 1:1000 | Block_7 | -6.2045 | 9.178 | TRUE | 5.3005 | 4.8587 | -0.93417 | 0.054371 | 0.92035 | -1.1741 | 0.21347 | 1.3027
P54108 | Cysteine-rich secretory protein 3 | CRISP3 | 1:1000 | Block_7 | -1.6736 | 3.4384 | TRUE | 7.1377 | 6.9175 | -0.47833 | 0.13577 | 0.52567 | -0.51119 | 0.014197 | 0.51582
P07333 | Macrophage colony-stimulating factor 1 receptor | CSF1R | 1:1000 | Block_7 | -3.7033 | 6.5382 | TRUE | 7.3355 | 3.6881 | -0.48776 | 0.10436 | 0.81089 | -0.28066 | 0.38839 | 1.0747
P0DML2 | Chorionic somatomammotropin hormone 1 | CSH1 | 1:1000 | Block_7 | 0.64765 | 11.514 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P01034 | Cystatin-C | CST3 | 1:1000 | Block_7 | -6.17 | 0.066314 | FALSE | 5.3849 | 5.5292 | -0.45701 | 0.25987 | 0.79381 | 0.24585 | 0.96386 | 1.7693
Q01459 | Di-N-acetylchitobiase | CTBS | 1:1000 | Block_7 | -4 | 6.1125 | TRUE | 5.6459 | 5.898 | -0.22688 | 0.27847 | 0.94384 | -0.15916 | 0.4052 | 1.0394
P07339 | Cathepsin D | CTSD | 1:1000 | Block_7 | 1.2383 | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1.4675
Q9UBR2 | Cathepsin Z | CTSZ | 1:1000 | Block_7 | -2.7041 | 6.1067 | TRUE | 11.4137 | 4.3855 | -0.35573 | 0.21375 | 0.94103 | -0.13989 | 0.63265 | 1.3817
P09172 | Dopamine beta-hydroxylase | DBH | 1:1000 | Block_7 | -4.07 | 4.7338 | FALSE | 4.4682 | 5.0189 | -1.265 | 0.018068 | 1.2228 | -1.6656 | -0.25259 | 1.1044
P59665 | Neutrophil defensin 1 | DEFA1 | 1:1000 | Block_7 | -0.90698 | 6.3843 | TRUE | 8.1415 | 6.7348 | -0.5694 | 0.13874 | 1.1802 | 0.019099 | 0.95197 | 3.4675
P27487 | Dipeptidyl peptidase 4 | DPP4 | 1:1000 | Block_7 | -3.3865 | 5.0845 | TRUE | 5.2978 | 4.9589 | -0.17847 | 0.31105 | 0.81292 | -0.35167 | 0.30306 | 0.96839
Q12805 | EGF-containing fibulin-like extracellular matrix protein 1 | EFEMP1 | 1:1000 | Block_7 | -3.6939 | 4.6215 | TRUE | 5.3546 | 8.0088 | -0.46967 | 0.51557 | 1.2069 | 0.18009 | 1.0771 | 2.1009
P08246 | Neutrophil elastase | ELANE | 1:1000 | Block_7 | -2.3163 | 7.4808 | TRUE | 10.4983 | 32.5152 | -0.91239 | -0.36697 | 0.10521 | -0.7921 | -0.09018 | 1.7325
P00742 | Coagulation factor X | F10 | 1:1000 | Block_7 | -3.2436 | 5.3699 | TRUE | 5.0936 | 4.534 | -0.48488 | 0.089949 | 0.6013 | -0.43278 | 0.12151 | 0.57026
P03951 | Coagulation factor XI | F11 | 1:1000 | Block_7 | -5.3675 | NA | NA | 5.8108 | 3.1086 | -0.50447 | 0.20843 | 0.64756 | -0.14018 | 0.35157 | 0.80707
P08709 | Coagulation factor VII | F7 | 1:1000 | Block_7 | -2.3158 | 7.7103 | TRUE | 4.2206 | 7.8947 | 0.073321 | 0.74543 | 1.3152 | 0.33096 | 0.95635 | 1.5974
P00740 | Coagulation factor IX | F9 | 1:1000 | Block_7 | -4.8022 | 3.3903 | TRUE | 4.2617 | 8.0226 | 0.010029 | 0.48781 | 1.0759 | 0.18249 | 0.70435 | 1.1056
P12318 | Low affinity immunoglobulin gamma Fc region receptor II-a | FCGR2A | 1:1000 | Block_7 | -3.5328 | 5.3991 | TRUE | 4.1732 | 10.1138 | -0.73374 | 0.014834 | 0.99705 | -0.69403 | 0.36743 | 0.98244
O75015_P08637 | Low affinity immunoglobulin gamma Fc region receptor III-A/B | FCGR3B_FCGR3A | 1:1000 | Block_7 | -5.6562 | NA | NA | 4.7537 | 3.2445 | -0.84706 | 0.47309 | 1.0329 | -0.36787 | 0.66247 | 1.3707
O00602 | Ficolin-1 | FCN1 | 1:1000 | Block_7 | -3 | 9.6959 | FALSE | 9.0278 | 5.2804 | 0.38802 | 1.2072 | 1.8123 | 0.67663 | 1.7822 | 2.74
Q15485 | Ficolin-2 | FCN2 | 1:1000 | Block_7 | -4.5695 | 4.5802 | TRUE | 5.7916 | 7.3437 | -0.45243 | 0.36113 | 0.91296 | -0.52396 | 0.38008 | 1.0707
Q9UGM5 | Fetuin-B | FETUB | 1:1000 | Block_7 | -4.0986 | 4.5733 | TRUE | 7.5746 | 13.2944 | -0.53976 | 0.25279 | 0.81469 | -0.7115 | 0.2082 | 0.78988
Q08830 | Fibrinogen-like protein 1 | FGL1 | 1:1000 | Block_7 | -2.5756 | 5.5178 | FALSE | 10.269 | 7.6891 | -1.1067 | 0.42386 | 1.9365 | -0.50784 | 0.67622 | 2.21
Q92820 | Gamma-glutamyl hydrolase | GGH | 1:1000 | Block_7 | -3.41 | 4.7146 | TRUE | 9.2442 | 8.7014 | -0.80507 | -0.11828 | 0.55377 | -0.55866 | 0.11455 | 0.76602
P07359 | Platelet glycoprotein Ib alpha chain | GP1BA | 1:1000 | Block_7 | -5.2664 | 5.3795 | TRUE | 3.8777 | 4.4481 | -0.45698 | 0.26165 | 0.7175 | 0.042262 | 0.7542 | 1.2024
P06744 | Glucose-6-phosphate isomerase | GPI | 1:1000 | Block_7 | -1.0667 | 8.0511 | TRUE | 7.3467 | 5.0717 | -0.55036 | 0.84975 | 2.4102 | 0.93905 | 2.819 | 4.4098
P06396 | Gelsolin | GSN | 1:1000 | Block_7 | -3.5487 | 2.3626 | FALSE | 4.707 | 4.2428 | -0.27508 | 0.08243 | 0.56767 | -0.47649 | 0.087205 | 0.47548
P00390 | Glutathione reductase, mitochondrial | GSR | 1:1000 | Block_7 | -3.321 | 6.6888 | TRUE | 3.7278 | 6.0772 | -0.35789 | 0.32904 | 0.82698 | 0.16796 | 0.77903 | 1.5438
Q04756 | Hepatocyte growth factor activator | HGFAC | 1:1000 | Block_7 | -5.6 | 2.8189 | TRUE | 5.3761 | 8.0944 | -0.56638 | 0.18009 | 0.65764 | -0.30613 | 0.28539 | 0.73273
P98160 | Basement membrane-specific heparan sulfate proteoglycan core protein | HSPG2 | 1:1000 | Block_7 | -4.3246 | 7.8038 | TRUE | 6.5285 | 4.5726 | -0.285 | 0.2344 | 0.69764 | 0.1835 | 0.83391 | 1.7741
Q12794 | Hyaluronidase-1 | HYAL1 | 1:1000 | Block_7 | -4.0011 | 6.5893 | TRUE | 5.228 | 5.3449 | -0.41303 | 0.20363 | 0.73939 | -0.02406 | 0.45316 | 1.0144
P05362 | Intercellular adhesion molecule 1 | ICAM1 | 1:1000 | Block_7 | -6.0623 | 5.3233 | TRUE | 6.8388 | 5.3774 | -0.34515 | 0.18277 | 0.95413 | -0.047976 | 0.56429 | 1.4078
P17936 | Insulin-like growth factor-binding protein 3 | IGFBP3 | 1:1000 | Block_7 | -6.3719 | 3.7969 | TRUE | 5.0671 | 3.4747 | -0.92983 | 0.15334 | 0.68079 | -0.63602 | 0.2531 | 0.83947
P24592 | Insulin-like growth factor-binding protein 6 | IGFBP6 | 1:1000 | Block_7 | -5.9127 | 5.9017 | TRUE | 5.705 | 6.0459 | -0.53267 | 0.15531 | 0.97656 | 0.011758 | 0.79401 | 2.0974
O95998 | Interleukin-18-binding protein | IL18BP | 1:1000 | Block_7 | -3.2539 | 7.9421 | TRUE | 4.7507 | 5.3724 | -0.40837 | 0.32849 | 1.0926 | 0.081726 | 1.0107 | 1.9183
P08887 | Interleukin-6 receptor subunit alpha | IL6R | 1:1000 | Block_7 | -3.6064 | 8.9826 | TRUE | 6.4447 | 6.1569 | -0.28005 | 0.41105 | 1.1921 | -0.045324 | 0.64295 | 1.4238
P05556 | Integrin beta-1 | ITGB1 | 1:1000 | Block_7 | -2.1164 | 6.3384 | TRUE | 5.6559 | 9.1209 | -0.14868 | 0.38985 | 0.84518 | 0.04807 | 0.50775 | 1.0782
Q06033 | Inter-alpha-trypsin inhibitor heavy chain H3 | ITIH3 | 1:1000 | Block_7 | -3.2834 | 6.67 | TRUE | 7.5432 | 9.3085 | -0.74999 | -0.027145 | 1.0039 | -0.37478 | 0.58001 | 2.291
P10721 | Mast/stem cell growth factor receptor Kit | KIT | 1:1000 | Block_7 | -5.0867 | 5.9422 | TRUE | 4.9938 | 5.6747 | -0.45316 | 0.16616 | 0.64336 | -0.88931 | -0.27148 | 0.44006
P18428 | Lipopolysaccharide-binding protein | LBP | 1:1000 | Block_7 | -3.7804 | 6.1535 | FALSE | 11.1873 | 10.1916 | -1.2966 | -0.27272 | 1.0484 | -0.56829 | 0.5801 | 1.8666
P04180 | Phosphatidylcholine-sterol acyltransferase | LCAT | 1:1000 | Block_7 | -3.7164 | 3.8417 | TRUE | 6.252 | 7.7456 | -0.33092 | 0.16908 | 0.67598 | -0.1783 | 0.29918 | 0.76391
P80188 | Neutrophil gelatinase-associated lipocalin | LCN2 | 1:1000 | Block_7 | -3 | 5.3301 | TRUE | 5.9447 | 4.5269 | -0.6082 | 0.17178 | 0.75148 | -0.2478 | 0.72377 | 2.6994
P43121 | Cell surface glycoprotein MUC18 | MCAM | 1:1000 | Block_7 | -2.8469 | 7.5274 | TRUE | 5.1851 | 6.5901 | -0.61895 | 0.3878 | 0.89912 | -0.046879 | 0.57052 | 1.2982
P40925 | Malate dehydrogenase, cytoplasmic | MDH1 | 1:1000 | Block_7 | 0.19943 | 4.7883 | FALSE | NA | NA | NA | NA | 0.3117 | NA | 0.30995 | 0.69259
P22897 | Macrophage mannose receptor 1 | MRC1 | 1:1000 | Block_7 | -3.1677 | 7.4578 | TRUE | 5.0279 | 9.2834 | -0.6085 | 0.31093 | 1.2039 | -0.066465 | 0.79138 | 1.8713
P26927 | Hepatocyte growth factor-like protein | MST1 | 1:1000 | Block_7 | -4.947 | 4.7176 | TRUE | 6.065 | 4.547 | -0.17179 | 0.62799 | 1.0883 | -0.26583 | 0.69951 | 1.2669
P13591 | Neural cell adhesion molecule 1 | NCAM1 | 1:1000 | Block_7 | -2.6008 | 7.2403 | TRUE | 8.292 | 4.8893 | -0.55235 | -0.047817 | 0.61524 | -0.32116 | 0.21571 | 0.83291
Q0ZGT2 | Nexilin | NEXN | 1:1000 | Block_7 | 0.38928 | 7.9864 | TRUE | NA | NA | NA | 0.46242 | 3.38 | 0.95159 | 3.0745 | 4.646
P14543 | Nidogen-1 | NID1 | 1:1000 | Block_7 | -3.7653 | 7.6114 | TRUE | 5.4679 | 3.2715 | -0.8595 | 0.59687 | 1.2042 | 0.47355 | 1.0863 | 1.9288
Q92686 | Neurogranin | NRGN | 1:1000 | Block_7 | 0.33754 | 6.8857 | FALSE | NA | NA | NA | NA | 1.8888 | NA | 1.7805 | 3.7465
O14786 | Neuropilin-1 | NRP1 | 1:1000 | Block_7 | -3.1977 | 5.8158 | TRUE | 5.8935 | 6.428 | -0.35355 | 0.11453 | 0.83705 | 0.21925 | 0.79293 | 1.4183
Q99650 | Oncostatin-M-specific receptor subunit beta | OSMR | 1:1000 | Block_7 | -1.1364 | 5.9482 | TRUE | 4.5442 | 4.158 | -0.071343 | 0.26141 | 0.79184 | 0.069063 | 0.47124 | 0.98663
Q15113 | Procollagen C-endopeptidase enhancer 1 | PCOLCE | 1:1000 | Block_7 | -4.7781 | 4.6012 | TRUE | 4.3784 | 4.9338 | -0.63583 | 0.16303 | 0.82316 | -0.10282 | 0.77277 | 1.922
P12955 | Xaa-Pro dipeptidase | PEPD | 1:1000 | Block_7 | -2.6222 | 6.2261 | FALSE | 8.8382 | 12.674 | -0.48028 | 0.29017 | 0.9544 | -0.33458 | 0.23643 | 0.92414
P02776 | Platelet factor 4 | PF4 | 1:1000 | Block_7 | -0.60278 | 6.8618 | TRUE | 6.2318 | 11.1968 | 0.17685 | 1.5069 | 2.7794 | 0.812 | 2.9756 | 4.4376
P0DJD7 | Pepsin A-4 | PGA4 | 1:1000 | Block_7 | -5.4877 | 7.9297 | TRUE | 14.0655 | 7.9007 | -1.3786 | -0.085451 | 0.68335 | -1.309 | 0.1476 | 1.9153
Q6UXB8 | Peptidase inhibitor 16 | PI16 | 1:1000 | Block_7 | -2.1193 | 4.2677 | TRUE | 4.0293 | 5.69 | -0.41212 | 0.61366 | 1.0998 | -0.13346 | 0.62393 | 1.2995
P16233 | Pancreatic triacylglycerol lipase | PNLIP | 1:1000 | Block_7 | -2.5659 | 9.0865 | TRUE | 6.9382 | 8.4292 | -1.2437 | -0.4322 | 1.2916 | -1.1462 | 0.03737 | 1.2364
P27169 | Serum paraoxonase/arylesterase 1 | PON1 | 1:1000 | Block_7 | -6 | 1.8442 | TRUE | 6.8032 | 6.7487 | -0.095755 | 0.21114 | 0.50613 | -0.51239 | 0.15988 | 0.58195
P32119 | Peroxiredoxin-2 | PRDX2 | 1:1000 | Block_7 | 0.26585 | 6.1089 | TRUE | NA | 2.5377 | NA | 1.4154 | 2.0517 | 0.35291 | 1.2371 | 2.3999
P04070 | Vitamin K-dependent protein C | PROC | 1:1000 | Block_7 | -5.9937 | 3.8401 | TRUE | 7.4077 | 5.6472 | -0.30836 | 0.33133 | 0.8591 | -0.18024 | 0.49009 | 1.1002
P07478 | Trypsin-2 | PRSS2 | 1:1000 | Block_7 | -4.7363 | 3.1388 | TRUE | 4.2926 | 5.4425 | -0.88995 | 0.17247 | 1.359 | -0.54441 | 0.59198 | 2.2382
P41222 | Prostaglandin-H2 D-isomerase | PTGDS | 1:1000 | Block_7 | -3.8109 | 4.4165 | TRUE | 6.2273 | 11.2498 | -0.34059 | 0.2491 | 0.91904 | 0.081173 | 0.71851 | 1.9582
Q99969 | Retinoic acid receptor responder protein 2 | RARRES2 | 1:1000 | Block_7 | -4.6368 | 6.4868 | TRUE | 17.5066 | 8.4219 | -1.144 | -0.22668 | 0.75873 | -0.33553 | 0.57394 | 1.5437
P05451 | Lithostathine-1-alpha | REG1A | 1:1000 | Block_7 | -4.5077 | 6.7318 | TRUE | 6.1263 | 4.123 | -0.51238 | 0.24697 | 0.92511 | 0.0021112 | 0.94523 | 2.3836
P07998 | Ribonuclease pancreatic | RNASE1 | 1:1000 | Block_7 | -5.8 | 5.4911 | TRUE | 8.9272 | 12.4466 | -0.54368 | 0.3532 | 0.92218 | 0.10629 | 1.3546 | 2.764
P34096 | Ribonuclease 4 | RNASE4 | 1:1000 | Block_7 | -5.3 | 4.0767 | TRUE | 5.5341 | 8.3455 | -0.3704 | 0.29729 | 0.9467 | 0.039609 | 0.7996 | 1.7528
Q93091 | Ribonuclease K6 | RNASE6 | 1:1000 | Block_7 | -1.9094 | 9.7472 | TRUE | 4.8931 | 8.884 | -0.42555 | 0.21041 | 0.79986 | 0.024285 | 0.90978 | 1.9549
P35542 | Serum amyloid A-4 protein | SAA4 | 1:1000 | Block_7 | -3.5087 | 3.5459 | TRUE | 6.5041 | 3.3216 | -0.79849 | 0.46114 | 1.487 | -0.64367 | 0.31221 | 1.8004
P16109 | P-selectin | SELP | 1:1000 | Block_7 | -3.72 | 8.2207 | TRUE | 6.0065 | 7.4718 | -0.50962 | 0.88824 | 1.6687 | 0.54669 | 1.7789 | 2.9643
P04278 | Sex hormone-binding globulin | SHBG | 1:1000 | Block_7 | -4.7149 | 2.2607 | FALSE | 4.156 | 7.7431 | 0.079558 | 0.76853 | 1.9799 | -0.15106 | 0.91821 | 1.8221
P00441 | Superoxide dismutase [Cu-Zn] | SOD1 | 1:1000 | Block_7 | 0.93819 | 7.3937 | TRUE | NA | NA | NA | NA | 2.0108 | 1.0648 | 2.1397 | 3.3518
P08294 | Extracellular superoxide dismutase [Cu-Zn] | SOD3 | 1:1000 | Block_7 | -0.65409 | 4.3655 | FALSE | 5.3065 | 11.0091 | -0.60475 | 0.15763 | 0.65807 | -0.27811 | 0.2673 | 0.95321
Q14515 | SPARC-like protein 1 | SPARCL1 | 1:1000 | Block_7 | -5.1645 | 6.3248 | TRUE | 6.2442 | 7.0244 | -0.5345 | 0.14085 | 0.87601 | -0.2118 | 0.52766 | 1.303
A1L4H1 | Soluble scavenger receptor cysteine-rich domain-containing protein SSC5D | SSC5D | 1:1000 | Block_7 | -4.6519 | 3.9776 | TRUE | 7.1763 | 3.2166 | -0.54938 | 0.52634 | 1.5375 | -0.025803 | 0.89686 | 1.7379
P20061 | Transcobalamin-1 | TCN1 | 1:1000 | Block_7 | -2.39 | 9.4355 | TRUE | 5.3736 | 3.6204 | -0.43361 | 0.047226 | 0.54 | -0.35787 | 0.38051 | 1.5
P20062 | Transcobalamin-2 | TCN2 | 1:1000 | Block_7 | -4.27 | 6.5839 | TRUE | 6.5213 | 6.5222 | -0.35371 | 0.11151 | 0.58696 | -0.20859 | 0.49178 | 1.1809
Q07654 | Trefoil factor 3 | TFF3 | 1:1000 | Block_7 | -1.13 | 8.6639 | TRUE | 7.8768 | 8.5796 | -0.36487 | 0.11672 | 0.63226 | -0.061187 | 0.76237 | 2.2122
P10646 | Tissue factor pathway inhibitor | TFPI | 1:1000 | Block_7 | -4.3105 | 6.1547 | TRUE | 5.7467 | 4.7143 | -0.21057 | 0.30533 | 1.0437 | 0.23977 | 0.69534 | 1.3215
P02786 | Transferrin receptor protein 1 | TFRC | 1:1000 | Block_7 | -1.4392 | 5.3838 | FALSE | 7.2305 | 7.2614 | -0.80334 | -0.10378 | 0.90562 | -0.73924 | 0.094741 | 1.025
Q15582 | Transforming growth factor-beta-induced protein ig-h3 | TGFBI | 1:1000 | Block_7 | -5.0688 | 4.2593 | FALSE | 6.5575 | 5.587 | -0.49594 | 0.31995 | 1.0153 | -0.22566 | 0.33998 | 0.84369
O43493 | Trans-Golgi network integral membrane protein 2 | TGOLN2 | 1:1000 | Block_7 | -2.8136 | 2.4952 | TRUE | 5.9302 | 8.6613 | -0.30191 | 0.53362 | 1.1016 | 0.13898 | 0.90635 | 1.7121
P35443 | Thrombospondin-4 | THBS4 | 1:1000 | Block_7 | -3.6859 | 6.0878 | FALSE | 5.3831 | 4.2989 | -0.66038 | 0.33642 | 0.87109 | -0.3275 | 0.60448 | 1.4765
Q96H15 | T-cell immunoglobulin and mucin domain-containing protein 4 | TIMD4 | 1:1000 | Block_7 | -1.19 | 5.9037 | TRUE | 12.7671 | 8.2386 | -0.80222 | -0.089614 | 0.37477 | -0.51856 | 0.2566 | 1.2389
P01033 | Metalloproteinase inhibitor 1 | TIMP1 | 1:1000 | Block_7 | -3.89 | 6.0688 | TRUE | 4.8813 | 4.2008 | -0.38236 | 0.36567 | 0.86381 | 0.66091 | 1.1811 | 1.9974
P24821 | Tenascin | TNC | 1:1000 | Block_7 | -2.9641 | 8.0985 | TRUE | 7.2217 | 7.5878 | -0.46691 | 0.74677 | 1.2631 | 0.037123 | 1.0353 | 2.2739
Q86YW5 | Trem-like transcript 1 protein | TREML1 | 1:1000 | Block_7 | 0.3963 | 12.6509 | TRUE | 10.3411 | 17.8785 | NA | 1.8166 | 3.9693 | 1.0684 | 4.0051 | 5.5464
P10599 | Thioredoxin | TXN | 1:1000 | Block_7 | 2.3005 | 9.1113 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.4585
Q6EMK4 | Vasorin | VASN | 1:1000 | Block_7 | -3.0994 | 6.7337 | TRUE | 5.2086 | 7.1858 | -0.39713 | 0.29183 | 0.96077 | 0.040302 | 0.73664 | 1.3794
P50552 | Vasodilator-stimulated phosphoprotein | VASP | 1:1000 | Block_7 | 0.3226 | 4.604 | FALSE | 5.6972 | 1.484 | NA | 0.80233 | 4.0023 | 1.4678 | 4.0866 | 5.7716
P19320 | Vascular cell adhesion protein 1 | VCAM1 | 1:1000 | Block_7 | -2.1093 | 4.7896 | TRUE | 4.9067 | 6.1305 | -0.25211 | 0.17076 | 0.93042 | 0.043992 | 0.6503 | 1.482
O95497 | Pantetheinase | VNN1 | 1:1000 | Block_7 | -3.4807 | 5.3065 | TRUE | 6.8888 | 4.6331 | -1.2441 | 0.73121 | 1.2709 | -0.51713 | 0.67796 | 1.8163
P13686 | Tartrate-resistant acid phosphatase type 5 | ACP5 | 1:100 | Block_6 | -4.56 | 7.2956 | TRUE | 18.8839 | 14.5373 | -1.0277 | -0.23397 | 0.51373 | -0.57018 | 0.43342 | 1.3518
Q76LX8 | A disintegrin and metalloproteinase with thrombospondin motifs 13 | ADAMTS13 | 1:100 | Block_6 | -5.7585 | 5.7316 | TRUE | 2.5801 | 6.6811 | -0.097479 | 0.30812 | 0.65495 | -0.53167 | 0.28135 | 0.85575
Q86TH1 | ADAMTS-like protein 2 | ADAMTSL2 | 1:100 | Block_6 | -1.3318 | 6.277 | FALSE | 6.659 | 8.5832 | -0.3859 | 0.047721 | 0.55209 | -0.15388 | 0.48921 | 1.3329
P48960 | Adhesion G protein-coupled receptor E5 | ADGRE5 | 1:100 | Block_6 | -3.7208 | 5.7318 | TRUE | 4.3949 | 12.3705 | -0.43444 | 0.40385 | 0.8941 | -0.0051977 | 0.63787 | 1.3257
P02771 | Alpha-fetoprotein | AFP | 1:100 | Block_6 | 0.04235 | 9.8641 | TRUE | 10.8447 | 6.7265 | NA | 0.29778 | 0.9376 | NA | 0.17856 | 0.96597
Q9NZD4 | Alpha-hemoglobin-stabilizing protein | AHSP | 1:100 | Block_6 | -3.1913 | 8.138 | TRUE | 8.0478 | 11.2103 | -0.53611 | 1.1715 | 2.4975 | 0.093212 | 1.6594 | 3.4213
Q13740 | CD166 antigen | ALCAM | 1:100 | Block_6 | -5.5947 | 4.9594 | TRUE | 4.2943 | 7.4265 | -0.24465 | 0.34972 | 0.89529 | 0.043804 | 0.6462 | 1.2464
P00352 | Aldehyde dehydrogenase 1A1 | ALDH1A1 | 1:100 | Block_6 | -1.6594 | 6.8061 | TRUE | 9.1208 | 12.376 | -0.53755 | 0.51185 | 1.5772 | 0.016105 | 0.88029 | 2.143
Q9Y303 | N-acetylglucosamine-6-phosphate deacetylase | AMDHD2 | 1:100 | Block_6 | 0.9424 | 7.3193 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P02749 | Beta-2-glycoprotein 1 | APOH | 1:100 | Block_6 | 0.20457 | NA | NA | 9.1203 | 16.5541 | NA | 0.55532 | 1.239 | NA | 0.84776 | 1.6436
P05067 | Amyloid-beta precursor protein | APP | 1:100 | Block_6 | 0.82991 | 4.8888 | TRUE | NA | NA | NA | NA | 1.4664 | NA | 1.3371 | 2.2512
Q13510 | Acid ceramidase | ASAH1 | 1:100 | Block_6 | -0.64 | 6.4848 | TRUE | 53.2521 | 27.728 | -0.13246 | 0.4952 | 1.6976 | 0.56532 | 1.4806 | 2.8357
P20160 | Azurocidin | AZU1 | 1:100 | Block_6 | -3.59 | 3.9737 | TRUE | 12.1415 | 28.4013 | -1.915 | -1.2189 | -0.048493 | -1.7283 | -0.53799 | 1.7059
O43505 | Beta-1,4-glucuronyltransferase 1 | B4GAT1 | 1:100 | Block_6 | -3.4185 | 7.1483 | TRUE | 10.0556 | 13.5332 | -0.41807 | 0.19739 | 0.73879 | -0.17262 | 0.46227 | 1.063
P50895 | Basal cell adhesion molecule | BCAM | 1:100 | Block_6 | -1.5656 | 6.1514 | TRUE | 7.0776 | 11.4746 | -0.21255 | 0.29112 | 0.67516 | -0.00054363 | 0.59736 | 1.2373
P23560 | Brain-derived neurotrophic factor | BDNF | 1:100 | Block_6 | -2.0166 | 10.0005 | TRUE | 4.5498 | 14.424 | 0.59126 | 2.3748 | 3.5528 | 0.93271 | 3.7432 | 5.2131
P02818 | Osteocalcin | BGLAP | 1:100 | Block_6 | -4.2214 | 8.9126 | TRUE | 9.0476 | 11.4351 | -1.6131 | -0.39687 | 0.82596 | -2.9858 | -0.1289 | 2.0445
Q9UBW5 | Bridging integrator 2 | BIN2 | 1:100 | Block_6 | 0.6557 | 6.0209 | FALSE | NA | NA | NA | NA | 3.5303 | NA | 2.7472 | 5.0431
Q13867 | Bleomycin hydrolase | BLMH | 1:100 | Block_6 | -2.89 | 6.2742 | TRUE | 13.4717 | 21.2032 | -0.39348 | 0.65121 | 1.1355 | 0.47492 | 1.2369 | 1.9725
P30043 | Flavin reductase | BLVRB | 1:100 | Block_6 | -0.44383 | 7.4968 | TRUE | 6.3364 | 14.1397 | 0.89638 | 1.6719 | 2.5153 | 0.82672 | 1.678 | 3.0504
Q8TDL5 | BPI fold-containing family B member 1 | BPIFB1 | 1:100 | Block_6 | -8.18 | 1.9185 | FALSE | 8.5286 | 11.2452 | -2.4791 | -0.95769 | 0.25627 | -2.8736 | -1.1208 | 0.18828
P00918 | Carbonic anhydrase 2 | CA2 | 1:100 | Block_6 | -2.83 | 8.9647 | FALSE | 23.8713 | 17.8416 | 1.0173 | 3.1568 | 4.2606 | 2.3992 | 4.0736 | 5.3561
P07451 | Carbonic anhydrase 3 | CA3 | 1:100 | Block_6 | -4.3016 | 6.9865 | FALSE | 4.2025 | 6.1421 | -0.025362 | 0.91383 | 1.9673 | 0.37009 | 1.6206 | 3.236
P22748 | Carbonic anhydrase 4 | CA4 | 1:100 | Block_6 | -3.8567 | 8.2525 | TRUE | 2.7905 | 6.7058 | -0.36272 | 0.18649 | 0.78829 | -0.11693 | 0.46247 | 1.0123
P23280 | Carbonic anhydrase 6 | CA6 | 1:100 | Block_6 | -3.06 | 4.3074 | TRUE | 4.8986 | 10.4154 | -1.1851 | -0.52404 | 1.0099 | -2.0041 | -0.68515 | 0.57101
P42574 | Caspase-3 | CASP3 | 1:100 | Block_6 | -1.378 | 10.4755 | TRUE | 12.8118 | 11.2302 | -0.84134 | 1.5152 | 4.9867 | 2.414 | 5.0865 | 6.4375
Q8IUK8 | Cerebellin-2 | CBLN2 | 1:100 | Block_6 | -1.5675 | NA | NA | 5.4493 | 6.8829 | -0.57084 | 0.16799 | 0.7632 | -0.74759 | 0.034906 | 0.70046
Q16663 | C-C motif chemokine 15 | CCL15 | 1:100 | Block_6 | -4.7108 | 5.3425 | TRUE | 4.545 | 10.5043 | -0.65798 | -0.12575 | 0.82662 | -0.5362 | 0.51004 | 1.6383
O15467 | C-C motif chemokine 16 | CCL16 | 1:100 | Block_6 | -5.7533 | 5.221 | TRUE | 7.3237 | 10.6346 | -1.2052 | -0.071053 | 0.73528 | -0.65132 | 0.44342 | 1.5298
P48745 | CCN family member 3 | CCN3 | 1:100 | Block_6 | -2.5635 | 7.4719 | TRUE | 5.5651 | 13.6656 | -0.46495 | 0.25308 | 0.90336 | 0.26161 | 0.98405 | 2.3674
O76076 | CCN family member 5 | CCN5 | 1:100 | Block_6 | -5.6296 | NA | NA | 7.1872 | 10.3401 | 0.12163 | 0.85383 | 2.2995 | 0.50064 | 1.4829 | 2.4889
Q6YHK3 | CD109 antigen | CD109 | 1:100 | Block_6 | -2.9968 | 5.8241 | FALSE | 4.3912 | 13.6016 | -0.19574 | 0.29453 | 1.0277 | -0.03449 | 0.71336 | 1.3758
Q04900 | Sialomucin core protein 24 | CD164 | 1:100 | Block_6 | -1.39 | 6.3332 | TRUE | 14.1724 | 16.1267 | -0.079825 | 0.33865 | 0.79911 | 0.51449 | 1.0108 | 1.5609
Q8N6Q3 | CD177 antigen | CD177 | 1:100 | Block_6 | -3.58 | 7.0397 | TRUE | 9.9128 | 15.5764 | -0.90396 | 0.065084 | 1.1368 | -1.4196 | 0.85307 | 2.4987
Q9HCU0 | Endosialin | CD248 | 1:100 | Block_6 | -3.3835 | 6.4121 | TRUE | 5.3373 | 17.6199 | -0.45993 | 0.23016 | 1.1523 | -0.013216 | 0.8752 | 1.7083
Q6UXG3 | CMRF35-like molecule 9 | CD300LG | 1:100 | Block_6 | -1.1024 | 5.8792 | TRUE | 3.6497 | 7.8744 | -0.57825 | 0.066415 | 0.57211 | -0.50914 | 0.28435 | 1.3647
P13987 | CD59 glycoprotein | CD59 | 1:100 | Block_6 | -4 | 3.227 | TRUE | 8.8009 | 9.4104 | -0.14495 | 0.20574 | 0.72016 | 0.093846 | 0.68753 | 1.7213
P14209 | CD99 antigen | CD99 | 1:100 | Block_6 | -4.9254 | 4.9315 | TRUE | 3.073 | 9.5423 | -0.074611 | 0.2409 | 0.75729 | 0.19404 | 0.76524 | 1.5296
Q8TCZ2 | CD99 antigen-like protein 2 | CD99L2 | 1:100 | Block_6 | -4.2939 | 7.6888 | TRUE | 6.2941 | 13.4653 | -0.19288 | 0.31406 | 0.65159 | 0.15501 | 0.79951 | 1.4554
P33151 | Cadherin-5 | CDH5 | 1:100 | Block_6 | -3.3762 | 4.6738 | TRUE | 6.0196 | 17.6183 | -0.42451 | 0.31161 | 1.0478 | -0.088732 | 0.71613 | 1.36
P40199 | Carcinoembryonic antigen-related cell adhesion molecule 6 | CEACAM6 | 1:100 | Block_6 | 0.38205 | 7.2877 | TRUE | NA | NA | NA | NA | 0.86235 | NA | NA | 0.87057
P08217 | Chymotrypsin-like elastase family member 2A | CELA2A | 1:100 | Block_6 | -4.8272 | 7.2358 | TRUE | 6.9575 | 10.4576 | -1.2885 | 0.19647 | 1.5165 | -1.3421 | 0.12477 | 1.3928
P09093 | Chymotrypsin-like elastase family member 3A | CELA3A | 1:100 | Block_6 | -3.19 | 6.7967 | TRUE | 3.9405 | 8.7416 | -0.68509 | 0.23063 | 1.6474 | -0.75878 | 0.23358 | 1.4812
P23141 | Liver carboxylesterase 1 | CES1 | 1:100 | Block_6 | -3.3478 | 6.4781 | TRUE | 5.2261 | 6.8849 | -0.62664 | 0.13466 | 1.2357 | -0.75642 | 0.40616 | 1.797
P10645 | Chromogranin-A | CHGA | 1:100 | Block_6 | -2.2347 | 6.8733 | TRUE | 12.2044 | 19.229 | -0.84608 | 1.2736 | 2.0309 | 0.029477 | 1.6767 | 4.7343
Q13231 | Chitotriosidase-1 | CHIT1 | 1:100 | Block_6 | -7.5 | 4.9214 | TRUE | 2.9379 | 7.5158 | -0.93835 | 0.11584 | 0.77136 | -6.6256 | 0.63741 | 2.2748
O75339 | Cartilage intermediate layer protein 1 | CILP | 1:100 | Block_6 | -1.7844 | 7.6709 | TRUE | 14.6712 | 14.4408 | -0.78673 | 0.031024 | 1.2176 | -0.55051 | 0.58428 | 2.0261
P12277 | Creatine kinase B-type | CKB | 1:100 | Block_6 | -3.0002 | 3.1163 | FALSE | 7.2052 | 13.2113 | -0.7955 | 0.30094 | 1.2006 | -0.26941 | 1.2302 | 2.7808
Q9Y240 | C-type lectin domain family 11 member A | CLEC11A | 1:100 | Block_6 | -3.2597 | 7.2088 | TRUE | 5.6524 | 10.6278 | -0.56078 | 0.42906 | 1.1411 | 0.14115 | 0.74801 | 1.7606
Q86T13 | C-type lectin domain family 14 member A | CLEC14A | 1:100 | Block_6 | -1.2023 | 8.5064 | TRUE | 7.368 | 7.5469 | -0.47244 | 0.21334 | 0.73695 | -0.041671 | 0.7496 | 1.8498
Q9H2X3 | C-type lectin domain family 4 member M | CLEC4M | 1:100 | Block_6 | -3.34 | 4.8184 | TRUE | 6.9005 | 16.1383 | -0.37784 | 0.23757 | 0.69025 | -0.49679 | 0.28544 | 0.89117
P04118 | Colipase | CLPS | 1:100 | Block_6 | -6.1913 | 5.9346 | TRUE | 12.4647 | 18.7246 | -1.0451 | -0.00039219 | 1.9439 | -0.71994 | 0.34446 | 2.0126
Q12860 | Contactin-1 | CNTN1 | 1:100 | Block_6 | -2.6836 | 5.7188 | TRUE | 4.0551 | 6.5197 | -0.35191 | 0.29079 | 0.82885 | -0.151 | 0.48694 | 1.0997
Q9P232 | Contactin-3 | CNTN3 | 1:100 | Block_6 | -2.5643 | 4.4716 | TRUE | 6.1829 | 10.7418 | -0.84609 | 0.38761 | 0.95679 | -0.086681 | 0.64403 | 1.2649
Q8IWV2 | Contactin-4 | CNTN4 | 1:100 | Block_6 | -5.1408 | 6.0642 | TRUE | 5.3972 | 11.3962 | -0.25472 | 0.38648 | 0.65685 | -0.068426 | 0.51494 | 1.0825
P39059 | Collagen alpha-1(XV) chain | COL15A1 | 1:100 | Block_6 | -5.43 | 7.4304 | TRUE | 5.3068 | 7.1478 | -0.49384 | 0.15483 | 0.57086 | -0.2568 | 0.5903 | 1.5822
P39060 | Collagen alpha-1(XVIII) chain | COL18A1 | 1:100 | Block_6 | -7.5 | 4.1353 | TRUE | 4.5722 | 7.1661 | -0.56191 | -0.061199 | 0.36921 | 0.014048 | 0.58028 | 1.3591
P02452 | Collagen alpha-1(I) chain | COL1A1 | 1:100 | Block_6 | -5.2192 | 6.0354 | TRUE | 3.5605 | 7.5346 | -0.78244 | 0.0018589 | 0.7795 | -0.78661 | 0.30023 | 1.1145
P02461 | Collagen alpha-1(III) chain | COL3A1 | 1:100 | Block_6 | -1.29 | 7.2998 | TRUE | 11.1227 | 18.2552 | -0.28515 | 0.27504 | 0.98755 | -0.016219 | 0.67699 | 1.5122
P12111 | Collagen alpha-3(VI) chain | COL6A3 | 1:100 | Block_6 | -5.1572 | 5.1346 | TRUE | 4.2085 | 9.0783 | -0.47983 | 0.16781 | 0.70701 | 0.12979 | 1.0749 | 2.4732
Q9BWP8 | Collectin-11 | COLEC11 | 1:100 | Block_6 | -6.8639 | NA | NA | 4.2046 | 7.4381 | -0.74645 | -0.077366 | 0.70188 | -0.40773 | 0.276 | 0.91823
P31146 | Coronin-1A | CORO1A | 1:100 | Block_6 | -0.99968 | 10.7149 | FALSE | 12.9338 | 16.6308 | -0.8452 | 0.96561 | 3.2534 | 1.0475 | 3.9491 | 5.7259
P15085 | Carboxypeptidase A1 | CPA1 | 1:100 | Block_6 | -4 | 4.0687 | TRUE | 4.1682 | 13.8014 | -1.4804 | -0.010601 | 1.446 | -1.0792 | 0.017142 | 1.6531
P15086 | Carboxypeptidase B | CPB1 | 1:100 | Block_6 | -3.8061 | 7.4411 | TRUE | 3.0513 | 9.5418 | -1.2567 | -0.18359 | 1.3086 | -0.81695 | 0.15702 | 1.7852
Q9Y646 | Carboxypeptidase Q | CPQ | 1:100 | Block_6 | -3.7884 | 4.5511 | TRUE | 5.7963 | 9.6944 | -0.7852 | -0.17108 | 0.64811 | -0.72086 | -0.068194 | 0.69224
Q9NQ79 | Cartilage acidic protein 1 | CRTAC1 | 1:100 | Block_6 | -5.0212 | 5.6296 | FALSE | 12.8846 | 10.4081 | -0.66853 | 0.18 | 0.6644 | -0.72345 | 0.072989 | 0.7096
Q6UVK1 | Chondroitin sulfate proteoglycan 4 | CSPG4 | 1:100 | Block_6 | -1.9642 | 7.3995 | TRUE | 4.4824 | 8.401 | -0.30351 | 0.064545 | 0.70927 | -0.081568 | 0.58574 | 1.4088
Q15828 | Cystatin-M | CST6 | 1:100 | Block_6 | -2.63 | 4.2044 | TRUE | 4.459 | 13.8372 | -0.54617 | 0.48036 | 1.3424 | -0.11823 | 0.75998 | 1.7116
P04080 | Cystatin-B | CSTB | 1:100 | Block_6 | -2.51 | 6.4609 | TRUE | 5.7709 | 13.2947 | -0.30079 | 0.47643 | 1.0651 | 0.5525 | 1.6877 | 3.5909
Q96CG8 | Collagen triple helix repeat-containing protein 1 | CTHRC1 | 1:100 | Block_6 | -0.91617 | 8.81 | TRUE | 4.7424 | 10.7948 | 0.015516 | 0.82765 | 1.4464 | 0.43863 | 1.2912 | 2.2175
P16410 | Cytotoxic T-lymphocyte protein 4 | CTLA4 | 1:100 | Block_6 | 1.6379 | 10.8288 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P17538 | Chymotrypsinogen B | CTRB1 | 1:100 | Block_6 | -4.3066 | 7.4368 | TRUE | 5.2054 | 7.3764 | -0.56001 | 0.45765 | 1.3988 | -0.6996 | 0.44899 | 1.8771
P10619 | Lysosomal protective protein | CTSA | 1:100 | Block_6 | -1.1 | 0.6311 | FALSE | 21.2999 | 24.3754 | -0.8812 | 0.011618 | 0.9719 | -0.15251 | 1.0492 | 2.2367
P07858 | Cathepsin B | CTSB | 1:100 | Block_6 | -3.6713 | 5.7435 | TRUE | 8.9094 | 6.3967 | -0.51401 | 0.45132 | 1.2268 | -0.29854 | 0.82748 | 2.1942
Q9H2A7 | C-X-C motif chemokine 16 | CXCL16 | 1:100 | Block_6 | -3.1737 | 5.4607 | TRUE | 6.3313 | 10.5161 | -0.50608 | 0.13458 | 0.69593 | -0.0026741 | 0.51715 | 1.2259
P19876 | C-X-C motif chemokine 3 | CXCL3 | 1:100 | Block_6 | -1.71 | 11.4696 | TRUE | 10.5218 | 6.172 | -1.0764 | 1.324 | 3.0237 | 0.50603 | 2.8033 | 4.6491
Q9NRR1 | Cytokine-like protein 1 | CYTL1 | 1:100 | Block_6 | -2.1422 | 7.7625 | TRUE | 8.3038 | 8.7621 | -0.25707 | 0.32834 | 0.95886 | -0.37518 | 0.70095 | 2.0453
P07108 | Acyl-CoA-binding protein | DBI | 1:100 | Block_6 | -0.25709 | 5.1511 | TRUE | 3.6875 | 9.111 | -0.143 | 0.40679 | 1.3596 | 0.63132 | 1.9432 | 2.9824
P30046 | D-dopachrome decarboxylase | DDT | 1:100 | Block_6 | 0.60977 | 4.6865 | FALSE | NA | NA | NA | NA | 0.95112 | NA | 0.76762 | 1.4905
P60022 | Beta-defensin 1 | DEFB1 | 1:100 | Block_6 | -4.3526 | NA | NA | 4.9647 | 18.8336 | -1.1908 | -0.045507 | 0.79246 | -0.38356 | 0.44377 | 1.6511
Q9UBP4 | Dickkopf-related protein 3 | DKK3 | 1:100 | Block_6 | -4.6779 | 5.882 | TRUE | 6.3538 | 7.905 | -0.1149 | 0.36557 | 1.0386 | 0.3514 | 0.93516 | 1.8519
P80370 | Protein delta homolog 1 | DLK1 | 1:100 | Block_6 | -6.33 | 6.5638 | TRUE | 4.5154 | 12.2091 | -0.75373 | 0.096112 | 1.4591 | -0.12481 | 0.98348 | 2.0386
P24855 | Deoxyribonuclease-1 | DNASE1 | 1:100 | Block_6 | -1.5134 | NA | NA | 15.778 | 6.5327 | -0.47993 | 0.17089 | 0.96047 | -0.60364 | 0.49982 | 1.3131
Q07507 | Dermatopontin | DPT | 1:100 | Block_6 | -3.5028 | 3.3088 | TRUE | 5.1016 | 6.945 | -0.59914 | 0.22992 | 0.99531 | -0.13832 | 0.61626 | 1.2723
Q02487 | Desmocollin-2 | DSC2 | 1:100 | Block_6 | -3.7072 | 6.6517 | TRUE | 5.1446 | 11.4588 | -0.24463 | 0.40446 | 0.83447 | 0.17704 | 0.99992 | 1.9737
Q14126 | Desmoglein-2 | DSG2 | 1:100 | Block_6 | -3.139 | 4.8563 | TRUE | 4.5096 | 9.4631 | 0.19113 | 0.71881 | 1.0937 | 0.39643 | 0.99796 | 1.6056
P42126 | Enoyl-CoA delta isomerase 1, mitochondrial | ECI1 | 1:100 | Block_6 | -2.7187 | NA | NA | 4.3935 | 11.2566 | -1.5754 | -0.74368 | -0.051228 | -1.3023 | -0.28661 | 1.0367
P20827 | Ephrin-A1 | EFNA1 | 1:100 | Block_6 | -2.6879 | 5.4742 | TRUE | 6.376 | 9.6516 | -0.40655 | 0.27585 | 0.63383 | 0.079506 | 0.89646 | 2.1108
P00533 | Epidermal growth factor receptor | EGFR | 1:100 | Block_6 | -3.8688 | NA | NA | 5.4879 | 8.6508 | -0.16027 | 0.22569 | 0.77442 | -0.36688 | 0.23198 | 0.71411
P17813 | Endoglin | ENG | 1:100 | Block_6 | -2.1422 | 4.7574 | TRUE | 8.3517 | 8.9396 | -0.12972 | 0.37255 | 0.84706 | -0.11267 | 0.41191 | 0.90512
Q13822 | Ectonucleotide pyrophosphatase/phosphodiesterase family member 2 | ENPP2 | 1:100 | Block_6 | -2.9399 | 4.9107 | FALSE | 5.649 | 10.9567 | -0.48679 | 0.26978 | 0.91197 | -0.23296 | 0.36897 | 1.1334
P54760 | Ephrin type-B receptor 4 | EPHB4 | 1:100 | Block_6 | -1.4942 | 7.7322 | TRUE | 5.7565 | 9.3224 | -0.39855 | 0.077195 | 0.63969 | -0.07186 | 0.67509 | 1.5655
Q96AP7 | Endothelial cell-selective adhesion molecule | ESAM | 1:100 | Block_6 | -1.717 | 4.9288 | TRUE | 5.689 | 10.6012 | 0.034442 | 0.58054 | 1.1766 | 0.81753 | 1.3861 | 2.1937
Q9Y624 | Junctional adhesion molecule A | F11R | 1:100 | Block_6 | -3.62 | 7.7543 | TRUE | 5.9496 | 8.5696 | -0.21997 | 0.77932 | 2.4363 | 1.1212 | 2.1042 | 3.4407
P15090 | Fatty acid-binding protein, adipocyte | FABP4 | 1:100 | Block_6 | -3.1172 | 7.1569 | TRUE | 5.5837 | 8.6148 | -0.80094 | -0.13479 | 0.42808 | -0.61815 | 0.87347 | 2.8767
Q12884 | Prolyl endopeptidase FAP | FAP | 1:100 | Block_6 | -1.7829 | 4.9407 | TRUE | 3.1847 | 8.293 | -0.41253 | 0.091811 | 0.60979 | -0.3372 | 0.19039 | 0.91638
P25445 | Tumor necrosis factor receptor superfamily member 6 | FAS | 1:100 | Block_6 | -2.9958 | 7.296 | TRUE | 5.1384 | 7.174 | -0.48964 | 0.29886 | 0.74383 | 0.20544 | 0.84019 | 1.59
P98095 | Fibulin-2 | FBLN2 | 1:100 | Block_6 | -3.2188 | 8.3011 | TRUE | 3.3179 | 11.2068 | 0.05774 | 0.58803 | 1.1398 | 0.54083 | 1.4274 | 2.6063
P06734 | Low affinity immunoglobulin epsilon Fc receptor | FCER2 | 1:100 | Block_6 | -2.2768 | 8.1516 | TRUE | 5.418 | 11.3403 | -0.5812 | 0.18352 | 0.67399 | -0.36213 | 0.63683 | 2.1188
Q96RD9 | Fc receptor-like protein 5 | FCRL5 | 1:100 | Block_6 | -4.093 | 5.7593 | TRUE | 8.9416 | 8.6665 | -1.3512 | -0.32367 | 0.88977 | -1.2074 | -0.075431 | 1.0098
Q9BYJ0 | Fibroblast growth factor-binding protein 2 | FGFBP2 | 1:100 | Block_6 | -2.4976 | 5.0894 | TRUE | 5.8538 | 10.7252 | -0.91977 | 0.20051 | 0.99055 | -0.95085 | 0.03662 | 0.85471
P04066 | Tissue alpha-L-fucosidase | FUCA1 | 1:100 | Block_6 | -4.46 | 4.3708 | TRUE | 7.2873 | 9.9721 | -1.7723 | 0.14686 | 1.2064 | -0.79338 | 0.10296 | 0.88118
O15117 | FYN-binding protein 1 | FYB1 | 1:100 | Block_6 | -2.5726 | NA | NA | 12.8302 | 10.5807 | -0.83579 | 1.6678 | 6.0101 | 2.4159 | 5.9608 | 7.4553
Q14393 | Growth arrest-specific protein 6 | GAS6 | 1:100 | Block_6 | -3.7292 | 2.3138 | FALSE | 6.0403 | 8.7789 | -0.16927 | 0.62373 | 1.1225 | 0.081713 | 0.87108 | 1.5988
Q99988 | Growth/differentiation factor 15 | GDF15 | 1:100 | Block_6 | -1.3293 | 8.7455 | TRUE | 6.7471 | 10.7092 | -0.4594 | 0.27763 | 1.2822 | -0.027077 | 1.3987 | 3.0247
P19440 | Glutathione hydrolase 1 proenzyme | GGT1 | 1:100 | Block_6 | -2.5522 | 7.7687 | TRUE | 7.4554 | 8.9802 | -1.471 | -0.54207 | 0.67382 | -0.93924 | -0.20514 | 1.2079
P17900 | Ganglioside GM2 activator | GM2A | 1:100 | Block_6 | -1.09 | 5.5056 | TRUE | 10.03 | 16.6444 | -0.19201 | 0.30838 | 0.76855 | 0.12646 | 0.97893 | 2.3586
P22749 | Granulysin | GNLY | 1:100 | Block_6 | -4.0123 | 9.7257 | TRUE | 5.1912 | 10.4487 | -1.2744 | 0.026961 | 1.1375 | -0.37596 | 0.41279 | 1.7924
P17174 | Aspartate aminotransferase, cytoplasmic | GOT1 | 1:100 | Block_6 | -2.393 | 7.9276 | TRUE | 4.4462 | 7.85 | -0.47746 | 0.11904 | 0.88062 | -0.014682 | 0.58395 | 1.5952
Q9HCN6 | Platelet glycoprotein VI | GP6 | 1:100 | Block_6 | -3.8225 | 7.3221 | TRUE | 7.2136 | 9.6507 | -0.47872 | 1.2928 | 3.2744 | 1.3587 | 3.3346 | 4.4087
P28799 | Progranulin | GRN | 1:100 | Block_6 | -4.43 | 4.7088 | TRUE | 5.5263 | 11.2548 | -0.34138 | 0.27675 | 1.1933 | -0.18349 | 0.51565 | 1.1803
Q02747 | Guanylin | GUCA2A | 1:100 | Block_6 | -4.7191 | 5.8362 | TRUE | 8.2915 | 7.1773 | -0.077103 | 0.58846 | 1.0913 | -0.012975 | 1.1134 | 2.6176
P08236 | Beta-glucuronidase | GUSB | 1:100 | Block_6 | -1.3483 | 7.7026 | TRUE | 7.5177 | 6.3575 | -1.134 | 0.31868 | 1.4375 | -0.21385 | 0.81273 | 2.0632
P12081 | Histidine--tRNA ligase, cytoplasmic | HARS1 | 1:100 | Block_6 | 0.85328 | 10.0789 | TRUE | NA | NA | NA | NA | 2.7249 | NA | 2.7884 | 3.9709
Q8TDQ0 | Hepatitis A virus cellular receptor 2 | HAVCR2 | 1:100 | Block_6 | -0.84396 | 8.1563 | TRUE | 9.1724 | 10.6538 | -0.41519 | 0.22617 | 0.66921 | -0.080012 | 0.8703 | 2.1444
P32754 | 4-hydroxyphenylpyruvate dioxygenase | HPD | 1:100 | Block_6 | -2.5 | NA | NA | 4.9892 | 10.5513 | -1.8688 | -0.87604 | -0.089047 | -1.783 | -0.78901 | 1.1712
Q9Y251 | Heparanase | HPSE | 1:100 | Block_6 | -1 | 10.7975 | FALSE | 6.6065 | 5.9275 | -0.021319 | 2.0996 | 4.4217 | 1.8488 | 4.7042 | 6.2728
P04792 | Heat shock protein beta-1 | HSPB1 | 1:100 | Block_6 | -1.25 | 8.9464 | FALSE | 17.1524 | 9.9865 | -0.31555 | 1.9836 | 5.3719 | 1.9006 | 5.0918 | 6.96
Q9Y4L1 | Hypoxia up-regulated protein 1 | HYOU1 | 1:100 | Block_6 | -2.1189 | 3.11 | TRUE | 4.0574 | 4.9668 | -0.33998 | 0.20506 | 0.70817 | 0.20418 | 0.66126 | 1.0776
P13598 | Intercellular adhesion molecule 2 | ICAM2 | 1:100 | Block_6 | -5.46 | 5.153 | FALSE | 5.8913 | 12.188 | -0.65483 | 0.080632 | 0.62362 | -0.26298 | 0.44298 | 1.0816
P32942 | Intercellular adhesion molecule 3 | ICAM3 | 1:100 | Block_6 | -5 | 4.6249 | TRUE | 7.7708 | 6.7114 | -0.30553 | 0.20265 | 0.68779 | -0.36349 | 0.31846 | 0.83222
P48551 | Interferon alpha/beta receptor 2 | IFNAR2 | 1:100 | Block_6 | -1.1172 | NA | NA | 4.2864 | 10.7878 | -0.36675 | 0.29959 | 0.8201 | 0.074416 | 0.90914 | 1.9736
Q8TDY8 | Immunoglobulin superfamily DCC subclass member 4 | IGDCC4 | 1:100 | Block_6 | -1.71 | 6.0625 | TRUE | 9.7371 | 13.5749 | -0.40989 | 0.10461 | 0.76658 | -0.4871 | 0.11221 | 0.68998
P11717 | Cation-independent mannose-6-phosphate receptor | IGF2R | 1:100 | Block_6 | -3.71 | 4.8876 | TRUE | 4.6076 | 8.2385 | -0.59265 | 0.032968 | 0.65183 | -0.057489 | 0.4029 | 1.061
P08833 | Insulin-like growth factor-binding protein 1 | IGFBP1 | 1:100 | Block_6 | -5.3076 | 6.3372 | TRUE | 4.7083 | 8.3682 | -2.5246 | -0.81118 | 1.4931 | -1.9602 | 0.35852 | 1.7449
P18065 | Insulin-like growth factor-binding protein 2 | IGFBP2 | 1:100 | Block_6 | -4.6149 | 3.5122 | TRUE | 6.7945 | 10.4462 | -1.0398 | 0.20982 | 1.0357 | -0.34279 | 1.0362 | 2.2916
P22692 | Insulin-like growth factor-binding protein 4 | IGFBP4 | 1:100 | Block_6 | -5.2405 | 6.1165 | TRUE | 4.495 | 10.0994 | -0.89051 | 0.14695 | 0.96325 | -0.037747 | 1.0543 | 3.131
Q16270 | Insulin-like growth factor-binding protein 7 | IGFBP7 | 1:100 | Block_6 | -2.1919 | 6.1092 | TRUE | 6.6678 | 8.1991 | -0.40061 | 0.21749 | 0.64561 | 0.13871 | 0.76251 | 1.6754
Q96F46 | Interleukin-17 receptor A | IL17RA | 1:100 | Block_6 | -2.9316 | 9.2888 | TRUE | 7.0933 | 9.7727 | -0.42428 | 0.24841 | 0.60593 | -0.26658 | 0.43732 | 1.2106
P14778 | Interleukin-1 receptor type 1 | IL1R1 | 1:100 | Block_6 | -3.0213 | 7.2713 | TRUE | 5.099 | 10.783 | -0.43563 | 0.20803 | 0.76938 | -0.076381 | 0.48577 | 1.1427
Q9NPH3 | Interleukin-1 receptor accessory protein | IL1RAP | 1:100 | Block_6 | -3.4476 | 6.1018 | TRUE | 4.7156 | 12.7597 | -0.43221 | 0.49278 | 1.1564 | -0.37263 | 0.53772 | 1.302
Q01638 | Interleukin-1 receptor-like 1 | IL1RL1 | 1:100 | Block_6 | -3.17 | 6.4858 | TRUE | 9.0541 | 9.6747 | -1.1432 | -0.18263 | 0.88573 | -0.52224 | 0.4145 | 2.2147
P01589 | Interleukin-2 receptor subunit alpha | IL2RA | 1:100 | Block_6 | -2.99 | 7.9219 | TRUE | 8.6621 | 7.7908 | -0.55319 | 0.15252 | 0.79449 | -0.17335 | 0.86336 | 2.1257
P16871 | Interleukin-7 receptor subunit alpha | IL7R | 1:100 | Block_6 | -3.7055 | 6.0109 | TRUE | 5.7087 | 8.6233 | -0.75075 | 0.17229 | 0.90466 | -0.55483 | 0.41032 | 1.1743
Q9GZP8 | Immortalization up-regulated protein | IMUP | 1:100 | Block_6 | -0.32151 | NA | NA | 14.0092 | 6.9645 | -0.024011 | 0.80124 | 5.4133 | 1.5641 | 4.8981 | 7.3219
P55103 | Inhibin beta C chain | INHBC | 1:100 | Block_6 | -2.2306 | 5.5713 | TRUE | 15.3772 | 9.2566 | -0.54559 | 0.27783 | 0.89431 | -0.14431 | 0.49728 | 1.2487
Q6H9L7 | Isthmin-2 | ISM2 | 1:100 | Block_6 | 0.97705 | 7.86 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P08648 | Integrin alpha-5 | ITGA5 | 1:100 | Block_6 | -1.5943 | 5.6675 | TRUE | 7.3153 | 11.2052 | -0.19207 | 0.25211 | 0.82277 | -0.16129 | 0.52463 | 1.1391
P11215 | Integrin alpha-M | ITGAM | 1:100 | Block_6 | -1.3729 | 7.8641 | TRUE | 6.3976 | 13.4832 | -0.48583 | 0.20517 | 1.1535 | -0.36318 | 0.20713 | 1.0618
P05107 | Integrin beta-2 | ITGB2 | 1:100 | Block_6 | -7.45 | 5.9765 | TRUE | 4.7036 | 9.1153 | -0.61363 | 0.079053 | 0.83585 | -0.41603 | 0.2075 | 0.96429
O95965 | Integrin beta-like protein 1 | ITGBL1 | 1:100 | Block_6 | -2.48 | 8.6235 | TRUE | 7.641 | 5.8605 | -0.49519 | 0.11473 | 1.0085 | 0.10076 | 0.8445 | 2.3574
P07476 | Involucrin | IVL | 1:100 | Block_6 | -1.32 | 7.7895 | TRUE | 6.8187 | 7.5639 | -0.77353 | -0.084708 | 1.1185 | -0.25511 | 0.65322 | 1.5058
Q15726 | Metastasis-suppressor KiSS-1 | KISS1 | 1:100 | Block_6 | 1.679 | 14.2511 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P55268 | Laminin subunit beta-2 | LAMB2 | 1:100 | Block_6 | -3.5 | -1.3798 | FALSE | 7.9059 | 11.7425 | -0.39357 | 0.20869 | 1.0037 | -0.070396 | 0.499 | 1.0943
P11279 | Lysosome-associated membrane glycoprotein 1 | LAMP1 | 1:100 | Block_6 | 0.001 | 6.3628 | TRUE | 18.6316 | 15.3374 | NA | 0.16156 | 0.6626 | NA | 0.46897 | 1.1557
P13796 | Plastin-2 | LCP1 | 1:100 | Block_6 | -2.261 | 6.3468 | TRUE | 9.5239 | 13.2036 | -0.73041 | 0.083148 | 1.0259 | -0.26955 | 0.42766 | 1.3625
P01130 | Low-density lipoprotein receptor | LDLR | 1:100 | Block_6 | -3.27 | 7.5395 | TRUE | 38.5714 | 11.6494 | -0.59169 | 0.33747 | 1.5271 | -0.37648 | 0.77563 | 1.6253
P17931 | Galectin-3 | LGALS3 | 1:100 | Block_6 | 0.25254 | 9.7845 | TRUE | 8.8492 | 11.0276 | 0.26999 | 0.95968 | 1.4237 | 0.82218 | 1.7448 | 3.0759
Q8N149 | Leukocyte immunoglobulin-like receptor subfamily A member 2 | LILRA2 | 1:100 | Block_6 | -3.2983 | 4.6449 | TRUE | 5.409 | 8.4639 | -0.17549 | 0.41022 | 1.1772 | -0.027624 | 0.79683 | 1.4063
Q8N6C8 | Leukocyte immunoglobulin-like receptor subfamily A member 3 | LILRA3 | 1:100 | Block_6 | -1.58 | 3.9103 | FALSE | 6.5926 | 6.8858 | -1.5764 | 0.0073739 | 0.42144 | -0.47706 | 0.11565 | 0.5384
O75023 | Leukocyte immunoglobulin-like receptor subfamily B member 5 | LILRB5 | 1:100 | Block_6 | -3.0147 | 5.0265 | TRUE | 6.501 | 11.4741 | -1.492 | 3.53e-05 | 0.71264 | -1.6892 | 0.20962 | 0.78896
Q14767 | Latent-transforming growth factor beta-binding protein 2 | LTBP2 | 1:100 | Block_6 | -3.3798 | 6.8928 | TRUE | 4.3808 | 8.8102 | -0.38159 | 0.37082 | 1.1481 | 0.098661 | 1.1676 | 2.3536
P02788 | Lactotransferrin | LTF | 1:100 | Block_6 | -2 | NA | NA | 4.4493 | 14.13 | -0.68611 | 0.012834 | 0.96409 | -0.58184 | 0.38045 | 2.0568
Q9H1U4 | Multiple epidermal growth factor-like domains protein 9 | MEGF9 | 1:100 | Block_6 | -7.01 | 5.1713 | TRUE | 5.474 | 5.8553 | -0.41285 | 0.23764 | 0.73171 | -0.080304 | 0.47027 | 0.98317
Q14696 | LRP chaperone MESD | MESD | 1:100 | Block_6 | 0.80755 | 4.8047 | FALSE | NA | NA | NA | 1.7486 | 5.328 | 1.5097 | 5.3813 | 7.2675
P08581 | Hepatocyte growth factor receptor | MET | 1:100 | Block_6 | -4.7317 | 5.0361 | TRUE | 7.7002 | 7.9704 | -0.18609 | 0.30952 | 0.63619 | -0.058594 | 0.4442 | 0.88549
Q13361 | Microfibrillar-associated protein 5 | MFAP5 | 1:100 | Block_6 | -2.5742 | 4.3178 | TRUE | 5.5053 | 11.2983 | -0.2881 | 0.36813 | 0.77861 | -0.29283 | 0.59832 | 1.3987
Q08431 | Lactadherin | MFGE8 | 1:100 | Block_6 | -2.046 | 7.9155 | FALSE | 8.2058 | 15.1671 | -0.59072 | 0.16151 | 1.093 | -0.26513 | 0.49424 | 1.5496
P14174 | Macrophage migration inhibitory factor | MIF | 1:100 | Block_6 | -0.44506 | 8.9277 | TRUE | 6.1157 | 10.7404 | -0.16184 | 1.8539 | 3.6774 | 1.0545 | 3.8241 | 5.3819
P08254 | Stromelysin-1 | MMP3 | 1:100 | Block_6 | -5.1424 | 6.5554 | TRUE | 5.6907 | 11.4971 | -1.2756 | -0.22792 | 1.2887 | -1.129 | 0.55745 | 2.5933
P14780 | Matrix metalloproteinase-9 | MMP9 | 1:100 | Block_6 | -4.84 | 5.7012 | TRUE | 4.0595 | 8.3756 | -0.15789 | 1.3337 | 2.403 | 0.23285 | 1.9655 | 4.3394
P05164 | Myeloperoxidase | MPO | 1:100 | Block_6 | -3.5 | 1.7711 | TRUE | 8.0524 | 20.252 | -1.0723 | -0.40563 | 0.22258 | -0.58539 | 0.11262 | 1.3081
P08118 | Beta-microseminoprotein | MSMB | 1:100 | Block_6 | -0.5 | 6.0523 | TRUE | 9.2024 | 24.1793 | NA | 0.036769 | 0.76214 | NA | 0.058246 | 1.3298
P15941 | Mucin-1 | MUC1 | 1:100 | Block_6 | -0.5 | NA | NA | 9.3247 | 16.8817 | NA | 0.21289 | 1.1034 | NA | 0.35955 | 2.3865
Q02817 | Mucin-2 | MUC2 | 1:100 | Block_6 | -1.6206 | 10.9859 | TRUE | 5.086 | 6.7156 | -1.4186 | -0.23667 | 1.1828 | -1.4979 | 0.75972 | 2.5471
P35579 | Myosin-9 | MYH9 | 1:100 | Block_6 | 0.55404 | 9.7374 | TRUE | NA | NA | NA | 0.88331 | 3.9505 | 0.97885 | 2.4402 | 4.8721
Q99972 | Myocilin | MYOC | 1:100 | Block_6 | -4.3659 | 2.7506 | FALSE | 11.6795 | 11.5463 | -1.2377 | 0.17294 | 0.93717 | -0.77215 | -0.023623 | 1.0006
O15394 | Neural cell adhesion molecule 2 | NCAM2 | 1:100 | Block_6 | -3.0482 | 5.8796 | TRUE | 5.2386 | 11.178 | -0.20518 | 0.23061 | 0.54123 | -0.21731 | 0.3464 | 1.0107
Q14112 | Nidogen-2 | NID2 | 1:100 | Block_6 | -1.2412 | 5.6639 | FALSE | 8.5853 | 11.7655 | 0.078908 | 1.1366 | 3.0125 | 1.0555 | 2.2976 | 3.5829
Q15155 | Nodal modulator 1 | NOMO1 | 1:100 | Block_6 | -3.8334 | 6.0192 | TRUE | 7.2639 | 9.2821 | -0.65301 | -0.0017408 | 0.62261 | -0.38822 | 0.29007 | 0.85286
P46531 | Neurogenic locus notch homolog protein 1 | NOTCH1 | 1:100 | Block_6 | -5.3615 | 5.2686 | TRUE | 4.2222 | 9.2309 | -0.2293 | 0.3561 | 0.71819 | 0.0045543 | 0.52305 | 1.0524
Q04721 | Neurogenic locus notch homolog protein 2 | NOTCH2 | 1:100 | Block_6 | -4.8088 | 6.2391 | TRUE | 7.1458 | 6.8816 | -0.47273 | 0.21682 | 0.79 | -0.30652 | 0.36863 | 0.92234
Q9UM47 | Neurogenic locus notch homolog protein 3 | NOTCH3 | 1:100 | Block_6 | -3.35 | 6.5831 | TRUE | 7.6004 | 15.244 | -0.74954 | 0.15253 | 0.7935 | -0.32001 | 0.31933 | 1.2293
P61916 | NPC intracellular cholesterol transporter 2 | NPC2 | 1:100 | Block_6 | -3.64 | 4.4816 | TRUE | 6.3477 | 15.1108 | -0.35919 | 0.18485 | 0.78696 | 0.17199 | 0.90159 | 2.1659
Q15818 | Neuronal pentraxin-1 | NPTX1 | 1:100 | Block_6 | -2.5143 | 8.1825 | FALSE | 24.9316 | 9.0121 | -0.26569 | 0.86127 | 1.5499 | 0.023773 | 0.64776 | 1.7198
P47972 | Neuronal pentraxin-2 | NPTX2 | 1:100 | Block_6 | -2.0347 | 8.4349 | TRUE | 5.7329 | 4.9573 | -0.25842 | 0.29965 | 0.99295 | -0.010207 | 0.7295 | 1.6735
O60462 | Neuropilin-2 | NRP2 | 1:100 | Block_6 | -1.8703 | 5.9255 | TRUE | 9.4782 | 7.6074 | -0.42724 | 0.24024 | 0.78936 | 0.021616 | 0.57772 | 1.2097
Q9UKK9 | ADP-sugar pyrophosphatase | NUDT5 | 1:100 | Block_6 | -2.16 | 9.4963 | TRUE | 30.7009 | 24.9643 | -0.34367 | 0.96837 | 2.3095 | 1.2006 | 2.2852 | 3.6875
Q86UD1 | Out at first protein homolog | OAF | 1:100 | Block_6 | -2.987 | NA | NA | 5.9019 | 13.0948 | -0.23599 | 0.41876 | 1.1744 | 0.15325 | 0.9877 | 1.6582
P20774 | Mimecan | OGN | 1:100 | Block_6 | -3.7465 | 7.0539 | TRUE | 4.393 | 11.7246 | 0.16085 | 0.69502 | 1.6246 | 0.24329 | 1.5087 | 3.1598
P19021 | Peptidyl-glycine alpha-amidating monooxygenase | PAM | 1:100 | Block_6 | -4.811 | 5.532 | TRUE | 18.4556 | 8.8582 | -0.59147 | 0.16574 | 0.51066 | -0.41464 | 0.30158 | 0.9869
Q6UXH9 | Inactive serine protease PAMR1 | PAMR1 | 1:100 | Block_6 | -3.2094 | 4.5569 | TRUE | 7.0619 | 8.1658 | -0.32083 | 0.24268 | 0.75365 | -0.15106 | 0.44915 | 1.1032
Q99497 | Parkinson disease protein 7 | PARK7 | 1:100 | Block_6 | -2 | 7.2517 | TRUE | 10.4053 | 8.7368 | -0.44759 | 1.6597 | 3.5426 | 1.298 | 3.6662 | 5.0394
Q8NBP7 | Proprotein convertase subtilisin/kexin type 9 | PCSK9 | 1:100 | Block_6 | -2.64 | 3.6688 | FALSE | 7.171 | 20.1119 | -0.97537 | -0.15766 | 0.61677 | -0.55147 | 0.3831 | 1.8851
P04085 | Platelet-derived growth factor subunit A | PDGFA | 1:100 | Block_6 | -1.7201 | 9.9966 | TRUE | 4.301 | 8.7429 | -0.24441 | 1.3911 | 2.8295 | 0.90016 | 2.9653 | 4.4344
P09619 | Platelet-derived growth factor receptor beta | PDGFRB | 1:100 | Block_6 | -1.8255 | 4.7629 | TRUE | 4.7858 | 11.7794 | -0.22786 | 0.43462 | 1.0202 | -0.2233 | 0.46804 | 1.053
P13667 | Protein disulfide-isomerase A4 | PDIA4 | 1:100 | Block_6 | -0.91926 | 7.2088 | TRUE | 12.1216 | 10.4665 | -0.31145 | 0.85978 | 2.6845 | 1.1302 | 3.1502 | 4.503
O00151 | PDZ and LIM domain protein 1 | PDLIM1 | 1:100 | Block_6 | 0.88539 | NA | NA | NA | NA | NA | 1.2346 | 5.6804 | 1.7404 | 5.7083 | 7.6446
Q5VY43 | Platelet endothelial aggregation receptor 1 | PEAR1 | 1:100 | Block_6 | -4 | 5.5915 | TRUE | 4.7116 | 6.7194 | -0.21471 | 0.32734 | 0.8246 | 0.19484 | 0.74624 | 1.3073
P30086 | Phosphatidylethanolamine-binding protein 1 | PEBP1 | 1:100 | Block_6 | -2.6 | 7.4203 | TRUE | 6.4924 | 12.6969 | 0.0076217 | 1.1769 | 2.2494 | 1.2035 | 2.8776 | 4.0334
P16284 | Platelet endothelial cell adhesion molecule | PECAM1 | 1:100 | Block_6 | -6.2738 | 5.797 | TRUE | 3.7202 | 8.8525 | -0.35536 | 0.56036 | 1.3584 | 0.4219 | 1.0175 | 1.9095
O75594 | Peptidoglycan recognition protein 1 | PGLYRP1 | 1:100 | Block_6 | -4.59 | 4.8979 | TRUE | 6.5172 | 7.7747 | -0.73274 | -0.0003413 | 0.89122 | -0.45161 | 0.66969 | 2.5153
P19957 | Elafin | PI3 | 1:100 | Block_6 | -1.78 | 5.3095 | TRUE | 4.443 | 9.6539 | -0.80381 | -0.072442 | 0.75316 | -0.4094 | 0.57014 | 1.7466
P01833 | Polymeric immunoglobulin receptor | PIGR | 1:100 | Block_6 | -3.5 | 6.6317 | TRUE | 3.8782 | 8.537 | -0.78014 | 0.078218 | 1.2487 | -0.60326 | 0.31415 | 1.2497
Q96FE7 | Phosphoinositide-3-kinase-interacting protein 1 | PIK3IP1 | 1:100 | Block_6 | -3.3263 | 7.491 | TRUE | 4.2312 | 11.1128 | -0.13977 | 0.33571 | 0.97699 | 0.21197 | 1.037 | 2.2607
Q9UKJ1 | Paired immunoglobulin-like type 2 receptor alpha | PILRA | 1:100 | Block_6 | -3.8305 | 8.4815 | TRUE | 6.8425 | 10.5852 | -0.42513 | 0.34186 | 0.89426 | -0.36476 | 0.8581 | 1.9559
P04054 | Phospholipase A2 | PLA2G1B | 1:100 | Block_6 | -3.5602 | 8.1814 | TRUE | 12.1507 | 7.627 | -0.61507 | 0.07716 | 0.93963 | -0.60658 | 0.19699 | 1.3563
P14555 | Phospholipase A2, membrane associated | PLA2G2A | 1:100 | Block_6 | -3.55 | 5.525 | TRUE | 34.1975 | 43.17 | -0.78512 | -0.14912 | 0.89712 | -1.2866 | -0.16732 | 1.9395
Q13093 | Platelet-activating factor acetylhydrolase | PLA2G7 | 1:100 | Block_6 | -2.8137 | 2.9306 | FALSE | 7.6256 | 10.2196 | -0.41004 | 0.30788 | 0.83531 | -0.32276 | 0.29149 | 0.97173
P00749 | Urokinase-type plasminogen activator | PLAU | 1:100 | Block_6 | -3.6936 | 8.4762 | TRUE | 4.1335 | 9.778 | -0.23394 | 0.37976 | 0.88719 | 0.016748 | 0.54276 | 1.0818
P55058 | Phospholipid transfer protein | PLTP | 1:100 | Block_6 | -5.4326 | 6.3155 | TRUE | 5.3716 | 15.8716 | 0.089617 | 0.71908 | 1.3157 | -0.13738 | 0.69142 | 1.3627
O15031 | Plexin-B2 | PLXNB2 | 1:100 | Block_6 | -5.27 | 4.6228 | TRUE | 37.2026 | 28.1459 | -0.71753 | 0.13878 | 0.91787 | 0.01639 | 0.55477 | 1.2431
Q15063 | Periostin | POSTN | 1:100 | Block_6 | -4 | 6.1232 | FALSE | 7.9089 | 9.9635 | 0.048624 | 0.54184 | 1.1625 | 0.3532 | 1.2135 | 2.0705
P23284 | Peptidyl-prolyl cis-trans isomerase B | PPIB | 1:100 | Block_6 | -0.12108 | 8.0039 | TRUE | NA | 1.7351 | NA | 1.0015 | 4.3048 | 1.4894 | 4.4143 | 5.8029
P30405 | Peptidyl-prolyl cis-trans isomerase F, mitochondrial | PPIF | 1:100 | Block_6 | 0.86695 | 6.7943 | FALSE | NA | NA | NA | NA | 1.0546 | NA | NA | 1.106
Q96QH2 | PML-RARA-regulated adapter molecule 1 | PRAM1 | 1:100 | Block_6 | -3 | NA | NA | 8.0927 | 12.1756 | -2.2241 | -1.2691 | 0.761 | -1.7981 | -0.40367 | 4.7427
Q96NZ9 | Proline-rich acidic protein 1 | PRAP1 | 1:100 | Block_6 | -3.4007 | 4.9397 | TRUE | 5.377 | 12.8222 | -0.16559 | 0.81123 | 1.6146 | 0.10636 | 1.0137 | 1.896
P42785 | Lysosomal Pro-X carboxypeptidase | PRCP | 1:100 | Block_6 | -4 | 4.527 | TRUE | 4.9542 | 8.5608 | -0.63518 | 0.030597 | 0.8464 | -0.17249 | 0.48163 | 1.3045
P13727 | Bone marrow proteoglycan | PRG2 | 1:100 | Block_6 | -2.517 | 6.7206 | TRUE | 4.7876 | 11.6053 | -0.039954 | 0.87339 | 2.2748 | 0.01621 | 1.5856 | 2.8174
Q9Y2Y8 | Proteoglycan 3 | PRG3 | 1:100 | Block_6 | -3.7268 | 8.6105 | TRUE | 5.0995 | 8.651 | -0.41115 | 0.15694 | 1.7338 | -0.68173 | 0.54953 | 1.5099
P24158 | Myeloblastin | PRTN3 | 1:100 | Block_6 | -3.78 | 6.2747 | TRUE | 8.5717 | 14.7481 | -0.77061 | 0.1402 | 1.0067 | -0.15986 | 0.59773 | 1.8089
P07602 | Prosaposin | PSAP | 1:100 | Block_6 | -3.04 | 4.6205 | TRUE | 7.5172 | 11.6751 | -0.19405 | 0.45616 | 0.88357 | 0.65111 | 1.2071 | 2.2168
P08575 | Receptor-type tyrosine-protein phosphatase C | PTPRC | 1:100 | Block_6 | 0.14589 | 6.1999 | TRUE | 4.9432 | 14.8025 | NA | 0.28706 | 0.73026 | NA | 0.37125 | 0.92962
P10586 | Receptor-type tyrosine-protein phosphatase F | PTPRF | 1:100 | Block_6 | -3.7619 | 5.914 | TRUE | 5.1904 | 9.6487 | -0.26737 | 0.30921 | 0.84711 | 0.055968 | 0.64619 | 1.2907
Q13332 | Receptor-type tyrosine-protein phosphatase S | PTPRS | 1:100 | Block_6 | -4.0318 | 5.6995 | TRUE | 4.7289 | 7.5531 | -0.5294 | 0.14604 | 0.73298 | -0.043387 | 0.50029 | 1.0423
Q16769 | Glutaminyl-peptide cyclotransferase | QPCT | 1:100 | Block_6 | -5.44 | 0.84103 | FALSE | 6.5971 | 7.781 | -0.078617 | 0.45184 | 1.0025 | 0.2256 | 0.79252 | 1.4437
P48304 | Lithostathine-1-beta | REG1B | 1:100 | Block_6 | -3.4322 | 5.794 | TRUE | 6.2967 | 1.6636 | -0.68655 | 0.04044 | 0.80303 | -0.14309 | 0.89562 | 2.3092
Q06141 | Regenerating islet-derived protein 3-alpha | REG3A | 1:100 | Block_6 | -3.7493 | 4.793 | TRUE | 5.0771 | 10.4056 | -1.2611 | -0.35111 | 0.1287 | -0.76607 | 0.60992 | 2.271
Q969Z4 | Tumor necrosis factor receptor superfamily member 19L | RELT | 1:100 | Block_6 | -4.95 | 7.8775 | TRUE | 6.4674 | 6.4915 | -0.39817 | 0.11186 | 0.64082 | 0.14869 | 0.89896 | 2.1349
Q9HD89 | Resistin | RETN | 1:100 | Block_6 | -6.09 | 6.9754 | TRUE | 4.4331 | 8.0899 | -0.66651 | 0.065975 | 0.70726 | -0.30192 | 0.7777 | 2.1267
O00584 | Ribonuclease T2 | RNASET2 | 1:100 | Block_6 | -4.77 | 5.8944 | TRUE | 4.4138 | 7.9621 | -0.3019 | 0.4496 | 0.75838 | 0.48216 | 0.99809 | 1.7368
Q6UWP8 | Suprabasin | SBSN | 1:100 | Block_6 | -1.31 | 3.4487 | TRUE | 8.815 | 24.9747 | -0.59691 | -0.030809 | 0.43581 | -0.49254 | 0.13937 | 0.93479
Q14162 | Scavenger receptor class F member 1 | SCARF1 | 1:100 | Block_6 | -3.2 | 6.5104 | TRUE | 6.598 | 15.9444 | -0.24609 | 0.58213 | 1.1908 | 0.62069 | 1.4434 | 2.3151
P16581 | E-selectin | SELE | 1:100 | Block_6 | -6.4284 | 4.7482 | TRUE | 3.6713 | 7.4478 | -1.3351 | -0.28399 | 0.74498 | -0.87528 | -0.017854 | 1.1706
P14151 | L-selectin | SELL | 1:100 | Block_6 | -5.5011 | NA | NA | 5.9145 | 4.324 | -0.31333 | 0.13932 | 0.45078 | -0.30913 | 0.19052 | 0.58037
O75326 | Semaphorin-7A | SEMA7A | 1:100 | Block_6 | -4.181 | 6.3999 | TRUE | 4.5827 | 5.9781 | -0.25702 | 0.26739 | 0.71376 | -0.029859 | 0.50223 | 1.2369
Q86U17 | Serpin A11 | SERPINA11 | 1:100 | Block_6 | -3.4205 | 4.9904 | TRUE | 6.0395 | 11.2177 | 0.29833 | 1.489 | 2.1961 | 0.077195 | 1.3516 | 2.3192
P30740 | Leukocyte elastase inhibitor | SERPINB1 | 1:100 | Block_6 | -1.31 | 7.8872 | FALSE | 8.366 | 12.9896 | -0.82462 | 0.39391 | 2.1414 | 0.52478 | 2.5451 | 4.8181
P05121 | Plasminogen activator inhibitor 1 | SERPINE1 | 1:100 | Block_6 | -4.7082 | -2.4601 | FALSE | 4.1294 | 5.5554 | -1.4349 | 0.19491 | 1.6262 | 0.1878 | 1.8687 | 2.9142
Q8N114 | Protein shisa-5 | SHISA5 | 1:100 | Block_6 | -0.67025 | 7.6259 | TRUE | 11.5131 | 11.9736 | -0.14971 | 0.24732 | 0.70785 | 0.15668 | 0.87074 | 2.2838
O15389 | Sialic acid-binding Ig-like lectin 5 | SIGLEC5 | 1:100 | Block_6 | -3.5501 | 5.2245 | TRUE | 10.9397 | 17.2597 | -1.0052 | 0.032834 | 0.97509 | -0.79596 | 0.9454 | 2.2739
P78324 | Tyrosine-protein phosphatase non-receptor type substrate 1 | SIRPA | 1:100 | Block_6 | -5.67 | 4.7771 | TRUE | 4.1601 | 6.8878 | -0.42177 | 0.28564 | 0.85737 | 0.0052824 | 0.67182 | 1.396
P02730 | Band 3 anion transport protein | SLC4A1 | 1:100 | Block_6 | 0.25 | 10.2063 | TRUE | NA | NA | NA | 2.9911 | 4.8678 | NA | 2.1471 | 4.3796
P03973 | Antileukoproteinase | SLPI | 1:100 | Block_6 | -4.5 | NA | NA | 4.6752 | 13.6459 | -0.86574 | 0.15618 | 0.86137 | -0.35256 | 0.51344 | 1.5183
P55000 | Secreted Ly-6/uPAR-related protein 1 | SLURP1 | 1:100 | Block_6 | -0.59 | 4.5201 | TRUE | 11.6996 | 18.7517 | -0.42861 | 0.094611 | 0.46138 | -0.15815 | 0.37887 | 1.4383
P04179 | Superoxide dismutase [Mn], mitochondrial | SOD2 | 1:100 | Block_6 | -4.0591 | 5.3766 | FALSE | 5.0386 | 10.3262 | -0.64021 | 0.33097 | 1.526 | 0.017763 | 0.97961 | 2.3219
P00995 | Serine protease inhibitor Kazal-type 1 | SPINK1 | 1:100 | Block_6 | -2.8977 | 5.8511 | TRUE | 5.5169 | 9.6187 | -1.1386 | -0.046938 | 0.58282 | -0.50565 | 0.81655 | 2.45
Q9NQ38 | Serine protease inhibitor Kazal-type 5 | SPINK5 | 1:100 | Block_6 | -1.2517 | 7.1974 | TRUE | 6.0051 | 9.2185 | -0.77312 | -0.31495 | 0.45981 | -0.58933 | 0.087007 | 0.85755
O43278 | Kunitz-type protease inhibitor 1 | SPINT1 | 1:100 | Block_6 | -2.3305 | 6.5888 | TRUE | 8.1706 | 14.6068 | -0.38014 | 0.041023 | 0.69106 | -0.37482 | 0.3273 | 1.0436
P10451 | Osteopontin | SPP1 | 1:100 | Block_6 | -3.4456 | 4.1163 | TRUE | 5.2637 | 10.3102 | -0.17001 | 0.94576 | 1.3919 | 0.51987 | 1.4382 | 2.4581
P15907 | Beta-galactoside alpha-2,6-sialyltransferase 1 | ST6GAL1 | 1:100 | Block_6 | -4.7183 | 6.0758 | FALSE | 6.0361 | 8.5588 | -0.67875 | 0.22611 | 0.95988 | 0.17401 | 0.87448 | 1.4446
Q8WWQ8 | Stabilin-2 | STAB2 | 1:100 | Block_6 | -4.6053 | 7.2139 | FALSE | 6.8604 | 7.1412 | -0.50159 | 0.1727 | 0.72102 | -0.16689 | 0.44992 | 1.1522
O76061 | Stanniocalcin-2 | STC2 | 1:100 | Block_6 | -3.5239 | 6.0061 | TRUE | 4.1266 | 4.8052 | -0.3606 | 0.62053 | 1.214 | -0.0531 | 1.1272 | 1.8507
P31948 | Stress-induced-phosphoprotein 1 | STIP1 | 1:100 | Block_6 | 0.29268 | 9.5883 | TRUE | NA | NA | 0.2972 | 1.1426 | 2.9697 | 1.0897 | 3.3092 | 4.6529
Q9UGT4 | Sushi domain-containing protein 2 | SUSD2 | 1:100 | Block_6 | -2.5626 | 7.4824 | TRUE | 7.1284 | 8.769 | -0.52083 | 0.24351 | 0.97647 | -0.2884 | 0.39263 | 1.1434
P37837 | Transaldolase | TALDO1 | 1:100 | Block_6 | -1.7532 | 5.9301 | TRUE | 5.0292 | 6.6459 | -0.20565 | 0.61958 | 1.5498 | 0.81288 | 1.9691 | 3.0562
P04155 | Trefoil factor 1 | TFF1 | 1:100 | Block_6 | -2.3437 | 7.6432 | TRUE | 6.1922 | 8.8603 | -1.7535 | -0.84638 | 0.64944 | -0.95977 | 0.41157 | 2.6873
Q03403 | Trefoil factor 2 | TFF2 | 1:100 | Block_6 | -3.98 | 7.9716 | TRUE | 12.5322 | 14.7642 | -1.4898 | -0.29256 | 0.96873 | -0.84812 | 0.10008 | 1.6538
Q03167 | Transforming growth factor beta receptor type 3 | TGFBR3 | 1:100 | Block_6 | -4.5 | 6.2866 | TRUE | 8.2509 | 6.6533 | -0.1585 | 1.0389 | 1.5726 | -0.70755 | 1.2863 | 2.0933
P35442 | Thrombospondin-2 | THBS2 | 1:100 | Block_6 | -3.1148 | 8.2786 | FALSE | 4.6982 | 8.6075 | -0.35057 | 0.24062 | 0.92347 | 0.098715 | 0.91494 | 2.5656
P35590 | Tyrosine-protein kinase receptor Tie-1 | TIE1 | 1:100 | Block_6 | -5.0504 | 5.0356 | TRUE | 6.202 | 7.9657 | -0.44607 | 0.15844 | 0.84554 | -0.1218 | 0.51714 | 1.143
Q99727 | Metalloproteinase inhibitor 4 | TIMP4 | 1:100 | Block_6 | -1.4863 | 6.7654 | TRUE | 5.6778 | 9.2174 | -0.7424 | 0.088972 | 0.77688 | -0.65196 | 0.21926 | 1.263
P63313 | Thymosin beta-10 | TMSB10 | 1:100 | Block_6 | -1.8572 | 9.0579 | TRUE | 3.226 | 12.829 | -0.9188 | 1.1658 | 2.5651 | 0.29821 | 4.0653 | 6.0389
O14798 | Tumor necrosis factor receptor superfamily member 10C | TNFRSF10C | 1:100 | Block_6 | -2.1515 | 5.2462 | TRUE | 6.5371 | 14.6016 | -0.97275 | 0.013406 | 0.63935 | -0.66928 | 0.21747 | 0.92902
Q02223 | Tumor necrosis factor receptor superfamily member 17 | TNFRSF17 | 1:100 | Block_6 | -6 | 8.644 | TRUE | 5.3117 | 13.1494 | -0.52856 | 0.080744 | 0.76647 | -0.81697 | 0.21156 | 1.3157
P19438 | Tumor necrosis factor receptor superfamily member 1A | TNFRSF1A | 1:100 | Block_6 | -2.3935 | 9.1726 | TRUE | 22.4854 | 10.4499 | -0.51762 | 0.11513 | 0.66307 | 0.03478 | 0.98827 | 2.4138
P20333 | Tumor necrosis factor receptor superfamily member 1B | TNFRSF1B | 1:100 | Block_6 | -6.1 | 6.8555 | TRUE | 5.3044 | 10.0382 | -0.6219 | 0.20199 | 0.735 | 0.037024 | 1.0075 | 2.4349
Q9Y275 | Tumor necrosis factor ligand superfamily member 13B | TNFSF13B | 1:100 | Block_6 | -2.4107 | 8.6184 | TRUE | 7.07 | 5.9956 | -0.27345 | 0.503 | 0.96427 | -0.034958 | 0.82901 | 2.0649
P32971 | Tumor necrosis factor ligand superfamily member 8 | TNFSF8 | 1:100 | Block_6 | 0.33 | 10.3144 | TRUE | 3.6854 | 27.0653 | NA | NA | 0.8467 | NA | NA | 1.079
Q9UQP3 | Tenascin-N | TNN | 1:100 | Block_6 | -1.998 | 7.5837 | TRUE | 8.6103 | 8.8289 | -0.34655 | 0.51841 | 1.171 | -0.017663 | 0.75749 | 1.4126
P22105 | Tenascin-X | TNXB | 1:100 | Block_6 | -3.07 | 3.7908 | TRUE | 4.5402 | 12.3724 | -0.4906 | 0.052424 | 0.67792 | -0.48286 | 0.12733 | 0.67497
O43280 | Trehalase | TREH | 1:100 | Block_6 | -2.2777 | 7.9221 | FALSE | 5.7558 | 8.8052 | NA | -0.24006 | 1.4317 | -1.8461 | 0.24627 | 1.2213
Q5T2D2 | Trem-like transcript 2 protein | TREML2 | 1:100 | Block_6 | -2.8907 | 7.0948 | TRUE | 6.2255 | 9.3939 | -0.40192 | 0.38301 | 1.1381 | 0.17941 | 1.008 | 1.692
Q6IBS0 | Twinfilin-2 | TWF2 | 1:100 | Block_6 | 0.12239 | 10.5769 | TRUE | 2.6746 | 2.1311 | NA | 1.2819 | 3.7853 | 0.54951 | 4.1701 | 5.7959
Q16881 | Thioredoxin reductase 1, cytoplasmic | TXNRD1 | 1:100 | Block_6 | -0.20445 | 8.3348 | TRUE | 5.9368 | 7.7504 | -0.14223 | 0.52118 | 1.3696 | 0.532 | 1.6603 | 2.6483
P07911 | Uromodulin | UMOD | 1:100 | Block_6 | -2.3111 | 5.2541 | TRUE | 4.5892 | 7.8059 | -0.57614 | 0.24259 | 0.98937 | -1.4614 | 0.083506 | 0.98599
Q9Y279 | V-set and immunoglobulin domain-containing protein 4 | VSIG4 | 1:100 | Block_6 | -4.4531 | 6.5405 | TRUE | 5.5812 | 13.5806 | -0.74853 | 0.043971 | 0.67217 | -0.13942 | 1.404 | 3.1494
P04275 | von Willebrand factor | VWF | 1:100 | Block_6 | -2.7 | 6.0125 | FALSE | 26.0033 | 20.1833 | -1.4752 | 0.18883 | 1.7817 | -1.5457 | 0.69551 | 2.4455
P23381 | Tryptophan--tRNA ligase, cytoplasmic | WARS1 | 1:100 | Block_6 | -3.5647 | 6.7729 | TRUE | 5.0659 | 10.4029 | -0.52443 | 0.1797 | 0.76026 | -0.047995 | 0.72632 | 1.7008
P55808 | Glycoprotein Xg | XG | 1:100 | Block_6 | -2.35 | 9.512 | TRUE | 10.475 | 12.7337 | -0.083231 | 0.75327 | 1.296 | 0.29175 | 1.0792 | 2.3262
Q9NQ94 | APOBEC1 complementation factor | A1CF | 1:10 | Block_5 | 0.75633 | 8.4929 | FALSE | NA | NA | NA | NA | 0.83095 | NA | NA | 0.80673
P45954 | Short/branched chain specific acyl-CoA dehydrogenase, mitochondrial | ACADSB | 1:10 | Block_5 | 0.54503 | 8.8633 | FALSE | NA | 7.7934 | NA | 0.6365 | 2.6487 | NA | 1.1633 | 2.923
P16112 | Aggrecan core protein | ACAN | 1:10 | Block_5 | -2.4727 | 6.0289 | TRUE | 10.0799 | 3.7241 | -0.17962 | 0.26025 | 0.67903 | -0.15057 | 0.60693 | 1.0718
P22303 | Acetylcholinesterase | ACHE | 1:10 | Block_5 | -2.3843 | 5.4247 | TRUE | 8.5495 | 12.014 | -0.47708 | 0.22939 | 0.74171 | -0.28762 | 0.32521 | 1.1147
Q15067 | Peroxisomal acyl-coenzyme A oxidase 1 | ACOX1 | 1:10 | Block_5 | 0.50771 | 5.6547 | FALSE | NA | NA | NA | 0.77046 | 1.8772 | NA | 1.615 | 3.1726
Q03154 | Aminoacylase-1 | ACY1 | 1:10 | Block_5 | -3.9365 | 2.6303 | FALSE | 7.0128 | 4.6375 | -1.4179 | -0.2881 | 0.83218 | -0.82265 | 0.13241 | 1.4988
P07311 | Acylphosphatase-1 | ACYP1 | 1:10 | Block_5 | -0.94176 | 7.5406 | TRUE | 15.2261 | 8.3249 | 0.059495 | 1.545 | 3.3575 | 1.6849 | 3.3695 | 4.4551
P78325 | Disintegrin and metalloproteinase domain-containing protein 8 | ADAM8 | 1:10 | Block_5 | -2.118 | 7.0536 | TRUE | 10.4845 | 8.9886 | -0.08499 | 0.75725 | 1.281 | 0.25886 | 1.0333 | 1.9912
Q9UP79 | A disintegrin and metalloproteinase with thrombospondin motifs 8 | ADAMTS8 | 1:10 | Block_5 | -4.622 | 5.272 | TRUE | 10.8215 | 5.0085 | -0.29709 | 0.12026 | 0.89712 | -0.90326 | 0.054918 | 0.67195
Q6ZMM2 | ADAMTS-like protein 5 | ADAMTSL5 | 1:10 | Block_5 | -2.27 | 6.7653 | TRUE | 24.6349 | 13.7854 | -0.86974 | 0.14729 | 1.252 | -0.86843 | 0.0052491 | 0.91924
P35611 | Alpha-adducin | ADD1 | 1:10 | Block_5 | 0.00071384 | 8.2498 | TRUE | 7.3986 | 18.3334 | 1.0251 | 3.2103 | 4.6015 | 2.4111 | 3.9762 | 5.2226
P35612 | Beta-adducin | ADD2 | 1:10 | Block_5 | 0.15407 | 6.228 | TRUE | NA | NA | 0.27393 | 1.2327 | 1.9619 | 0.25805 | 0.99534 | 1.8659
Q6QNK2 | Adhesion G-protein coupled receptor D1 | ADGRD1 | 1:10 | Block_5 | -3.93 | 5.767 | TRUE | 9.1184 | 10.3872 | -0.48073 | -0.017085 | 0.67025 | -0.37849 | 0.43503 | 1.6995
Q9UHX3 | Adhesion G protein-coupled receptor E2 | ADGRE2 | 1:10 | Block_5 | -3 | 4.2559 | TRUE | 7.8828 | 8.5008 | -0.54959 | 0.35572 | 0.84524 | -0.14676 | 0.55717 | 1.28
Q8IZF2 | Adhesion G protein-coupled receptor F5 | ADGRF5 | 1:10 | Block_5 | -2.84 | 4.1507 | FALSE | 10.7044 | 4.4582 | -0.84981 | 0.19242 | 0.60157 | 0.024437 | 0.44187 | 0.84505
Q8IZP9 | Adhesion G-protein coupled receptor G2 | ADGRG2 | 1:10 | Block_5 | -5.448 | 6.0681 | TRUE | 8.2875 | 2.7721 | -0.53219 | 0.050173 | 0.541 | -0.72426 | 0.074883 | 0.76406
P00325 | All-trans-retinol dehydrogenase [NAD(+)] ADH1B | ADH1B | 1:10 | Block_5 | -2.5 | 5.1016 | FALSE | 17.9574 | 5.4433 | -1.9137 | -0.50152 | 1.8159 | -1.2025 | -0.01381 | 1.9556
Q8IUX7 | Adipocyte enhancer-binding protein 1 | AEBP1 | 1:10 | Block_5 | -1.4956 | NA | NA | 13.4991 | 11.5311 | -0.081184 | 0.68577 | 1.2566 | 0.60396 | 1.3733 | 2.683
Q8N556 | Actin filament-associated protein 1 | AFAP1 | 1:10 | Block_5 | 0.8126 | 8.2865 | TRUE | NA | NA | NA | NA | 5.1303 | NA | 2.0764 | 7.3786
Q15109 | Advanced glycosylation end product-specific receptor | AGER | 1:10 | Block_5 | -4.2451 | 6.03 | TRUE | 10.4401 | 6.9253 | -0.18925 | 0.30151 | 0.99348 | 0.02765 | 0.88794 | 1.8721
O00468 | Agrin | AGRN | 1:10 | Block_5 | -4.4222 | 5.0983 | TRUE | 14.6117 | 8.7141 | -0.34674 | 0.31458 | 0.69558 | 0.03033 | 0.80483 | 1.9853
P21549 | Alanine--glyoxylate aminotransferase | AGXT | 1:10 | Block_5 | -2.1299 | 4.294 | TRUE | 7.0515 | 4.3819 | -1.8256 | -1.1219 | -0.22206 | -1.8161 | -0.83702 | 0.81069
P23526 | Adenosylhomocysteinase | AHCY | 1:10 | Block_5 | -0.14 | 8.7899 | TRUE | 14.7782 | 7.9024 | 0.46113 | 1.8279 | 2.9135 | 2.0758 | 3.1114 | 4.3597
P00568 | Adenylate kinase isoenzyme 1 | AK1 | 1:10 | Block_5 | -0.56094 | 7.5465 | TRUE | 11.9762 | 15.8675 | 0.011887 | 3.0159 | 5.0132 | 2.2488 | 4.5297 | 5.6966
Q02952 | A-kinase anchor protein 12 | AKAP12 | 1:10 | Block_5 | -1.0046 | 6.3789 | TRUE | 8.1466 | 4.4753 | -0.40737 | -0.014065 | 0.52076 | -0.31761 | 0.2969 | 0.83895
P10696 | Alkaline phosphatase, germ cell type | ALPG | 1:10 | Block_5 | -5.426 | 8.3299 | TRUE | 8.6279 | 5.5297 | -2.3652 | -0.52703 | 1.3753 | -3.6377 | -2.387 | -0.696
Q4VCS5 | Angiomotin | AMOT | 1:10 | Block_5 | -0.32533 | 8.5747 | TRUE | 10.3814 | 10.5656 | -0.0627 | 0.48769 | 1.4394 | -0.035771 | 0.76412 | 1.5478
Q15389 | Angiopoietin-1 | ANGPT1 | 1:10 | Block_5 | -0.725 | 8.1357 | TRUE | 7.2993 | 7.3397 | 0.3663 | 1.6053 | 2.9553 | 1.1027 | 3.1428 | 4.4322
O95841 | Angiopoietin-related protein 1 | ANGPTL1 | 1:10 | Block_5 | -1.5682 | 5.2253 | FALSE | 9.4231 | 9.2538 | 0.044053 | 0.46974 | 0.97935 | -0.2245 | 0.46866 | 1.2343
Q9UKU9 | Angiopoietin-related protein 2 | ANGPTL2 | 1:10 | Block_5 | -3.7719 | 0.41683 | FALSE | 12.0445 | 7.753 | -0.36409 | 0.18186 | 1.2444 | 0.061783 | 0.89476 | 1.7446
Q9BY76 | Angiopoietin-related protein 4 | ANGPTL4 | 1:10 | Block_5 | -1.9544 | 5.8184 | FALSE | 12.4263 | 8.7841 | -0.66918 | 0.49742 | 1.4415 | -0.29721 | 0.83382 | 1.7605
P04083 | Annexin A1 | ANXA1 | 1:10 | Block_5 | 1.1522 | 6.4168 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.2874
P12429 | Annexin A3 | ANXA3 | 1:10 | Block_5 | -2.7436 | 3.5909 | TRUE | 8.5336 | 5.537 | -1.9035 | 0.43692 | 2.238 | 0.021279 | 2.2894 | 4.0278
P09525 | Annexin A4 | ANXA4 | 1:10 | Block_5 | 0.25085 | 7.3795 | TRUE | 41.5638 | 17.9294 | NA | 0.88777 | 2.4655 | 0.53896 | 2.9303 | 4.8499
P51693 | Amyloid beta precursor like protein 1 | APLP1 | 1:10 | Block_5 | -4.3018 | 4.9651 | TRUE | 10.3355 | 8.9268 | -0.071397 | 0.33927 | 1.059 | -0.36037 | 0.2666 | 1.121
Q0VD83 | Apolipoprotein B receptor | APOBR | 1:10 | Block_5 | -2.12 | 3.5134 | TRUE | 7.5771 | 9.2366 | -0.89835 | -0.055927 | 0.745 | -0.94514 | 0.13884 | 0.84968
Q8NEU8 | DCC-interacting protein 13-beta | APPL2 | 1:10 | Block_5 | -0.30605 | 9.5074 | FALSE | 15.0352 | 12.2476 | NA | 1.6291 | 5.1356 | 1.7442 | 4.9049 | 6.8394
Q92619 | Rho GTPase-activating protein 45 | ARHGAP45 | 1:10 | Block_5 | 1.07 | 6.9615 | FALSE | NA | NA | NA | NA | 2.6753 | NA | 2.6823 | 4.4074
Q9BXN1 | Asporin | ASPN | 1:10 | Block_5 | -0.040841 | 6.3263 | FALSE | 17.8699 | 9.6517 | 0.65573 | 1.2035 | 1.6156 | 0.48767 | 1.2302 | 2.0343
O00244 | Copper transport protein ATOX1 | ATOX1 | 1:10 | Block_5 | -1.6041 | 6.201 | TRUE | 15.6101 | 5.8053 | -0.42309 | 0.48681 | 2.7561 | 1.6291 | 3.0989 | 3.9878
Q9UII2 | ATPase inhibitor, mitochondrial | ATP5IF1 | 1:10 | Block_5 | -0.31699 | 7.6567 | TRUE | 25.6672 | 14.5726 | -0.10099 | 2.5116 | 7.0746 | 3.236 | 6.2724 | 7.6114
P15291 | Beta-1,4-galactosyltransferase 1 | B4GALT1 | 1:10 | Block_5 | -4.3956 | 4.6983 | FALSE | 11.2649 | 3.9863 | -0.3646 | 0.43709 | 0.90918 | -0.012019 | 0.79506 | 1.6722
P54687 | Branched-chain-amino-acid aminotransferase, cytosolic | BCAT1 | 1:10 | Block_5 | -0.36378 | 7.9149 | FALSE | 18.0776 | 16.4874 | -0.32799 | 0.65224 | 1.1035 | 0.75579 | 1.446 | 2.5424
Q8TD16 | Protein bicaudal D homolog 2 | BICD2 | 1:10 | Block_5 | 0.19389 | 8.5794 | FALSE | NA | NA | NA | 0.52999 | 3.7491 | 0.89445 | 3.5221 | 5.1297
O95393 | Bone morphogenetic protein 10 | BMP10 | 1:10 | Block_5 | -1.4673 | 5.8718 | TRUE | 17.2084 | 9.5715 | -0.96077 | -0.36019 | 0.20571 | -0.85617 | -0.19021 | 0.52376
Q8N8U9 | BMP-binding endothelial regulator protein | BMPER | 1:10 | Block_5 | -2.1 | 8.3562 | TRUE | 29.1111 | 7.9659 | -0.76567 | 0.044938 | 0.55359 | -0.31535 | 0.52002 | 1.2762
Q12982 | BCL2/adenovirus E1B 19 kDa protein-interacting protein 2 | BNIP2 | 1:10 | Block_5 | 0.60212 | 7.2074 | TRUE | NA | 3.0572 | NA | 0.62798 | 2.1843 | NA | 2.2792 | 3.498
P17213 | Bactericidal permeability-increasing protein | BPI | 1:10 | Block_5 | -0.44563 | 6.6559 | FALSE | 8.1921 | NA | NA | -0.23599 | 0.71794 | NA | 0.30647 | 3.2062
Q8N4F0 | BPI fold-containing family B member 2 | BPIFB2 | 1:10 | Block_5 | -1.4426 | 4.3881 | FALSE | 8.482 | 5.2551 | -1.1636 | -0.11664 | 0.92869 | -1.0844 | -0.0044205 | 1.9196
O60885 | Bromodomain-containing protein 4 | BRD4 | 1:10 | Block_5 | 0.30988 | 8.8794 | TRUE | 8.1059 | 3.3818 | NA | NA | 0.33537 | NA | NA | 0.73771
Q10588 | ADP-ribosyl cyclase/cyclic ADP-ribose hydrolase 2 | BST1 | 1:10 | Block_5 | -2.5956 | 3.2333 | TRUE | 8.0826 | 8.9679 | -0.40031 | 0.12456 | 0.49577 | -0.35268 | 0.20911 | 0.60011
Q10589 | Bone marrow stromal antigen 2 | BST2 | 1:10 | Block_5 | -3.76 | 8.7438 | TRUE | 9.912 | 4.4712 | -1.115 | -0.35352 | 0.33827 | -0.87561 | -0.095657 | 1.8683
Q7KYR7 | Butyrophilin subfamily 2 member A1 | BTN2A1 | 1:10 | Block_5 | -1.7408 | 4.9962 | TRUE | 12.6449 | 7.8489 | -0.32127 | 0.17718 | 0.70911 | 0.056534 | 0.71173 | 1.6581
Q9BXJ0 | Complement C1q tumor necrosis factor-related protein 5 | C1QTNF5 | 1:10 | Block_5 | -3.2672 | 3.6198 | FALSE | 10.2183 | 5.6354 | -0.34413 | 0.20714 | 0.73746 | -0.22686 | 0.4349 | 1.1202
P0C862 | Complement C1q and tumor necrosis factor-related protein 9A | C1QTNF9 | 1:10 | Block_5 | -2.7224 | 7.7014 | FALSE | 10.9165 | 9.4722 | -0.74401 | 0.1938 | 1.0941 | -0.48824 | 0.40057 | 1.2066
Q9BSF0 | Small membrane A-kinase anchor protein | C2orf88 | 1:10 | Block_5 | 0.24767 | NA | NA | NA | 0.63318 | NA | 1.1962 | 5.3682 | 1.8695 | 4.6225 | 6.2745
P20851 | C4b-binding protein beta chain | C4BPB | 1:10 | Block_5 | -2.1728 | 6.6024 | TRUE | 8.6063 | 5.0642 | -0.48237 | -0.064811 | 0.45195 | -0.34283 | 0.36246 | 0.9556
Q8N1Q1 | Carbonic anhydrase 13 | CA13 | 1:10 | Block_5 | 0.52597 | 7.417 | TRUE | 15.0798 | 3.1591 | NA | 1.7644 | 4.4542 | 1.3167 | 4.6389 | 6.1105
Q9HB71 | Calcyclin-binding protein | CACYBP | 1:10 | Block_5 | -0.7757 | 8.51 | FALSE | 16.7108 | 7.3183 | -0.76081 | 0.81731 | 4.6898 | 1.7846 | 5.2305 | 6.6553
Q8WVQ1 | Soluble calcium-activated nucleotidase 1 | CANT1 | 1:10 | Block_5 | -2.784 | 4.6396 | TRUE | 10.944 | 3.7864 | -0.51335 | 0.10547 | 0.40826 | -0.022847 | 0.45563 | 0.95773
P27352 | Cobalamin binding intrinsic factor | CBLIF | 1:10 | Block_5 | -4.2443 | 8.3428 | TRUE | 8.8711 | 5.8565 | -0.83057 | 0.049164 | 1.1861 | -1.1583 | 0.26909 | 1.7826
Q6UXH8 | Collagen and calcium-binding EGF domain-containing protein 1 | CCBE1 | 1:10 | Block_5 | -5 | 8.6638 | TRUE | 13.5105 | 10.6389 | -0.56336 | 0.15121 | 0.58472 | -0.045882 | 0.50454 | 1.7622
Q76M96 | Coiled-coil domain-containing protein 80 | CCDC80 | 1:10 | Block_5 | -4.1404 | 8.3151 | FALSE | 10.9213 | 4.7224 | -0.58773 | 0.89634 | 1.5246 | 0.54449 | 1.5934 | 2.7544
Q99616 | C-C motif chemokine 13 | CCL13 | 1:10 | Block_5 | -4.9 | 9.0158 | TRUE | 10.0874 | 11.5985 | -0.43957 | 0.8537 | 1.9805 | 0.86663 | 1.8389 | 2.9236
Q92583 | C-C motif chemokine 17 | CCL17 | 1:10 | Block_5 | -1.7383 | 8.114 | TRUE | 10.023 | 8.9488 | -0.58922 | 0.86964 | 2.7353 | 0.56222 | 2.5291 | 4.1497
Q99731 | C-C motif chemokine 19 | CCL19 | 1:10 | Block_5 | -5.3104 | 8.2297 | TRUE | 9.4542 | 8.1626 | -0.88739 | -0.15024 | 1.5793 | -0.41109 | 0.94086 | 2.3025
P13500 | C-C motif chemokine 2 | CCL2 | 1:10 | Block_5 | -4.5849 | 7.6081 | TRUE | 8.7376 | 7.3064 | -0.54918 | -0.0094818 | 0.77831 | -0.13624 | 0.59903 | 1.5578
O00585 | C-C motif chemokine 21 | CCL21 | 1:10 | Block_5 | -2.2379 | 7.8451 | TRUE | 11.5832 | 13.7884 | -1.2159 | -0.011061 | 0.62087 | -0.67032 | 0.3356 | 1.7602
O00626 | C-C motif chemokine 22 | CCL22 | 1:10 | Block_5 | -5.4 | 5.2764 | TRUE | 15.8961 | 11.7987 | -0.73202 | 0.056794 | 0.79194 | -0.98081 | 0.53286 | 1.5808
P55773 | C-C motif chemokine 23 | CCL23 | 1:10 | Block_5 | -6.3832 | 6.8776 | TRUE | 8.2754 | 8.2165 | -0.63133 | -0.032108 | 0.69682 | -0.14727 | 0.72478 | 1.9094
O00175 | C-C motif chemokine 24 | CCL24 | 1:10 | Block_5 | -6.4644 | 5.9664 | TRUE | 7.8998 | 8.7552 | -1.4684 | -0.66635 | 1.3084 | -1.1115 | 0.49481 | 1.9662
Q9Y4X3 | C-C motif chemokine 27 | CCL27 | 1:10 | Block_5 | -5.3539 | 5.8133 | FALSE | 11.4657 | 12.3823 | -0.57183 | 0.51288 | 1.142 | 0.18716 | 0.96421 | 1.9348
P13236 | C-C motif chemokine 4 | CCL4 | 1:10 | Block_5 | -1.7817 | 9.1681 | TRUE | 17.0503 | 9.9115 | -0.71366 | 0.38325 | 0.76835 | 0.13157 | 0.79751 | 1.7397
P29279 | CCN family member 2 | CCN2 | 1:10 | Block_5 | -2.38 | 4.7508 | TRUE | 8.8811 | 7.3106 | -0.17753 | 0.76543 | 1.8485 | 1.194 | 1.8878 | 2.7564
P29017 | T-cell surface glycoprotein CD1c | CD1C | 1:10 | Block_5 | -1.8478 | 4.2416 | FALSE | 5.5739 | 5.8516 | -0.63671 | 0.16056 | 0.46493 | -0.9018 | -0.14814 | 0.48921
P26842 | CD27 antigen | CD27 | 1:10 | Block_5 | -6.2417 | 5.3355 | TRUE | 12.5776 | 9.0019 | -0.57018 | 0.34088 | 0.89039 | -0.15437 | 0.93255 | 2.4378
Q5ZPR3 | CD276 antigen | CD276 | 1:10 | Block_5 | -3.3956 | 5.5261 | FALSE | 10.5288 | 7.9596 | -0.26946 | 0.4744 | 0.98865 | 0.22543 | 1.0477 | 2.1397
Q9Y5K6 | CD2-associated protein | CD2AP | 1:10 | Block_5 | -0.79854 | 10.3538 | TRUE | 17.0629 | 9.7495 | 0.10818 | 1.4364 | 3.2714 | 1.661 | 3.1491 | 4.6301
Q9UGN4 | CMRF35-like molecule 8 | CD300A | 1:10 | Block_5 | -3.9745 | 3.4996 | TRUE | 9.7186 | 7.4579 | -0.47609 | 0.036052 | 0.57911 | 0.011121 | 0.68412 | 1.7251
Q08708 | CMRF35-like molecule 6 | CD300C | 1:10 | Block_5 | -1.6531 | 7.3377 | TRUE | 12.9781 | 6.2992 | -0.4867 | 0.1964 | 0.84011 | 0.18526 | 1.0815 | 2.2088
P20138 | Myeloid cell surface antigen CD33 | CD33 | 1:10 | Block_5 | -2.3101 | 5.1538 | TRUE | 12.1981 | 6.8547 | -0.9695 | 0.4113 | 0.96555 | -1.0106 | 0.25715 | 1.2159
P28906 | Hematopoietic progenitor cell antigen CD34 | CD34 | 1:10 | Block_5 | -1.5089 | 6.6038 | FALSE | 6.6572 | 7.9424 | -0.36194 | 0.10966 | 0.56764 | -0.54219 | 0.063583 | 0.52852
P25942 | Tumor necrosis factor receptor superfamily member 5 | CD40 | 1:10 | Block_5 | -5.0512 | 7.8647 | TRUE | 12.2836 | 5.4878 | -0.33605 | 0.4788 | 1.6659 | 0.76688 | 1.5472 | 2.7488
P09326 | CD48 antigen | CD48 | 1:10 | Block_5 | -4.1981 | 4.3846 | FALSE | 8.6197 | 6.8806 | -0.52473 | 0.19655 | 0.71797 | -0.30995 | 0.4304 | 1.12
P19256 | Lymphocyte function-associated antigen 3 | CD58 | 1:10 | Block_5 | -4.8627 | 3.3541 | TRUE | 7.719 | 8.5079 | -0.23112 | 0.24536 | 0.57992 | -0.15778 | 0.30527 | 0.63599
Q07108 | Early activation antigen CD69 | CD69 | 1:10 | Block_5 | -3.39 | 10.192 | TRUE | 14.3639 | 7.1978 | -0.8128 | 0.75965 | 4.0798 | 1.0272 | 3.9005 | 5.6208
P21854 | B-cell differentiation antigen CD72 | CD72 | 1:10 | Block_5 | -0.53 | 8.8302 | TRUE | 26.9088 | 11.1904 | -0.35294 | 0.48726 | 1.3557 | -0.20758 | 0.8366 | 1.9871
P04233 | HLA class II histocompatibility antigen gamma chain | CD74 | 1:10 | Block_5 | -2.0802 | 3.161 | TRUE | 6.4075 | 7.1507 | -0.36471 | 0.24328 | 1.0177 | -0.12063 | 0.86363 | 1.7751
P33681 | T-lymphocyte activation antigen CD80 | CD80 | 1:10 | Block_5 | -6.1221 | 8.3201 | TRUE | 9.8631 | 7.4391 | -0.4395 | 0.30075 | 0.97373 | 0.11978 | 0.72252 | 1.7477
P01732 | T-cell surface glycoprotein CD8 alpha chain | CD8A | 1:10 | Block_5 | -3.3073 | NA | NA | 13.8353 | 6.0984 | -0.27038 | 0.52024 | 1.5257 | -0.55198 | 0.97467 | 2.1295
P32320 | Cytidine deaminase | CDA | 1:10 | Block_5 | -4.1412 | 5.811 | FALSE | 7.5883 | 6.3472 | -0.9586 | -0.075153 | 0.63236 | -0.6256 | 0.51015 | 2.2174
Q16543 | Hsp90 co-chaperone Cdc37 | CDC37 | 1:10 | Block_5 | 0.096134 | 7.5392 | TRUE | NA | 13.1567 | NA | 0.72871 | 4.0081 | 1.1574 | 4.2554 | 5.7142
P55291 | Cadherin-15 | CDH15 | 1:10 | Block_5 | -1.3391 | 8.224 | FALSE | 9.1787 | 7.4516 | -0.55638 | -0.074155 | 1.751 | -0.29661 | 0.57067 | 2.329
Q12864 | Cadherin-17 | CDH17 | 1:10 | Block_5 | -2.7293 | 4.5851 | TRUE | 12.1721 | 7.2528 | -1.4536 | 0.043427 | 1.2873 | -1.4059 | -0.25743 | 1.5672
P19022 | Cadherin-2 | CDH2 | 1:10 | Block_5 | -4.0879 | 4.0749 | FALSE | 8.2813 | 7.9062 | 0.18934 | 1.1683 | 1.6994 | 0.61487 | 1.6559 | 2.4472
P22223 | Cadherin-3 | CDH3 | 1:10 | Block_5 | -2.6786 | 6.0688 | FALSE | 9.2874 | 5.7676 | -0.69353 | 0.0091411 | 0.76549 | -0.25209 | 0.52337 | 1.866
P55285 | Cadherin-6 | CDH6 | 1:10 | Block_5 | -1.6673 | 4.9014 | TRUE | 10.0063 | 6.8057 | -0.67028 | -0.029055 | 0.36667 | -0.25168 | 0.47116 | 1.2381
Q9HBB8 | Cadherin-related family member 5 | CDHR5 | 1:10 | Block_5 | -4.887 | 5.4272 | TRUE | 6.3409 | 4.5413 | -0.57967 | -0.015409 | 0.67937 | -0.60855 | 0.11746 | 0.86361
Q4KMG0 | Cell adhesion molecule-related/down-regulated by oncogenes | CDON | 1:10 | Block_5 | -5.0528 | 5.1904 | TRUE | 9.6376 | 4.9859 | -0.25613 | 0.30287 | 0.84592 | -0.30466 | 0.36037 | 0.985
Q9UKY7 | Protein CDV3 homolog | CDV3 | 1:10 | Block_5 | -0.094887 | 8.9356 | TRUE | 8.2729 | 4.4681 | -0.046891 | 0.80013 | 3.5391 | 1.494 | 3.803 | 5.0944
P13688 | Carcinoembryonic antigen-related cell adhesion molecule 1 | CEACAM1 | 1:10 | Block_5 | -3.048 | 3.6849 | TRUE | 7.7976 | 3.7509 | -0.27505 | 0.17055 | 0.63498 | -0.18835 | 0.33898 | 0.98272
Q9HCU4 | Cadherin EGF LAG seven-pass G-type receptor 2 | CELSR2 | 1:10 | Block_5 | -3.33 | 6.3306 | TRUE | 16.1988 | 7.1069 | -0.13523 | 0.99125 | 1.6845 | 0.56316 | 1.4915 | 2.2177
Q9UPN4 | Centrosomal protein of 131 kDa | CEP131 | 1:10 | Block_5 | 0.39409 | 12.2442 | TRUE | NA | 3.7171 | NA | 0.65455 | 3.9164 | NA | 2.25 | 5.5667
Q99674 | Cell growth regulator with EF hand domain protein 1 | CGREF1 | 1:10 | Block_5 | -2.5936 | 4.0461 | TRUE | 10.7073 | 7.993 | -0.3924 | 0.37516 | 1.2602 | 0.12933 | 0.95423 | 2.3164
O15335 | Chondroadherin | CHAD | 1:10 | Block_5 | -3.3356 | 3.4915 | TRUE | 10.1017 | 7.1333 | -0.93853 | -0.095649 | 0.54388 | -1.6952 | -0.079753 | 1.0168
P05060 | Secretogranin-1 | CHGB | 1:10 | Block_5 | -3.0734 | 7.2429 | FALSE | 9.16 | 5.9539 | -0.12198 | 0.94783 | 1.4632 | 0.32791 | 1.3951 | 2.7598
Q96FZ7 | Charged multivesicular body protein 6 | CHMP6 | 1:10 | Block_5 | -1.4504 | 7.322 | TRUE | 16.2563 | 6.8295 | -0.50562 | 0.26337 | 1.6921 | 0.57328 | 1.7485 | 2.7588
Q9BU40 | Chordin-like protein 1 | CHRDL1 | 1:10 | Block_5 | -4.1491 | 3.8293 | FALSE | 11.974 | 3.7503 | 0.47942 | 1.4135 | 1.9725 | 0.804 | 1.7337 | 2.4163
Q6WN34 | Chordin-like protein 2 | CHRDL2 | 1:10 | Block_5 | -2.9829 | 5.2995 | TRUE | 8.9178 | 5.4863 | -1.039 | 0.24885 | 1.3955 | -0.43178 | 0.68167 | 1.7056
Q6FI81 | Anamorsin | CIAPIN1 | 1:10 | Block_5 | -1.3217 | 6.8563 | FALSE | 14.0825 | 16.5746 | 0.017761 | 1.3187 | 3.4678 | 2.0132 | 3.847 | 5.0808
Q14008 | Cytoskeleton-associated protein 5 | CKAP5 | 1:10 | Block_5 | -0.3727 | NA | NA | 3.1462 | 8.2871 | NA | 0.21379 | 3.2124 | 0.59774 | 3.1442 | 4.8576
P12532 | Creatine kinase U-type, mitochondrial | CKMT1A_CKMT1B | 1:10 | Block_5 | -6.12 | 6.1758 | TRUE | 17.8438 | 8.3678 | -0.725 | 0.55469 | 1.7893 | -0.23729 | 1.1153 | 2.2248
Q05315 | Galectin-10 | CLC | 1:10 | Block_5 | -1.186 | 5.0146 | TRUE | 5.526 | 4.248 | -0.74073 | 0.18014 | 0.97554 | -0.28174 | 0.73247 | 3.6834
Q9P126 | C-type lectin domain family 1 member B | CLEC1B | 1:10 | Block_5 | -4 | 9.7144 | TRUE | 13.2737 | 6.6228 | -0.37951 | 1.1377 | 3.6832 | 1.6389 | 3.537 | 5.2334
Q9NY25 | C-type lectin domain family 5 member A | CLEC5A | 1:10 | Block_5 | -3.477 | 9.9789 | TRUE | 7.0918 | 12.216 | -0.63637 | 0.27646 | 0.65323 | -0.039491 | 0.57135 | 1.9339
P09497 | Clathrin light chain B | CLTB | 1:10 | Block_5 | 0.46418 | 7.3436 | TRUE | NA | NA | NA | 0.4738 | 3.4259 | 0.8034 | 2.6315 | 4.6076
Q15846 | Clusterin-like protein 1 | CLUL1 | 1:10 | Block_5 | -3.3499 | 5.8119 | TRUE | 7.7684 | 3.5508 | -0.22185 | 0.45661 | 1.5711 | -0.33217 | 0.5623 | 1.4332
Q7Z7K0 | COX assembly mitochondrial protein homolog | CMC1 | 1:10 | Block_5 | 2.16 | 9.8511 | FALSE | NA | NA | NA | NA | 3.4535 | NA | 3.2593 | 5.3497
Q99439 | Calponin-2 | CNN2 | 1:10 | Block_5 | 0.65866 | NA | NA | 14.5793 | 3.61 | NA | 1.7497 | 5.4594 | 1.2693 | 5.1806 | 6.41
Q6PJW8 | Consortin | CNST | 1:10 | Block_5 | 0.43692 | 7.6843 | TRUE | NA | NA | NA | NA | 1.736 | NA | 0.92841 | 2.5142
O94779 | Contactin-5 | CNTN5 | 1:10 | Block_5 | -2.0956 | 4.7517 | TRUE | 10.6395 | 7.031 | -0.3825 | 0.20009 | 1.0062 | -0.6584 | 0.18724 | 0.9362
O43405 | Cochlin | COCH | 1:10 | Block_5 | -3.5718 | 6.7155 | TRUE | 9.5828 | 6.1031 | -0.78552 | 0.092937 | 0.60695 | 0.36786 | 1.0644 | 1.7668
P02462 | Collagen alpha-1(IV) chain | COL4A1 | 1:10 | Block_5 | -3.8676 | 7.8935 | TRUE | 9.3509 | 4.6814 | -0.43533 | 0.29109 | 0.92666 | 0.059949 | 0.76379 | 2.2323
P20908 | Collagen alpha-1(V) chain | COL5A1 | 1:10 | Block_5 | -1.4093 | 10.3046 | TRUE | 16.977 | 13.5889 | -0.55531 | 0.19383 | 1.1488 | 0.17405 | 1.0437 | 2.0853
Q5KU26 | Collectin-12 | COLEC12 | 1:10 | Block_5 | -2.9249 | 5.7116 | TRUE | 9.4992 | 5.0119 | -0.30901 | 0.17873 | 0.54476 | -0.026657 | 0.87464 | 1.8433
Q86X83 | COMM domain-containing protein 2 | COMMD2 | 1:10 | Block_5 | -0.66679 | NA | NA | 10.4725 | 7.6322 | -0.34699 | 0.24603 | 1.5096 | 0.73191 | 1.8717 | 3.3729
P14854 | Cytochrome c oxidase subunit 6B1 | COX6B1 | 1:10 | Block_5 | 1.4332 | NA | NA | NA | NA | NA | NA | 3.2919 | NA | 1.973 | 3.2838
P48052 | Carboxypeptidase A2 | CPA2 | 1:10 | Block_5 | -3.5045 | NA | NA | 9.3899 | 6.1636 | -0.87452 | 0.55902 | 1.6393 | -0.62 | 0.51439 | 1.7475
Q9UI42 | Carboxypeptidase A4 | CPA4 | 1:10 | Block_5 | -3.75 | 4.2532 | TRUE | 27.8901 | 23.8954 | -0.46538 | -0.13131 | 0.34066 | -0.61955 | 0.29894 | 1.1125
P14384 | Carboxypeptidase M | CPM | 1:10 | Block_5 | -4.16 | 4.2459 | TRUE | 10.1208 | 8.3734 | -0.46527 | 0.34817 | 0.7511 | -0.095231 | 0.54297 | 1.3413
Q9BRF8 | Serine/threonine-protein phosphatase CPPED1 | CPPED1 | 1:10 | Block_5 | -2.2888 | 6.9561 | TRUE | 12.8246 | 3.3867 | -0.67422 | 1.1046 | 1.915 | 0.45888 | 2.0669 | 4.3178
Q96SM3 | Probable carboxypeptidase X1 | CPXM1 | 1:10 | Block_5 | -0.27845 | 6.8796 | TRUE | 7.2834 | 2.8717 | -0.027497 | 0.45608 | 0.94643 | 0.40982 | 1.3997 | 2.4279
Q8N436 | Inactive carboxypeptidase-like protein X2 | CPXM2 | 1:10 | Block_5 | -0.87604 | 7.418 | FALSE | 13.1076 | 10.7432 | -0.082522 | 0.44346 | 1.0388 | 0.12886 | 1.0542 | 2.2195
Q5IJ48 | Protein crumbs homolog 2 | CRB2 | 1:10 | Block_5 | -2.7635 | 7.4093 | TRUE | 8.2673 | 10.8202 | -0.11223 | 0.28539 | 0.85164 | -0.27694 | 0.37706 | 1.0607
O75629 | Protein CREG1 | CREG1 | 1:10 | Block_5 | -3.3 | 3.7775 | TRUE | 46.7593 | 33.1218 | -0.71447 | 0.16013 | 0.55031 | -0.14774 | 0.41966 | 1.2803
Q96HD1 | Protein disulfide isomerase CRELD1 | CRELD1 | 1:10 | Block_5 | -6.24 | 4.5133 | TRUE | 12.6976 | 10.3728 | -0.62355 | 0.24882 | 0.81032 | -0.0014292 | 0.60605 | 1.2813
Q6UXH1 | Protein disulfide isomerase CRELD2 | CRELD2 | 1:10 | Block_5 | -5.57 | 6.299 | TRUE | 11.7722 | 7.6761 | -0.4609 | 0.65712 | 1.3315 | 0.77571 | 1.6789 | 2.4589
P24387 | Corticotropin-releasing factor-binding protein | CRHBP | 1:10 | Block_5 | -5 | 2.4691 | FALSE | 14.0138 | 4.4767 | -0.65941 | 0.086859 | 0.52997 | -0.25325 | 0.465 | 1.1671
Q9NZV1 | Cysteine-rich motor neuron 1 protein | CRIM1 | 1:10 | Block_5 | -2.8938 | 6.902 | TRUE | 10.2044 | 7.3156 | -0.30756 | 0.389 | 0.93542 | 0.16668 | 0.62244 | 1.4753
P46109 | Crk-like protein | CRKL | 1:10 | Block_5 | -1.0452 | 7.5844 | TRUE | 12.2057 | 10.8291 | -0.55022 | 1.3996 | 5.0321 | 2.181 | 5.17 | 6.244
O75462 | Cytokine receptor-like factor 1 | CRLF1 | 1:10 | Block_5 | -2.3292 | 4.9019 | TRUE | 10.8077 | 5.5623 | -0.65902 | 0.024095 | 0.57688 | -0.047172 | 0.61792 | 1.3531
P53674 | Beta-crystallin B1 | CRYBB1 | 1:10 | Block_5 | -3.4624 | 10.145 | TRUE | 14.8821 | 6.6066 | -0.54077 | 0.70244 | 2.4015 | 1.451 | 2.9291 | 4.0467
P09603 | Macrophage colony-stimulating factor 1 | CSF1 | 1:10 | Block_5 | -4.8866 | 8.452 | TRUE | 8.5227 | 4.6795 | -0.79834 | -0.20762 | 0.30175 | -0.239 | 0.55288 | 1.5008
P01037 | Cystatin-SN | CST1 | 1:10 | Block_5 | -2.94 | 8.1739 | FALSE | 14.6644 | 12.5295 | -1.4045 | -0.44803 | 1.0876 | -0.69271 | 0.56185 | 3.1354
O76096 | Cystatin-F | CST7 | 1:10 | Block_5 | -3.2681 | 7.9639 | TRUE | 11.9861 | 10.6519 | -0.4136 | 0.62314 | 1.3521 | 0.23083 | 1.5453 | 3.4849
Q99895 | Chymotrypsin-C | CTRC | 1:10 | Block_5 | -5.5 | 5.5614 | TRUE | 14.4344 | 3.8157 | -0.87289 | 0.61701 | 1.8164 | -1.2565 | 0.10094 | 1.5971
P53634 | Dipeptidyl peptidase 1 | CTSC | 1:10 | Block_5 | -1.2834 | 6.3242 | TRUE | 12.7085 | 8.0009 | 0.18603 | 1.5429 | 2.5893 | 0.56701 | 1.6665 | 2.7274
P14091 | Cathepsin E | CTSE | 1:10 | Block_5 | -3.8787 | NA | NA | 10.619 | 8.9726 | -0.57373 | 0.80541 | 1.843 | -0.39927 | 0.90416 | 2.073
Q9UBX1 | Cathepsin F | CTSF | 1:10 | Block_5 | -2.1993 | 4.3956 | TRUE | 24.4149 | 11.9329 | -1.1014 | -0.12505 | 0.51585 | -1.1962 | -0.2714 | 0.53738
P07711 | Procathepsin L | CTSL | 1:10 | Block_5 | -1.8936 | 6.5458 | FALSE | 11.2475 | 7.6166 | -0.21622 | 0.50235 | 0.81528 | 0.3016 | 1.1053 | 2.112
P25774 | Cathepsin S | CTSS | 1:10 | Block_5 | -4.8359 | 4.2316 | TRUE | 5.1547 | 6.4905 | -0.0089619 | 0.3443 | 0.5657 | 0.17839 | 0.57081 | 1.0769
P09341 | Growth-regulated alpha protein | CXCL1 | 1:10 | Block_5 | -1.52 | 10.4446 | TRUE | 34.0951 | 16.7835 | -0.74295 | 0.69882 | 2.4598 | 0.61527 | 2.115 | 4.1189
P02778 | C-X-C motif chemokine 10 | CXCL10 | 1:10 | Block_5 | -2.5367 | 7.6231 | TRUE | 25.8974 | 7.8303 | -0.93708 | 0.14893 | 1.7388 | -0.11199 | 1.0382 | 2.6608
O14625 | C-X-C motif chemokine 11 | CXCL11 | 1:10 | Block_5 | -1.1314 | 9.7013 | TRUE | 10.8582 | 10.0479 | 0.059011 | 1.4909 | 2.7126 | 1.0996 | 2.5549 | 3.7831
O43927 | C-X-C motif chemokine 13 | CXCL13 | 1:10 | Block_5 | -4.3494 | 7.4307 | TRUE | 8.73 | 10.8725 | -1.5604 | -0.4602 | 0.46538 | -0.9822 | 0.085041 | 1.9506
P42830 | C-X-C motif chemokine 5 | CXCL5 | 1:10 | Block_5 | -2.7509 | 7.4308 | TRUE | 10.5243 | 9.5314 | -1.3631 | 1.0466 | 3.125 | -0.2833 | 2.4825 | 4.9222
P80162 | C-X-C motif chemokine 6 | CXCL6 | 1:10 | Block_5 | -2.98 | 10.2266 | TRUE | 17.4751 | 15.5569 | -1.3491 | 1.4776 | 2.3356 | 1.3826 | 2.9044 | 4.5235
Q07325 | C-X-C motif chemokine 9 | CXCL9 | 1:10 | Block_5 | -1.7058 | 8.7946 | TRUE | 7.7093 | 7.9785 | -0.39384 | 0.49765 | 2.2448 | 0.13582 | 1.5467 | 3.8516
Q9Y4D1 | Disheveled-associated activator of morphogenesis 1 | DAAM1 | 1:10 | Block_5 | 2 | 7.5535 | FALSE | NA | NA | NA | NA | 3.683 | NA | 3.105 | 5.0002
Q14118 | Dystroglycan 1 | DAG1 | 1:10 | Block_5 | -2.7041 | 6.5575 | TRUE | 11.1383 | 4.9846 | -0.37716 | 0.68392 | 1.8432 | 1.0457 | 2.1204 | 3.2176
P51397 | Death-associated protein 1 | DAP | 1:10 | Block_5 | -0.83357 | 9.2652 | TRUE | 9.509 | 12.0376 | -0.54599 | 1.6971 | 6.0154 | 1.3398 | 6.0138 | 7.4636
Q9UJU6 | Drebrin-like protein | DBNL | 1:10 | Block_5 | -0.96021 | 9.2685 | FALSE | 19.0175 | 14.2082 | -0.75342 | 1.0959 | 5.3624 | 1.8506 | 5.8916 | 6.9638
Q96PD2 | Discoidin, CUB and LCCL domain-containing protein 2 | DCBLD2 | 1:10 | Block_5 | -3.0585 | 8.2168 | TRUE | 11.2965 | 5.085 | -0.24146 | 0.53873 | 0.93586 | 0.25962 | 1.0821 | 2.043
P32321 | Deoxycytidylate deaminase | DCTD | 1:10 | Block_5 | 2 | 8.7206 | TRUE | NA | NA | NA | NA | 4.1964 | NA | 4.5389 | 6.3542
Q9H773 | dCTP pyrophosphatase 1 | DCTPP1 | 1:10 | Block_5 | -3.8922 | 7.3662 | TRUE | 8.2368 | 10.1157 | -0.45996 | 0.64874 | 1.2333 | 0.39747 | 1.2247 | 2.5262
Q7Z4W1 | L-xylulose reductase | DCXR | 1:10 | Block_5 | -3 | 3.9718 | TRUE | 14.8002 | 7.6623 | -1.3229 | -0.3732 | 0.59008 | -0.86583 | 0.18489 | 1.5834
P20711 | Aromatic-L-amino-acid decarboxylase | DDC | 1:10 | Block_5 | -5.4366 | 6.6679 | TRUE | 8.4806 | 9.7491 | -0.50614 | 0.20544 | 1.1058 | -0.44216 | 0.65318 | 1.7211
Q5TDH0 | Protein DDI1 homolog 2 | DDI2 | 1:10 | Block_5 | -1.6 | 7.6392 | TRUE | 22.8133 | 8.6135 | 1.0975 | 2.6464 | 3.6294 | 1.0889 | 2.5682 | 3.8546
Q08345 | Epithelial discoidin domain-containing receptor 1 | DDR1 | 1:10 | Block_5 | -4.7668 | 6.1242 | TRUE | 7.9323 | 8.7841 | -0.33913 | 0.18845 | 0.83815 | -0.12377 | 0.34402 | 1.0823
Q9NTK1 | Protein DEPP1 | DEPP1 | 1:10 | Block_5 | 0.56561 | NA | NA | NA | NA | NA | NA | 0.67917 | NA | NA | NA
Q9NR28 | Diablo IAP-binding mitochondrial protein | DIABLO | 1:10 | Block_5 | 1.1114 | 8.7366 | TRUE | NA | 15.11 | NA | 1.6392 | 4.4187 | 1.3592 | 3.9652 | 5.9375
Q9H7Y0 | Divergent protein kinase domain 2B | DIPK2B | 1:10 | Block_5 | -1.8452 | 7.0979 | FALSE | 12.99 | 5.5925 | -0.51843 | 0.028739 | 0.52009 | -0.38198 | 0.36635 | 0.95017
O94907 | Dickkopf-related protein 1 | DKK1 | 1:10 | Block_5 | -3.9504 | 7.9848 | TRUE | 11.7348 | 11.6949 | -0.40766 | 1.1795 | 2.3439 | 1.0508 | 2.4803 | 3.4364
O00548 | Delta-like protein 1 | DLL1 | 1:10 | Block_5 | -4.3321 | 5.8353 | TRUE | 10.2557 | 4.359 | -0.50229 | 0.20476 | 0.62302 | -0.065849 | 0.7741 | 1.8072
P25686 | DnaJ homolog subfamily B member 2 | DNAJB2 | 1:10 | Block_5 | 0.17212 | 7.8104 | TRUE | NA | 10.4069 | NA | 1.8738 | 3.3304 | 1.6788 | 3.3454 | 4.6628
Q99615 | DnaJ homolog subfamily C member 7 | DNAJC7 | 1:10 | Block_5 | -0.3101 | NA | NA | 4.6544 | 7.8692 | NA | 0.39582 | 2.6091 | 0.73027 | 3.0381 | 4.2919
Q8NFT8 | Delta and Notch-like epidermal growth factor-related receptor | DNER | 1:10 | Block_5 | -4.2732 | 3.7176 | TRUE | 7.0288 | 7.9545 | -0.65987 | -0.0099386 | 0.65235 | -0.72091 | -0.12506 | 0.45113
O43598 | 2'-deoxynucleoside 5'-phosphate N-hydrolase 1 | DNPH1 | 1:10 | Block_5 | -2.4886 | 5.6401 | TRUE | 13.0461 | 7.4429 | 0.26678 | 1.9905 | 3.2598 | 1.5278 | 3.1895 | 4.082
O60496 | Docking protein 2 | DOK2 | 1:10 | Block_5 | -0.12937 | 10.266 | TRUE | 10.0379 | 10.2717 | NA | 0.88585 | 5.4505 | 1.5271 | 5.1528 | 7.2435
Q9UHL4 | Dipeptidyl peptidase 2 | DPP7 | 1:10 | Block_5 | -2.7096 | 4.3935 | TRUE | 15.0492 | 12.1973 | -1.1061 | -0.19085 | 1.0368 | -0.56908 | 0.62809 | 1.9387
Q8TEA8 | D-aminoacyl-tRNA deacylase 1 | DTD1 | 1:10 | Block_5 | 0.39137 | 8.5398 | FALSE | NA | 51.4296 | NA | 1.0285 | 5.3173 | 1.6176 | 4.9118 | 6.5635
P23919 | Thymidylate kinase | DTYMK | 1:10 | Block_5 | 0.41951 | 7.0387 | FALSE | NA | NA | 0.61676 | 2.0109 | 4.0786 | 1.9464 | 4.3504 | 5.5478
Q96DC8 | Enoyl-CoA hydratase domain-containing protein 3, mitochondrial | ECHDC3 | 1:10 | Block_5 | -4.6228 | 7.6004 | TRUE | 8.2853 | 5.9471 | -3.3076 | -2.0694 | -0.44078 | -2.7908 | -1.2702 | 0.041029
P29692 | Elongation factor 1-delta | EEF1D | 1:10 | Block_5 | -1.213 | 8.1165 | FALSE | 8.7714 | 5.7369 | -1.099 | 1.5472 | 2.9535 | 1.3695 | 3.0353 | 5.2278
O95967 | EGF-containing fibulin-like extracellular matrix protein 2 | EFEMP2 | 1:10 | Block_5 | -1.2555 | 9.0341 | TRUE | 14.3793 | 7.054 | -0.45256 | 0.81626 | 1.6647 | 0.015865 | 1.3305 | 2.5571
P01133 | Pro-epidermal growth factor | EGF | 1:10 | Block_5 | 0.47005 | 10.2879 | TRUE | 11.6869 | 16.3476 | NA | 2.339 | 4.1491 | NA | 4.4252 | 6.1536
Q8IUX8 | Epidermal growth factor-like protein 6 | EGFL6 | 1:10 | Block_5 | -1.3522 | 11.3826 | TRUE | 13.9238 | 9.6283 | -1.0457 | 0.42997 | 2.7213 | 0.39821 | 2.4542 | 3.9458
Q8NDI1 | EH domain-binding protein 1 | EHBP1 | 1:10 | Block_5 | 0.24155 | 11.2184 | TRUE | NA | 7.7308 | NA | 0.57825 | 2.6722 | 0.71913 | 2.4389 | 3.9965
Q13541 | Eukaryotic translation initiation factor 4E-binding protein 1 | EIF4EBP1 | 1:10 | Block_5 | 0.55819 | 10.0254 | TRUE | 5.7623 | 4.2932 | 0.84048 | 2.0741 | 3.6835 | 1.5574 | 4.0232 | 5.8903
P15502 | Elastin | ELN | 1:10 | Block_5 | 0.67376 | 2.989 | FALSE | NA | NA | NA | NA | 0.87828 | NA | NA | 1.1173
Q15006 | ER membrane protein complex subunit 2 | EMC2 | 1:10 | Block_5 | 0.66415 | 9.3778 | FALSE | 6.6466 | 12.9975 | NA | NA | 0.88119 | NA | NA | NA
P06733 | Alpha-enolase | ENO1 | 1:10 | Block_5 | -0.8048 | 4.7092 | FALSE | 17.0283 | 14.5904 | 0.28357 | 2.1367 | 3.3022 | 1.4184 | 3.7753 | 4.7055
P09104 | Gamma-enolase | ENO2 | 1:10 | Block_5 | -2.8837 | 5.1032 | FALSE | 12.8052 | 6.9125 | 0.31798 | 1.5614 | 4.0011 | 1.9254 | 4.0217 | 5.1761
P22413 | Ectonucleotide pyrophosphatase/phosphodiesterase family member 1 | ENPP1 | 1:10 | Block_5 | -2.64 | 7.173 | FALSE | 10.487 | 3.9319 | -0.1343 | 0.43398 | 1.1756 | 0.16142 | 0.98436 | 1.5204
Q9UJA9 | Ectonucleotide pyrophosphatase/phosphodiesterase family member 5 | ENPP5 | 1:10 | Block_5 | -4.4543 | 3.7605 | TRUE | 9.1992 | 4.4278 | -0.90897 | -0.079762 | 0.54343 | -1.0565 | -0.13304 | 0.48496
Q6UWV6 | Ectonucleotide pyrophosphatase/phosphodiesterase family member 7 | ENPP7 | 1:10 | Block_5 | -4.7088 | 5.6059 | TRUE | 12.7783 | 6.6604 | -1.3346 | -0.35752 | 0.36352 | -1.6283 | -0.28022 | 1.201
P16422 | Epithelial cell adhesion molecule | EPCAM | 1:10 | Block_5 | -4.1164 | 6.0437 | TRUE | 7.5182 | 7.7155 | -1.7761 | -0.22459 | 1.0233 | -1.745 | -0.26602 | 1.0481
Q9UM22 | Mammalian ependymin-related protein 1 | EPDR1 | 1:10 | Block_5 | -1.4529 | 7.5747 | TRUE | 11.6939 | 13.4779 | -0.17077 | 0.93902 | 2.2843 | 1.1235 | 2.4903 | 3.8141
P21709 | Ephrin type-A receptor 1 | EPHA1 | 1:10 | Block_5 | -4.7232 | 5.7364 | TRUE | 7.579 | 2.4325 | -0.74673 | 0.054701 | 0.49264 | -0.04048 | 0.67054 | 1.3687
P54764 | Ephrin type-A receptor 4 | EPHA4 | 1:10 | Block_5 | -6.11 | 3.5636 | TRUE | 12.2722 | 14.9848 | -0.55757 | 0.13057 | 0.8083 | -0.27678 | 0.65688 | 1.4243
P04626 | Receptor tyrosine-protein kinase erbB-2 | ERBB2 | 1:10 | Block_5 | -3.4797 | 4.7529 | TRUE | 47.7261 | 26.1115 | -0.3119 | 0.11495 | 0.76343 | -0.18135 | 0.30226 | 0.83942
P21860 | Receptor tyrosine-protein kinase erbB-3 | ERBB3 | 1:10 | Block_5 | -5.5826 | 4.1186 | TRUE | 9.0334 | 7.0956 | -0.60271 | 0.020461 | 0.46532 | -0.38563 | 0.10194 | 0.65132
Q15303 | Receptor tyrosine-protein kinase erbB-4 | ERBB4 | 1:10 | Block_5 | -3.3561 | 6.0238 | TRUE | 11.9629 | 11.4832 | -0.38698 | 0.11302 | 0.63158 | -0.28073 | 0.2136 | 0.77237
P30040 | Endoplasmic reticulum resident protein 29 | ERP29 | 1:10 | Block_5 | 0.17246 | 6.779 | TRUE | 17.0953 | 2.9888 | NA | 1.3434 | 4.3372 | 1.435 | 4.7222 | 6.2629
Q9NQ30 | Endothelial cell-specific molecule 1 | ESM1 | 1:10 | Block_5 | -2.9408 | 6.8397 | TRUE | 10.6635 | 8.635 | 0.024803 | 0.93138 | 1.6831 | 0.77756 | 1.6376 | 2.3536
Q9BSJ8 | Extended synaptotagmin-1 | ESYT1 | 1:10 | Block_5 | -0.35433 | NA | NA | 10.0453 | 6.6633 | NA | 0.73748 | 3.5708 | 0.84825 | 2.5549 | 5.5348
P15311 | Ezrin | EZR | 1:10 | Block_5 | -1.4605 | 5.4438 | TRUE | 10.8564 | 3.4753 | -0.3847 | 0.28067 | 0.9806 | 0.19511 | 0.80174 | 2.0809
P25116 | Proteinase-activated receptor 1 | F2R | 1:10 | Block_5 | -2.6862 | NA | NA | 10.3574 | 6.5729 | 0.64641 | 1.3778 | 2.1658 | 1.6561 | 3.0307 | 4.0084
P07148 | Fatty acid-binding protein, liver | FABP1 | 1:10 | Block_5 | -6.4116 | 5.4058 | FALSE | 14.5476 | 16.2222 | -1.8305 | -0.81151 | 1.5853 | -0.95375 | 0.96404 | 3.491
P12104 | Fatty acid-binding protein, intestinal | FABP2 | 1:10 | Block_5 | -5.1059 | 5.5078 | TRUE | 9.854 | 7.3608 | -0.82156 | -0.22417 | 1.3463 | -0.86981 | 0.53097 | 1.9802
P05413 | Fatty acid-binding protein, heart | FABP3 | 1:10 | Block_5 | -8.55 | 4.8435 | TRUE | 12.7885 | 6.0244 | -0.48114 | 0.24464 | 0.95425 | -0.25755 | 1.1221 | 2.5812
Q01469 | Fatty acid-binding protein 5 | FABP5 | 1:10 | Block_5 | -1.08 | 6.4743 | FALSE | 12.8349 | 17.8119 | -0.46718 | 0.33941 | 1.3226 | 0.72033 | 2.0303 | 3.3027
Q13158 | FAS-associated death domain protein | FADD | 1:10 | Block_5 | 0.30584 | 4.9815 | FALSE | 3.6915 | 0.50853 | NA | 0.99268 | 3.1745 | 0.80755 | 3.5719 | 5.0932
Q96MK3 | Pseudokinase FAM20A | FAM20A | 1:10 | Block_5 | -2.0414 | 6.9518 | TRUE | 10.0666 | 9.4027 | -0.61563 | 0.10026 | 0.62211 | -0.17899 | 0.37553 | 1.0853
Q92520 | Protein FAM3C | FAM3C | 1:10 | Block_5 | -5.523 | 4.1693 | TRUE | 10.7644 | 11.7397 | -0.42157 | 0.10269 | 0.69953 | 0.14589 | 0.87604 | 1.7058
Q53RD9 | Fibulin-7 | FBLN7 | 1:10 | Block_5 | -3.6987 | NA | NA | 22.9355 | 10.4575 | -0.88198 | -0.021104 | 0.66471 | -0.40556 | 0.24418 | 1.0078
P09467 | Fructose-1,6-bisphosphatase 1 | FBP1 | 1:10 | Block_5 | -4.1692 | 6.509 | FALSE | 17.5235 | 15.5688 | -2.5428 | -1.4514 | 0.48506 | -2.7023 | -0.79948 | 1.4369
Q8WWV6 | High affinity immunoglobulin alpha and immunoglobulin mu Fc receptor | FCAMR | 1:10 | Block_5 | -3.6117 | 5.9983 | FALSE | 8.47 | 7.3782 | -1.0132 | 0.32512 | 1.1883 | -0.4029 | 0.63054 | 2.053
Q96LA6 | Fc receptor-like protein 1 | FCRL1 | 1:10 | Block_5 | -5.2064 | 5.9783 | FALSE | 8.4103 | 6.5388 | -0.88948 | -0.029376 | 0.6037 | -1.2679 | 0.0019897 | 0.96858
O95750 | Fibroblast growth factor 19 | FGF19 | 1:10 | Block_5 | -3.0632 | 6.5902 | TRUE | 10.2162 | 6.262 | -1.2255 | -0.087249 | 1.307 | -1.4704 | -0.089975 | 1.25
Q14512 | Fibroblast growth factor-binding protein 1 | FGFBP1 | 1:10 | Block_5 | -5.21 | 7.9909 | TRUE | 15.2935 | 14.7633 | -0.2208 | 0.40333 | 1.3485 | -0.27934 | 0.45819 | 1.2554
P21802 | Fibroblast growth factor receptor 2 | FGFR2 | 1:10 | Block_5 | -0.059873 | 7.7411 | TRUE | 10.0978 | 13.2167 | NA | 0.24731 | 0.58076 | NA | 0.42246 | 0.94691
P22455 | Fibroblast growth factor receptor 4 | FGFR4 | 1:10 | Block_5 | -4.3778 | 4.6496 | TRUE | 11.7658 | 6.6884 | -1.0685 | -0.0081084 | 0.95437 | -0.75655 | 0.084376 | 1.1836
Q9Y3D6 | Mitochondrial fission 1 protein | FIS1 | 1:10 | Block_5 | -3.8056 | 5.4544 | FALSE | 8.9367 | 10.5857 | -0.47569 | 1.1601 | 2.5516 | 1.1701 | 2.7686 | 4.072
P35916 | Vascular endothelial growth factor receptor 3 | FLT4 | 1:10 | Block_5 | -0.97329 | 1.9305 | FALSE | 10.4363 | 10.2947 | -0.44521 | 0.14039 | 0.83673 | -0.30109 | 0.37952 | 1.0155
Q96RU3 | Formin-binding protein 1 | FNBP1 | 1:10 | Block_5 | -0.7194 | 7.895 | TRUE | 6.4722 | 3.6137 | -0.69348 | 0.10034 | 0.96698 | -0.11505 | 1.2293 | 2.9002
P15328 | Folate receptor alpha | FOLR1 | 1:10 | Block_5 | -5.77 | 6.0305 | TRUE | 9.2027 | 8.6279 | -0.32626 | 0.38986 | 0.78913 | 0.16239 | 0.80831 | 1.6625
P14207 | Folate receptor beta | FOLR2 | 1:10 | Block_5 | -4.5445 | 3.5346 | TRUE | 12.6147 | 6.5551 | -0.20614 | 0.2139 | 0.58833 | -0.16786 | 0.54607 | 1.1944
P41439 | Folate receptor gamma | FOLR3 | 1:10 | Block_5 | -8.7765 | 3.7263 | TRUE | 8.7368 | 6.5878 | -6.6805 | -6.145 | 2.2073 | -6.2276 | -5.4204 | 2.2562
P19883 | Follistatin | FST | 1:10 | Block_5 | -4.7258 | 6.7459 | TRUE | 7.9649 | 6.2476 | -0.32894 | 0.6362 | 1.346 | -0.22157 | 0.58745 | 1.4906
O95633 | Follistatin-related protein 3 | FSTL3 | 1:10 | Block_5 | -3.8049 | 4.2264 | TRUE | 9.9541 | 6.3088 | -0.54487 | 0.16495 | 0.64407 | -0.0051739 | 0.91634 | 2.1497
Q9C0B1 | Alpha-ketoglutarate-dependent dioxygenase FTO | FTO | 1:10 | Block_5 | -2.188 | NA | NA | 10.8509 | 4.1922 | -0.71354 | 0.1872 | 0.82336 | 0.36496 | 1.0479 | 2.2203
A2VDF0 | Fucose mutarotase | FUOM | 1:10 | Block_5 | -2.4105 | 5.4444 | TRUE | 12.4804 | 11.7108 | -1.1822 | -0.43225 | 0.51962 | -1.2295 | -0.06157 | 1.3617
P09958 | Furin | FURIN | 1:10 | Block_5 | -2.2875 | 6.0957 | TRUE | 11.1613 | 9.0404 | -0.57486 | 0.023824 | 0.87445 | -0.19945 | 0.40813 | 1.091
P21217_Q11128 | 3-galactosyl-N-acetylglucosaminide 4-alpha-L-fucosyltransferase FUT3_4-galactosyl-N-acetylglucosaminide 3-alpha-L-fucosyltransferase FUT5 | FUT3_FUT5 | 1:10 | Block_5 | -1.7542 | 6.3442 | FALSE | 7.6587 | 7.5376 | -0.61866 | 0.14487 | 0.71654 | -1.0419 | 0.074506 | 1.1883
P22466 | Galanin peptides | GAL | 1:10 | Block_5 | -4.9322 | 4.8624 | FALSE | 15.2231 | 17.0859 | -0.98259 | 0.80605 | 2.033 | -0.86237 | 0.74408 | 2.2127
Q96C23 | Galactose mutarotase | GALM | 1:10 | Block_5 | -2.3843 | 5.7982 | TRUE | 7.6577 | 7.7045 | -0.31501 | 0.59911 | 1.7867 | 1.139 | 2.1267 | 3.2353
Q7Z7M9 | Polypeptide N-acetylgalactosaminyltransferase 5 | GALNT5 | 1:10 | Block_5 | -1.6775 | 9.4061 | FALSE | 11.4303 | 10.7805 | -0.25399 | 0.97324 | 1.4881 | 0.34466 | 1.3583 | 2.2943
P04406 | Glyceraldehyde-3-phosphate dehydrogenase | GAPDH | 1:10 | Block_5 | -0.60462 | 5.8369 | FALSE | 16.4396 | 29.0343 | NA | 0.86731 | 2.8327 | NA | 2.0255 | 4.7677
Q99501 | GAS2-like protein 1 | GAS2L1 | 1:10 | Block_5 | 0.45429 | 9.6885 | TRUE | NA | NA | NA | 0.5893 | 2.3613 | NA | 2.1485 | 3.8611
P32455 | Guanylate-binding protein 1 | GBP1 | 1:10 | Block_5 | -3.2849 | -0.85977 | FALSE | 13.3168 | 9.8388 | -1.4491 | -0.58268 | 0.35717 | -0.94406 | 0.21389 | 2.3473
P30047 | GTP cyclohydrolase 1 feedback regulatory protein | GCHFR | 1:10 | Block_5 | -2.4067 | 6.2793 | TRUE | 24.8526 | 17.1299 | -0.82451 | -0.079689 | 0.90914 | -0.054953 | 0.69024 | 2.3546
Q9UK05 | Growth/differentiation factor 2 | GDF2 | 1:10 | Block_5 | -4.9437 | 8.8005 | TRUE | 13.9788 | 5.2872 | -1.7833 | -0.9025 | -0.32068 | -1.8686 | -0.94274 | 0.22685
O43681 | ATPase GET3 | GET3 | 1:10 | Block_5 | -0.89679 | 8.7457 | FALSE | 10.662 | 2.2015 | -0.31308 | 0.92668 | 2.7751 | 0.97912 | 2.6162 | 3.6331
Q9BVM4 | Gamma-glutamylaminecyclotransferase | GGACT | 1:10 | Block_5 | 0.3 | 8.2113 | TRUE | 16.4223 | 12.4805 | NA | 1.5794 | 3.4803 | 1.159 | 3.7565 | 5.6118
O75223 | Gamma-glutamylcyclotransferase | GGCT | 1:10 | Block_5 | 0.15 | 7.9153 | FALSE | NA | NA | NA | 1.0202 | 2.842 | 1.0372 | 2.6474 | 4.1489
P01241 | Somatotropin | GH1 | 1:10 | Block_5 | -3.5256 | 5.2737 | TRUE | 9.7288 | 7.7898 | NA | -1.5906 | 1.997 | -3.3939 | -1.2184 | 1.6612
P10912 | Growth hormone receptor | GHR | 1:10 | Block_5 | -5.4262 | 1.9232 | TRUE | 6.6455 | 4.0875 | -0.91718 | 0.045804 | 0.57299 | -0.83899 | -0.068569 | 0.81132
Q9UBU3 | Appetite-regulating hormone | GHRL | 1:10 | Block_5 | -4.0856 | 7.9073 | TRUE | 17.8563 | 11.3128 | -0.062595 | 1.4188 | 2.8694 | -0.15725 | 1.5771 | 3.7473
Q8NHV1 | GTPase IMAP family member 7 | GIMAP7 | 1:10 | Block_5 | -1.5281 | 9.7034 | TRUE | 26.947 | 8.8932 | -0.91214 | -0.031516 | 0.77666 | -0.42852 | 0.33606 | 1.5098
Q9HC38 | Glyoxalase domain-containing protein 4 | GLOD4 | 1:10 | Block_5 | -2.4031 | 7.2921 | TRUE | 20.37 | 14.0132 | 0.17787 | 0.98492 | 2.4233 | 1.288 | 2.3796 | 3.0355
P35754 | Glutaredoxin-1 | GLRX | 1:10 | Block_5 | -1.18 | 7.6773 | FALSE | 25.2414 | 19.7702 | -0.09981 | 1.0041 | 1.9646 | 0.63885 | 2.0372 | 4.0984
Q86SX6 | Glutaredoxin-related protein 5, mitochondrial | GLRX5 | 1:10 | Block_5 | 0.53155 | 12.253 | TRUE | NA | 15.5007 | NA | 1.5245 | 4.3632 | 1.9997 | 3.8888 | 5.7425
O60234 | Glia maturation factor gamma | GMFG | 1:10 | Block_5 | 0.0047973 | 8.5917 | TRUE | 12.7465 | 14.5479 | 0.12081 | 2.1134 | 4.5402 | 0.96211 | 4.6633 | 6.646
Q9P2T1 | GMP reductase 2 | GMPR2 | 1:10 | Block_5 | -2.89 | 10.692 | TRUE | 31.3116 | 23.5829 | -0.39581 | 1.779 | 3.7603 | 1.547 | 3.8321 | 5.4219
Q8TDQ7 | Glucosamine-6-phosphate isomerase 2 | GNPDA2 | 1:10 | Block_5 | -2.1979 | 5.8859 | TRUE | 10.011 | 9.212 | -0.49824 | 0.68089 | 1.5314 | 0.58583 | 1.3461 | 2.2824
Q6P4E1 | Protein GOLM2 | GOLM2 | 1:10 | Block_5 | -2.236 | 7.5985 | TRUE | 9.0539 | 6.2827 | -0.14693 | 0.50028 | 0.96989 | 0.13324 | 0.68044 | 1.3321
P40197 | Platelet glycoprotein V | GP5 | 1:10 | Block_5 | -2.0712 | 7.2793 | TRUE | 6.8974 | 6.5845 | -0.13749 | 1.1656 | 1.8775 | 0.7992 | 2.0102 | 2.3405
Q14956 | Transmembrane glycoprotein NMB | GPNMB | 1:10 | Block_5 | -3.6 | 4.5193 | TRUE | 10.1777 | 6.8546 | 0.082135 | 0.51179 | 0.96465 | 0.11398 | 0.7306 | 1.2989
P07203 | Glutathione peroxidase 1 | GPX1 | 1:10 | Block_5 | -3.7859 | NA | NA | 11.3791 | 7.2256 | -0.6245 | 1.6277 | 2.0967 | 0.75864 | 3.0294 | 4.6146
O75791 | GRB2-related adapter protein 2 | GRAP2 | 1:10 | Block_5 | -0.034535 | 6.9752 | FALSE | 15.5825 | 6.8824 | NA | 1.1323 | 6.0934 | 1.7332 | 5.785 | 7.9475
Q9UBQ7 | Glyoxylate reductase/hydroxypyruvate reductase | GRHPR | 1:10 | Block_5 | -0.17951 | 6.5123 | TRUE | 19.5083 | 9.29 | NA | 0.64398 | 2.6648 | NA | 2.2546 | 4.0919
P08263 | Glutathione S-transferase A1 | GSTA1 | 1:10 | Block_5 | -5.3342 | 5.965 | TRUE | 8.5746 | 7.9335 | -2.31 | -0.90902 | 0.38282 | -1.7781 | -0.41794 | 1.5384
P46952 | 3-hydroxyanthranilate 3,4-dioxygenase | HAAO | 1:10 | Block_5 | -2.4059 | 8.5546 | TRUE | 11.1944 | 6.2387 | -1.4733 | -0.73064 | 0.38178 | -0.39704 | 0.82586 | 2.1762
P02008 | Hemoglobin subunit zeta | HBZ | 1:10 | Block_5 | -0.47834 | 6.0144 | TRUE | 7.8237 | 6.7461 | NA | 0.30891 | 2.1698 | NA | 0.54233 | 2.467
Q9ULI3 | Protein HEG homolog 1 | HEG1 | 1:10 | Block_5 | -2.8443 | 5.4554 | TRUE | 8.3908 | 2.2924 | -0.40886 | 0.075469 | 0.62276 | -0.27185 | 0.26765 | 0.84749
Q9BQS7 | Hephaestin | HEPH | 1:10 | Block_5 | -5.32 | 6.8699 | TRUE | 8.7904 | 8.505 | -0.69678 | 0.14029 | 0.62053 | -0.35352 | 0.20738 | 0.91879
Q8NDA2 | Hemicentin-2 | HMCN2 | 1:10 | Block_5 | -2.3077 | 6.0959 | FALSE | 13.763 | 13.2124 | -1.3239 | -0.50229 | 0.23053 | -1.5646 | -0.61949 | 0.76244
P09601 | Heme oxygenase 1 | HMOX1 | 1:10 | Block_5 | -5.89 | 6.7155 | TRUE | 10.1495 | 6.3343 | -0.54888 | 0.75569 | 1.1478 | 0.35969 | 1.1306 | 2.2708
Q9NSB8 | Homer protein homolog 2 | HOMER2 | 1:10 | Block_5 | 0.56004 | NA | NA | 1.3691 | 46.1236 | NA | 1.4331 | 4.5112 | 1.3446 | 3.7712 | 6.1458
P0DMV8 | Heat shock 70 kDa protein 1A | HSPA1A | 1:10 | Block_5 | 0.65 | 7.0967 | FALSE | NA | 7.096 | NA | 0.80711 | 2.3034 | NA | 2.2304 | 4.0683
O43464 | Serine protease HTRA2, mitochondrial | HTRA2 | 1:10 | Block_5 | -1.7452 | 5.1466 | TRUE | 9.3534 | 3.8583 | -0.25599 | 0.85638 | 2.4231 | 0.76804 | 2.1758 | 3.4847
Q2TAA2 | Isoamyl acetate-hydrolyzing esterase 1 homolog | IAH1 | 1:10 | Block_5 | -4.0333 | NA | NA | 8.1183 | 5.8799 | 0.072174 | 1.7535 | 2.9937 | 1.2169 | 3.1124 | 4.3045
Q9UMF0 | Intercellular adhesion molecule 5 | ICAM5 | 1:10 | Block_5 | -1.7206 | 6.6774 | TRUE | 13.792 | 6.1927 | -0.37022 | 0.27855 | 0.84739 | -0.40815 | 0.50671 | 1.3598
P15260 | Interferon gamma receptor 1 | IFNGR1 | 1:10 | Block_5 | -2.6029 | 6.0794 | TRUE | 13.2616 | 5.1884 | -0.39526 | 0.12884 | 0.51853 | 0.12439 | 0.64031 | 1.4018
Q8WX77 | Insulin-like growth factor-binding protein-like 1 | IGFBPL1 | 1:10 | Block_5 | -1.5787 | 4.2246 | TRUE | 9.5005 | 4.42 | -0.53576 | 0.10169 | 0.61571 | -0.1026 | 0.67589 | 1.8278
Q6UXB1 | Insulin growth factor-like family member 3 | IGFL3 | 1:10 | Block_5 | -3 | NA | NA | 17.8851 | 9.7748 | -1.081 | 0.032999 | 0.81335 | -0.4078 | 0.6769 | 1.7274
Q969P0 | Immunoglobulin superfamily member 8 | IGSF8 | 1:10 | Block_5 | -2.9206 | 8.005 | TRUE | 10.3005 | 4.078 | -0.40717 | 0.36229 | 0.78032 | 0.37634 | 1.0242 | 2.0414
Q9P2J2 | Protein turtle homolog A | IGSF9 | 1:10 | Block_5 | -2.7577 | 10.0746 | TRUE | 9.9396 | 8.0759 | -2.1451 | -0.17723 | 1.6914 | -1.3327 | 0.13985 | 1.8674
Q08334 | Interleukin-10 receptor subunit beta | IL10RB | 1:10 | Block_5 | -2.94 | 4.5628 | TRUE | 11.0763 | 4.3071 | -0.5618 | 0.25163 | 0.66978 | 0.067484 | 0.70357 | 1.6286
P78552 | Interleukin-13 receptor subunit alpha-1 | IL13RA1 | 1:10 | Block_5 | -2.5 | 2.7057 | TRUE | 12.4647 | 9.342 | -0.39071 | 0.12449 | 0.65357 | -0.31312 | 0.1518 | 0.69659
Q14116 | Interleukin-18 | IL18 | 1:10 | Block_5 | -3.92 | 9.7683 | TRUE | 9.2097 | 8.5709 | -0.98218 | 0.0072845 | 0.57861 | 0.17991 | 1.2447 | 4.1123
Q9UHD0 | Interleukin-19 | IL19 | 1:10 | Block_5 | -1.9772 | 10.5314 | FALSE | 11.9772 | 6.1971 | -1.592 | -0.90054 | 0.91113 | -0.95303 | 0.41976 | 2.0882
P27930 | Interleukin-1 receptor type 2 | IL1R2 | 1:10 | Block_5 | -4.5764 | 3.6668 | TRUE | 3.8366 | 5.1792 | -0.484 | 0.12346 | 0.33545 | -0.29648 | 0.21293 | 0.59749
P18510 | Interleukin-1 receptor antagonist protein | IL1RN | 1:10 | Block_5 | -1.03 | 6.3566 | TRUE | 43.6593 | 34.7666 | -0.48998 | 0.035292 | 0.60061 | -0.027823 | 0.78444 | 2.7669
P40189 | Interleukin-6 receptor subunit beta | IL6ST | 1:10 | Block_5 | -3.2017 | 3.2229 | FALSE | 6.3658 | 3.9058 | -0.33025 | 0.010462 | 0.37476 | -0.1572 | 0.24452 | 0.74475
P09529 | Inhibin beta B chain | INHBB | 1:10 | Block_5 | -1.5843 | 5.9197 | TRUE | 15.4151 | 6.8964 | -0.60518 | 0.035603 | 0.94781 | -0.11767 | 0.78657 | 2.3027
P01308 | Insulin C-peptide | INS-CPEPTIDE | 1:10 | Block_5 | -7.3022 | NA | NA | 13.5326 | 8.7802 | -2.1769 | 0.5727 | 1.8652 | -3.5091 | 0.92623 | 2.6539
Q14653 | Interferon regulatory factor 3 | IRF3 | 1:10 | Block_5 | -1.906 | NA | NA | 13.0121 | 7.1077 | -0.3884 | 1.2318 | 3.8222 | 1.7445 | 4.1119 | 5.6434
P18084 | Integrin beta-5 | ITGB5 | 1:10 | Block_5 | -3.0321 | 5.8662 | FALSE | 11.0878 | 4.2492 | -0.25411 | 0.64706 | 1.3142 | 0.030457 | 0.70019 | 1.2985
P57087 | Junctional adhesion molecule B | JAM2 | 1:10 | Block_5 | -2.0652 | 6.3887 | TRUE | 11.2245 | 6.9572 | -0.21589 | 0.34837 | 0.72167 | 0.087103 | 0.82014 | 1.8929
Q9BX67 | Junctional adhesion molecule C | JAM3 | 1:10 | Block_5 | -0.71804 | 9.8992 | TRUE | 17.7384 | 13.1745 | -0.14243 | 0.69562 | 1.9509 | 0.82996 | 1.5748 | 2.7656
P35968 | Vascular endothelial growth factor receptor 2 | KDR | 1:10 | Block_5 | -3.5359 | 5.4165 | TRUE | 6.9253 | 3.5201 | -0.67428 | 0.13542 | 0.52788 | -0.24232 | 0.30518 | 0.82755
P21583 | Kit ligand | KITLG | 1:10 | Block_5 | -5.0276 | 8.0091 | TRUE | 9.6507 | 5.1973 | -0.48703 | 0.35719 | 1.0759 | -0.19988 | 0.70389 | 1.4519
P06870 | Kallikrein-1 | KLK1 | 1:10 | Block_5 | -4.1884 | 5.9533 | TRUE | 9.5471 | 5.8396 | -2.6116 | -1.5698 | -0.63992 | -2.3034 | -0.86298 | 2.1425
O43240 | Kallikrein-10 | KLK10 | 1:10 | Block_5 | -3.69 | 6.8446 | TRUE | 68.9745 | 37.4845 | -0.74 | -0.1594 | 0.67769 | -0.8769 | -0.0039272 | 0.82947
Q9UBX7 | Kallikrein-11 | KLK11 | 1:10 | Block_5 | -1.5 | 6.7806 | TRUE | 32.8806 | 27.8354 | -0.33894 | -0.030747 | 0.40489 | -0.14116 | 0.43936 | 1.2739
Q92876 | Kallikrein-6 | KLK6 | 1:10 | Block_5 | -6.09 | 3.6425 | TRUE | 10.0507 | 6.2921 | -0.3047 | 0.25347 | 0.80336 | -0.045595 | 0.73959 | 1.4798
P49862 | Kallikrein-7 | KLK7 | 1:10 | Block_5 | -2 | 3.627 | TRUE | 39.4057 | 31.5385 | -0.93358 | -0.42424 | 0.48135 | -1.7981 | -0.35351 | 0.38342
O60259 | Kallikrein-8 | KLK8 | 1:10 | Block_5 | -3.91 | 5.0533 | TRUE | 14.3597 | 9.5796 | -0.83386 | -0.12096 | 0.54714 | -0.60538 | 0.11888 | 0.89457
Q16773 | Kynurenine--oxoglutarate transaminase 1 | KYAT1 | 1:10 | Block_5 | -3.7 | 9.9711 | TRUE | 19.2099 | 17.7248 | -0.9725 | 0.80729 | 2.5823 | 0.72868 | 2.4333 | 3.9794
P32004 | Neural cell adhesion molecule L1 | L1CAM | 1:10 | Block_5 | -4.7147 | 4.4173 | TRUE | 10.2879 | 5.6324 | -0.4503 | 0.081539 | 0.69202 | -0.49361 | 0.054964 | 0.79939
Q96EM0 | Trans-3-hydroxy-L-proline dehydratase | L3HYPDH | 1:10 | Block_5 | -1.7085 | 9.2607 | TRUE | 13.6606 | 11.3124 | -0.21747 | 1.1181 | 4.352 | 1.291 | 3.0468 | 4.8673
Q53H82 | Endoribonuclease LACTB2 | LACTB2 | 1:10 | Block_5 | -0.32543 | 7.2186 | TRUE | 4.7869 | 3.471 | -0.12729 | 0.86925 | 2.5832 | 1.1093 | 3.0371 | 4.0627
Q6GTX8 | Leukocyte-associated immunoglobulin-like receptor 1 | LAIR1 | 1:10 | Block_5 | -4.6434 | 4.521 | TRUE | 10.0913 | 8.5373 | -1.501 | -0.97227 | 0.0014409 | -1.0633 | -0.013882 | 1.2783
Q16363 | Laminin subunit alpha-4 | LAMA4 | 1:10 | Block_5 | -4.7306 | 8.9896 | TRUE | 14.7047 | 7.3692 | -0.41001 | 0.24124 | 0.85097 | 0.17732 | 1.1376 | 1.8946
P13473 | Lysosome-associated membrane glycoprotein 2 | LAMP2 | 1:10 | Block_5 | -4.4241 | 3.973 | TRUE | 6.4504 | 4.4183 | -0.42252 | 0.077232 | 0.43051 | -0.19384 | 0.22421 | 0.58241
Q14739 | Delta(14)-sterol reductase LBR | LBR | 1:10 | Block_5 | -1.8962 | 5.5784 | FALSE | 7.2738 | 2.5558 | -1.4423 | -0.49236 | 1.0721 | -0.035707 | 1.2881 | 3.2201
Q13094 | Lymphocyte cytosolic protein 2 | LCP2 | 1:10 | Block_5 | -0.53728 | 9.0199 | FALSE | 12.2771 | 8.5127 | 0.14543 | 3.5518 | 6.77 | 2.0066 | 6.7729 | 8.3513
O14960 | Leukocyte cell-derived chemotaxin-2 | LECT2 | 1:10 | Block_5 | -7.4472 | 4.6339 | FALSE | 10.8825 | 10.4028 | -4.4692 | -3.0648 | -2.1264 | -3.9384 | -2.6489 | -0.92495
Q6P5S2 | Protein LEG1 homolog | LEG1 | 1:10 | Block_5 | -4.48 | 5.4222 | TRUE | 12.081 | 13.5756 | -1.9971 | -0.55091 | 0.14599 | -1.852 | -0.94761 | 0.13579
P41159 | Leptin | LEP | 1:10 | Block_5 | -4.3192 | 4.59 | TRUE | 7.3272 | 10.1483 | -2.3124 | -0.4394 | 0.76926 | -2.2105 | 0.05575 | 1.9689
P48357 | Leptin receptor | LEPR | 1:10 | Block_5 | -2.7644 | 5.1507 | FALSE | 6.387 | 7.2335 | -0.18706 | 0.1482 | 0.77019 | -0.25254 | 0.26385 | 0.89945
P09382 | Galectin-1 | LGALS1 | 1:10 | Block_5 | -4.0885 | 5.6548 | TRUE | 6.3804 | 8.8931 | 0.002811 | 0.51996 | 0.9846 | 0.28898 | 1.4613 | 2.31
P56470 | Galectin-4 | LGALS4 | 1:10 | Block_5 | -4.2842 | 7.5789 | TRUE | 9.7927 | 5.7139 | -0.89671 | 0.075901 | 1.003 | 0.041087 | 1.1285 | 2.5256
O00214 | Galectin-8 | LGALS8 | 1:10 | Block_5 | -3.31 | 9.0429 | TRUE | 11.8356 | 10.9776 | -0.45614 | 0.58628 | 2.1355 | 0.89581 | 2.3303 | 3.8994
O00182 | Galectin-9 | LGALS9 | 1:10 | Block_5 | -7.7 | 1.2849 | FALSE | 9.7372 | 7.3748 | -0.29266 | 0.19848 | 0.82929 | 0.21754 | 0.99854 | 2.1143
Q99538 | Legumain | LGMN | 1:10 | Block_5 | -5.63 | 5.0549 | TRUE | 11.6924 | 8.4346 | -0.62154 | 0.14232 | 0.82519 | -0.13628 | 0.42802 | 1.048
Q9H008 | Phospholysine phosphohistidine inorganic pyrophosphate phosphatase | LHPP | 1:10 | Block_5 | -4 | 7.6817 | TRUE | 15.1665 | 7.2299 | -0.0012769 | 1.0911 | 2.2992 | 0.98945 | 2.1983 | 3.211
A6NI73 | Leukocyte immunoglobulin-like receptor subfamily A member 5 | LILRA5 | 1:10 | Block_5 | -4.2506 | 5.5739 | TRUE | 10.5998 | 5.4487 | -1.0682 | -0.24402 | 0.70715 | -0.54333 | 0.29599 | 0.94472
Q6PI73 | Leukocyte immunoglobulin-like receptor subfamily A member 6 | LILRA6 | 1:10 | Block_5 | -3 | 5.9804 | TRUE | 20.8385 | 9.6219 | -1.3689 | 0.11776 | 1.441 | -0.82233 | 0.41355 | 1.537
P06858 | Lipoprotein lipase | LPL | 1:10 | Block_5 | -2.9187 | 3.2241 | FALSE | 8.519 | 7.7519 | -0.38224 | 0.86089 | 1.4441 | -0.60503 | 0.88249 | 2.0983
Q86VZ4 | Low-density lipoprotein receptor-related protein 11 | LRP11 | 1:10 | Block_5 | -4.9679 | 5.1172 | TRUE | 8.3813 | 4.4851 | -0.5864 | 0.22287 | 0.86497 | -0.23742 | 0.6735 | 1.7366
P30533 | Alpha-2-macroglobulin receptor-associated protein | LRPAP1 | 1:10 | Block_5 | -4.3152 | 7.5817 | TRUE | 13.5862 | 5.4759 | -0.55515 | 0.28409 | 1.0465 | 0.45381 | 1.0531 | 2.0014
P36941 | Tumor necrosis factor receptor superfamily member 3 | LTBR | 1:10 | Block_5 | -4.779 | 5.5463 | TRUE | 10.5446 | 7.8412 | -0.57434 | -0.0063596 | 0.43971 | -0.052831 | 0.62382 | 1.939
Q14210 | Lymphocyte antigen 6D | LY6D | 1:10 | Block_5 | -4.41 | 5.5494 | TRUE | 23.7969 | 20.7282 | -0.39568 | 0.012016 | 0.75104 | 0.089651 | 0.78258 | 1.9967
O95274 | Ly6/PLAUR domain-containing protein 3 | LYPD3 | 1:10 | Block_5 | -4.32 | 4.4572 | TRUE | 18.941 | 11.1895 | -0.63875 | -0.0013055 | 0.75051 | -0.63923 | 0.2042 | 0.9294
Q9NQ48 | Leucine zipper transcription factor-like protein 1 | LZTFL1 | 1:10 | Block_5 | 0.97 | 10.9155 | TRUE | 18.5626 | NA | NA | NA | 2.7973 | NA | 2.6743 | 4.0527
Q7Z304 | MAM domain-containing protein 2 | MAMDC2 | 1:10 | Block_5 | -0.77353 | 6.8144 | TRUE | 7.2924 | 8.7595 | -0.37124 | 0.027608 | 0.5592 | -0.39962 | 0.47743 | 1.6958
Q9Y2E5 | Epididymis-specific alpha-mannosidase | MAN2B2 | 1:10 | Block_5 | -1.2444 | 4.0566 | FALSE | 6.8859 | 6.6092 | -0.47186 | 0.15293 | 1.1199 | -0.7105 | 0.13855 | 1.0016
P55145 | Mesencephalic astrocyte-derived neurotrophic factor | MANF | 1:10 | Block_5 | 0.74735 | 10.5299 | TRUE | NA | 5.8303 | NA | 1.3649 | 6.5616 | 2.2509 | 6.6987 | 8.1823
Q9H8J5 | MANSC domain-containing protein 1 | MANSC1 | 1:10 | Block_5 | -3.4453 | 5.0324 | TRUE | 14.8193 | 9.9253 | -0.39777 | 0.23476 | 0.93098 | -0.0088731 | 0.79186 | 1.696
Q02750 | Dual specificity mitogen-activated protein kinase kinase 1 | MAP2K1 | 1:10 | Block_5 | 0.056197 | 8.0221 | FALSE | 9.8519 | 10.7468 | NA | 0.34424 | 1.2095 | 0.39926 | 1.346 | 2.6939
Q9UEW3 | Macrophage receptor MARCO | MARCO | 1:10 | Block_5 | -2.3015 | 5.263 | TRUE | 7.1484 | 9.3917 | -0.57456 | 0.13318 | 0.61406 | -0.31143 | 0.27401 | 0.78745
P56192 | Methionine--tRNA ligase, cytoplasmic | MARS1 | 1:10 | Block_5 | 0.52736 | 6.7178 | FALSE | NA | NA | NA | NA | 2.3268 | 0.63153 | 2.4827 | 4.1959
O00339 | Matrilin-2 | MATN2 | 1:10 | Block_5 | -7.5 | 4.9354 | TRUE | 10.1523 | 6.4317 | -0.48527 | 0.29191 | 0.65157 | 0.10536 | 0.59383 | 1.1519
Q8NI22 | Multiple coagulation factor deficiency protein 2 | MCFD2 | 1:10 | Block_5 | -3.2079 | 6.4288 | TRUE | 9.4532 | 5.5153 | -0.31739 | 0.31215 | 1.6673 | 0.80752 | 1.6886 | 2.2548
P21741 | Midkine | MDK | 1:10 | Block_5 | -0.32437 | 7.1601 | TRUE | 6.4537 | 6.4197 | 0.083228 | 0.66416 | 0.99791 | 0.70404 | 1.4349 | 2.6676
Q96KG7 | Multiple epidermal growth factor-like domains protein 10 | MEGF10 | 1:10 | Block_5 | -9.08 | 2.9332 | TRUE | 8.3369 | 6.5003 | -0.75334 | -0.09383 | 0.77017 | -0.79161 | -0.062491 | 0.93682
P08582 | Melanotransferrin | MELTF | 1:10 | Block_5 | -5 | 7.1208 | TRUE | 7.3693 | 3.7823 | -0.49577 | 0.099214 | 0.89532 | -0.74652 | 0.24521 | 0.79066
Q9BUN1 | Protein MENT | MENT | 1:10 | Block_5 | -3.8409 | 2.3536 | FALSE | 5.7735 | 3.4979 | -0.15926 | 0.15206 | 0.51585 | -0.41723 | 0.14515 | 0.57548
Q16820 | Meprin A subunit beta | MEP1B | 1:10 | Block_5 | -0.93728 | 3.4102 | FALSE | 8.143 | 8.0414 | NA | 0.48049 | 1.8871 | NA | -0.26755 | 1.7419
Q9NQ76 | Matrix extracellular phosphoglycoprotein | MEPE | 1:10 | Block_5 | -3.75 | 5.0707 | TRUE | 11.7222 | 6.7155 | 0.38553 | 1.0006 | 1.8352 | -0.14463 | 1.3644 | 1.9767
Q12866 | Tyrosine-protein kinase Mer | MERTK | 1:10 | Block_5 | -2.5 | 6.9841 | FALSE | 28.4979 | 39.711 | -0.4481 | 0.39221 | 0.99414 | 0.016772 | 1.0857 | 2.1129
P55083 | Microfibril-associated glycoprotein 4 | MFAP4 | 1:10 | Block_5 | -2.1184 | 7.1555 | FALSE | 8.0186 | 9.028 | -0.83017 | -0.28161 | 0.4346 | -0.50319 | 0.22097 | 1.0999
Q10469 | Alpha-1,6-mannosyl-glycoprotein 2-beta-N-acetylglucosaminyltransferase | MGAT2 | 1:10 | Block_5 | -3.1766 | 5.4504 | FALSE | 12.4073 | 6.785 | -0.22224 | 0.41998 | 1.1588 | -0.1271 | 0.48819 | 1.0659
Q99685 | Monoglyceride lipase | MGLL | 1:10 | Block_5 | 2.0036 | 7.6308 | FALSE | NA | NA | NA | NA | 2.2419 | NA | NA | 3.1358
Q16674 | Melanoma-derived growth regulatory protein | MIA | 1:10 | Block_5 | -4.96 | 4.5701 | TRUE | 13.2087 | 10.499 | -0.85523 | 0.04455 | 1.0909 | -0.33372 | 0.40927 | 1.2541
P03956 | Interstitial collagenase | MMP1 | 1:10 | Block_5 | -1.1797 | 8.173 | TRUE | 9.2279 | 13.6364 | -0.78365 | 0.60994 | 2.5265 | 0.063969 | 2.3573 | 4.4606
P09238 | Stromelysin-2 | MMP10 | 1:10 | Block_5 | -4.6973 | 7.0954 | TRUE | 7.5101 | 7.0864 | -1.1951 | -0.35288 | 0.64669 | -0.61622 | 0.35244 | 1.4467
P51511 | Matrix metalloproteinase-15 | MMP15 | 1:10 | Block_5 | 0.51592 | NA | NA | 7.1928 | 7.9029 | NA | NA | 1.1814 | NA | NA | 0.74663
P09237 | Matrilysin | MMP7 | 1:10 | Block_5 | -1.718 | 6.2393 | TRUE | 8.619 | 8.2736 | -1.5546 | -0.2501 | 0.78218 | -1.0616 | 0.49526 | 1.6732
P22894 | Neutrophil collagenase | MMP8 | 1:10 | Block_5 | -2.7632 | 8.3384 | FALSE | 10.2557 | 5.3659 | -1.2357 | -0.11091 | 0.89662 | -0.7503 | 0.79578 | 5.2148
Q9H8L6 | Multimerin-2 | MMRN2 | 1:10 | Block_5 | -4.4125 | 0.51169 | FALSE | 26.4464 | 13.8585 | -0.36042 | 0.25381 | 0.72519 | -0.13201 | 0.60958 | 1.521
P41218 | Myeloid cell nuclear differentiation antigen | MNDA | 1:10 | Block_5 | -2.6219 | 6.2783 | FALSE | 8.4128 | 6.9279 | NA | -1.647 | 0.63351 | NA | -1.3331 | 2.0285
O96007 | Molybdopterin synthase catalytic subunit | MOCS2 | 1:10 | Block_5 | -0.48 | 8.7854 | FALSE | 20.8168 | 21.5864 | -0.44267 | 0.70048 | 1.7856 | 0.42414 | 1.7353 | 2.9954
O95866 | Megakaryocyte and platelet inhibitory receptor G6b | MPIG6B | 1:10 | Block_5 | -1.2446 | 8.6002 | TRUE | 19.4482 | 14.7317 | -0.67895 | 1.8995 | 5.1038 | 1.4103 | 4.2893 | 5.5998
Q1L6U9 | Prostate-associated microseminoprotein | MSMP | 1:10 | Block_5 | -2.677 | 7.6957 | TRUE | 19.867 | 7.5839 | -1.0981 | -0.3514 | 0.65808 | -1.0984 | -0.10141 | 0.96202
P21757 | Macrophage scavenger receptor types I and II | MSR1 | 1:10 | Block_5 | -5.45 | 6.8058 | TRUE | 10.0978 | 5.2182 | -0.93587 | 0.06014 | 0.92337 | -0.18716 | 0.69313 | 1.7161
Q9UJ68 | Mitochondrial peptide methionine sulfoxide reductase | MSRA | 1:10 | Block_5 | -2.92 | 8.5118 | TRUE | 12.8449 | 12.8091 | -0.73439 | 0.79702 | 2.8664 | 0.70703 | 3.1681 | 4.6828
Q00872 | Myosin-binding protein C, slow-type | MYBPC1 | 1:10 | Block_5 | -0.81081 | 9.509 | FALSE | 10.703 | 7.0125 | NA | -0.34689 | 1.6596 | -0.79048 | 0.79927 | 3.0382
Q969H8 | Myeloid-derived growth factor | MYDGF | 1:10 | Block_5 | -0.96894 | 9.0776 | TRUE | 23.6592 | 22.4136 | -0.26387 | 1.5758 | 4.504 | 2.3239 | 5.1543 | 6.63
Q5VTT5 | Myomesin-3 | MYOM3 | 1:10 | Block_5 | -3.1168 | 6.2169 | TRUE | 11.6916 | 5.4758 | -1.8621 | -0.86239 | 1.2995 | -1.6815 | 0.071227 | 1.7507
Q8WU39 | Marginal zone B- and B1-cell-specific protein | MZB1 | 1:10 | Block_5 | -4.4021 | 4.9594 | TRUE | 11.3819 | 6.4835 | -1.0686 | -0.27311 | 0.6847 | -1.0002 | 0.057559 | 1.4307
O95544 | NAD kinase | NADK | 1:10 | Block_5 | -2.2929 | 8.199 | FALSE | 10.9649 | 4.3532 | -1.1176 | 0.10148 | 1.008 | -0.37309 | 0.68779 | 3.3854
Q9UK23 | N-acetylglucosamine-1-phosphodiester alpha-N-acetylglucosaminidase | NAGPA | 1:10 | Block_5 | -5.1761 | 4.1457 | TRUE | 8.4214 | 7.9108 | -0.14662 | 0.31069 | 0.65733 | -0.090541 | 0.40114 | 0.89867
P41271 | Neuroblastoma suppressor of tumorigenicity 1 | NBL1 | 1:10 | Block_5 | -3.2639 | 3.0605 | TRUE | 8.8896 | 3.9022 | -0.41367 | 0.050664 | 0.64051 | 0.078467 | 0.71994 | 1.4794
O14594 | Neurocan core protein | NCAN | 1:10 | Block_5 | -1.0333 | 6.8449 | TRUE | 13.7426 | 4.7463 | -0.1101 | 0.50331 | 1.1002 | -0.25974 | 0.50271 | 1.2223
O43639 | Cytoplasmic protein NCK2 | NCK2 | 1:10 | Block_5 | 0.36915 | 7.57 | FALSE | 15.3718 | 6.993 | NA | 1.2471 | 4.8371 | 1.2818 | 4.6209 | 6.0296
Q92692 | Nectin-2 | NECTIN2 | 1:10 | Block_5 | -4.133 | 4.9406 | TRUE | 8.2864 | 4.618 | -0.41226 | 0.24433 | 0.70223 | 0.26704 | 0.95586 | 2.046
Q96NY8 | Nectin-4 | NECTIN4 | 1:10 | Block_5 | -0.7407 | 7.2067 | FALSE | 14.9633 | 7.4122 | -0.39733 | 0.10522 | 0.9081 | -0.10537 | 0.38476 | 1.4772
Q92832 | Protein kinase C-binding protein NELL1 | NELL1 | 1:10 | Block_5 | -2.756 | 6.0294 | FALSE | 9.1613 | 7.6968 | -0.56995 | 0.1022 | 1.1327 | -0.75636 | 0.3326 | 1.0994
Q99435 | Protein kinase C-binding protein NELL2 | NELL2 | 1:10 | Block_5 | -2.7153 | 4.1553 | TRUE | 9.1528 | 6.193 | -0.47088 | 0.075728 | 0.53451 | -0.59195 | 0.042839 | 0.89249
Q15653 | NF-kappa-B inhibitor beta | NFKBIB | 1:10 | Block_5 | 0.91384 | 6.3117 | FALSE | NA | NA | NA | NA | 0.99373 | NA | NA | NA
P08138 | Tumor necrosis factor receptor superfamily member 16 | NGFR | 1:10 | Block_5 | 0.66168 | 7.753 | TRUE | NA | NA | NA | NA | 0.66448 | NA | NA | 1.0098
O14745 | Na(+)/H(+) exchange regulatory cofactor NHE-RF1 | NHERF1 | 1:10 | Block_5 | -1.1595 | 4.6723 | FALSE | 26.5627 | 9.5217 | -0.4718 | 1.4429 | 4.299 | 1.5629 | 5.1072 | 6.3302
Q5JS37 | NHL repeat-containing protein 3 | NHLRC3 | 1:10 | Block_5 | -5 | 6.5137 | TRUE | 9.778 | 9.0972 | 0.45474 | 1.2055 | 1.9065 | 0.66827 | 1.4749 | 2.427
Q9NQR4 | Omega-amidase NIT2 | NIT2 | 1:10 | Block_5 | -3.52 | 6.7227 | TRUE | 9.5232 | 5.9907 | -0.76642 | 1.7898 | 3.5829 | 1.0916 | 3.5691 | 4.9881
Q13232 | Nucleoside diphosphate kinase 3 | NME3 | 1:10 | Block_5 | -5.653 | 4.5602 | TRUE | 9.9668 | 8.6291 | -0.20902 | 0.26642 | 0.50332 | -0.20777 | 0.43897 | 1.0439
O60936 | Nucleolar protein 3 | NOL3 | 1:10 | Block_5 | -0.59852 | 8.3366 | TRUE | 9.7062 | 9.3911 | -0.27731 | 0.48434 | 0.9985 | 0.3728 | 1.2556 | 2.5969
Q9NQX5 | Neural proliferation differentiation and control protein 1 | NPDC1 | 1:10 | Block_5 | -4.4 | 7.6305 | TRUE | 9.3416 | 10.2518 | 0.00062629 | 0.46342 | 0.95956 | 0.33589 | 1.2 | 2.5644
Q9BXD5 | N-acetylneuraminate lyase | NPL | 1:10 | Block_5 | -2.074 | 8.8831 | TRUE | 10.2281 | 2.9458 | -0.01487 | 1.2595 | 2.6721 | 1.1351 | 2.3861 | 3.6065
O95502 | Neuronal pentraxin receptor | NPTXR | 1:10 | Block_5 | -5.8957 | 6.7288 | TRUE | 9.8689 | 3.8671 | -0.83704 | -0.37165 | 0.2908 | -0.75825 | -0.085331 | 0.64966
Q92823 | Neuronal cell adhesion molecule | NRCAM | 1:10 | Block_5 | -4 | 4.3552 | TRUE | 7.3694 | 6.6605 | -0.26025 | 0.27771 | 0.72308 | -0.089531 | 0.49551 | 1.3604
Q9UNZ2 | NSFL1 cofactor p47 | NSFL1C | 1:10 | Block_5 | -1.7794 | 11.2012 | FALSE | 16.6384 | 6.4411 | 0.80274 | 2.1908 | 4.6446 | 2.6208 | 4.4749 | 5.7634
P21589 | 5'-nucleotidase | NT5E | 1:10 | Block_5 | -2.3912 | 5.7367 | FALSE | 9.3053 | 7.4223 | -0.74331 | -0.13641 | 0.43761 | -0.67336 | 0.07116 | 1.0024
NT-proBNP | N-terminal prohormone of brain natriuretic peptide | NT-proBNP | 1:10 | Block_5 | -0.0073612 | 10.7099 | TRUE | 12.8887 | 9.4429 | NA | 0.81298 | 1.945 | 0.023924 | 2.1074 | 5.5996
Q16620 | BDNF/NT-3 growth factors receptor | NTRK2 | 1:10 | Block_5 | -1 | 3.4926 | TRUE | 10.3601 | 11.5677 | -0.36259 | 0.23047 | 0.65158 | -0.15317 | 0.23553 | 0.73467
Q16288 | NT-3 growth factor receptor | NTRK3 | 1:10 | Block_5 | -2.0776 | 4.2339 | TRUE | 9.7401 | 8.9081 | -0.4826 | 0.10384 | 0.50355 | -0.38131 | 0.14553 | 0.69992
Q6UX06 | Olfactomedin-4 | OLFM4 | 1:10 | Block_5 | -2.83 | 0.88105 | FALSE | 20.8613 | 9.3214 | -2.6806 | -1.3301 | 0.17898 | -1.8394 | 0.16658 | 2.7865
P78380 | Oxidized low-density lipoprotein receptor 1 | OLR1 | 1:10 | Block_5 | -1.5821 | 8.4041 | TRUE | 7.8059 | 7.1572 | -0.2184 | 0.42572 | 1.5497 | -0.14992 | 0.86684 | 3.6025
Q99983 | Osteomodulin | OMD | 1:10 | Block_5 | -2.6568 | 3.9719 | TRUE | 11.0906 | 4.9591 | -0.31363 | 0.89333 | 1.7279 | -0.42064 | 1.0024 | 1.9041
Q8IYS5 | Osteoclast-associated immunoglobulin-like receptor | OSCAR | 1:10 | Block_5 | -1.8593 | 3.9986 | TRUE | 11.4437 | 7.1515 | -0.6173 | 0.21617 | 0.72384 | -0.43322 | 0.47483 | 1.1657
P01178 | Oxytocin-neurophysin 1 | OXT | 1:10 | Block_5 | -3.6969 | 6.8224 | TRUE | 15.2174 | 11.8664 | -1.609 | 0.37124 | 2.3221 | -1.4409 | 0.77886 | 3.196
P07237 | Protein disulfide-isomerase | P4HB | 1:10 | Block_5 | -3.091 | 4.3156 | TRUE | 8.791 | 5.4148 | -0.32938 | 0.59273 | 1.0397 | 0.41697 | 1.0893 | 1.6798
Q9BUH6 | Protein PAXX | PAXX | 1:10 | Block_5 | -1.5443 | 8.0275 | TRUE | 15.2777 | 10.232 | -0.29465 | 0.61448 | 1.5256 | 0.48063 | 1.6173 | 2.7919
P61457 | Pterin-4-alpha-carbinolamine dehydratase | PCBD1 | 1:10 | Block_5 | -1.8 | 6.0009 | TRUE | 8.7112 | 7.7559 | -0.86492 | -0.28799 | 0.69723 | -0.34221 | 0.32951 | 1.5331
O14917 | Protocadherin-17 | PCDH17 | 1:10 | Block_5 | -1.1388 | 6.7472 | FALSE | 5.9863 | 8.5918 | -0.30061 | 0.40104 | 0.96733 | -0.19127 | 0.68191 | 1.6562
Q13442 | 28 kDa heat- and acid-stable phosphoprotein | PDAP1 | 1:10 | Block_5 | 1.1357 | 9.616 | TRUE | NA | NA | NA | 1.2401 | 4.6024 | 1.3246 | 4.7024 | 6.1834
O14737 | Programmed cell death protein 5 | PDCD5 | 1:10 | Block_5 | -2.7648 | 8.0543 | TRUE | 10.0796 | 5.3429 | -0.38921 | 0.97106 | 2.7115 | 1.8168 | 2.9824 | 4.1599
O75340 | Programmed cell death protein 6 | PDCD6 | 1:10 | Block_5 | -2.8033 | 3.5524 | FALSE | 12.0108 | 9.0617 | -0.36538 | 0.73839 | 1.1832 | 0.33552 | 1.2932 | 1.9156
O76074 | cGMP-specific 3',5'-cyclic phosphodiesterase | PDE5A | 1:10 | Block_5 | 0.5 | 5.5178 | FALSE | 16.318 | 2.9909 | NA | 0.9706 | 2.6637 | NA | 2.7018 | 4.2705
P01127 | Platelet-derived growth factor subunit B | PDGFB | 1:10 | Block_5 | -2.8484 | 6.8334 | TRUE | 11.5952 | 7.5458 | 0.016892 | 1.9772 | 3.6619 | 1.3289 | 3.5909 | 5.0714
Q9GZP0 | Platelet-derived growth factor D | PDGFD | 1:10 | Block_5 | -2.7716 | NA | NA | 8.2431 | 9.3092 | -1.0083 | 1.0402 | 2.4775 | 0.37195 | 2.6963 | 4.1104
P16234 | Platelet-derived growth factor receptor alpha | PDGFRA | 1:10 | Block_5 | -1.4546 | 4.7398 | TRUE | 11.7469 | 6.5036 | -0.22636 | 0.31954 | 0.69117 | -0.01721 | 0.55455 | 1.338
P30101 | Protein disulfide-isomerase A3 | PDIA3 | 1:10 | Block_5 | 0.2927 | 5.68 | FALSE | 7.8327 | 9.1866 | NA | 1.15 | 3.0525 | 1.2847 | 2.3519 | 3.9166
Q9NR12 | PDZ and LIM domain protein 7 | PDLIM7 | 1:10 | Block_5 | -0.078088 | 7.1492 | TRUE | 7.3777 | 2.4481 | NA | 0.60103 | 4.7825 | 1.6706 | 4.8734 | 6.3983
P01210 | Proenkephalin-A | PENK | 1:10 | Block_5 | -3.48 | 4.9706 | FALSE | 12.3611 | 9.3362 | -0.65935 | 0.047609 | 0.71822 | -0.13135 | 0.58747 | 1.7009
P52209 | 6-phosphogluconate dehydrogenase, decarboxylating | PGD | 1:10 | Block_5 | -2.3895 | NA | NA | 9.8665 | 6.7752 | -0.16475 | 1.1159 | 2.2372 | 1.1071 | 2.5518 | 3.9978
Q8IUZ5 | 5-phosphohydroxy-L-lysine phospho-lyase | PHYKPL | 1:10 | Block_5 | -2.4436 | 2.8322 | FALSE | 11.98 | 7.7669 | -0.48912 | 0.29225 | 0.8813 | 0.33599 | 1.0744 | 1.8825
Q9UKJ0 | Paired immunoglobulin-like type 2 receptor beta | PILRB | 1:10 | Block_5 | -3.8902 | 7.5717 | TRUE | 9.0636 | 8.3563 | -0.65977 | 0.3386 | 0.96851 | -0.64084 | 0.95537 | 2.1774
A6NC86 | phospholipase A2 inhibitor and Ly6/PLAUR domain-containing protein | PINLYP | 1:10 | Block_5 | -4 | 9.4093 | TRUE | 15.3205 | 6.238 | -1.0197 | 0.10117 | 1.1451 | -0.98355 | -0.042668 | 1.0152
P98161 | Polycystin-1 | PKD1 | 1:10 | Block_5 | -2.6209 | 7.6465 | TRUE | 9.6029 | 9.5382 | -0.43292 | 0.18109 | 0.77321 | -0.066513 | 0.49801 | 1.2132
P30613 | Pyruvate kinase PKLR | PKLR | 1:10 | Block_5 | -1.7574 | 8.7153 | TRUE | 17.8804 | 10.6206 | 0.9141 | 1.8413 | 2.7596 | 0.71864 | 1.9123 | 3.2266
Q9BZM1 | Group XIIA secretory phospholipase A2 | PLA2G12A | 1:10 | Block_5 | -1.752 | NA | NA | 14.6111 | 11.2812 | 0.011992 | 0.74202 | 3.0765 | 0.84203 | 2.7376 | 4.3975
Q8NCC3 | Phospholipase A2 group XV | PLA2G15 | 1:10 | Block_5 | -3.9335 | 4.8615 | TRUE | 5.2806 | 7.9945 | -0.33727 | 0.080539 | 0.58723 | -0.0099205 | 0.4398 | 1.089
P47712 | Cytosolic phospholipase A2 | PLA2G4A | 1:10 | Block_5 | -0.0070813 | 8.3142 | TRUE | 19.3481 | 14.4638 | 0.32698 | 1.8257 | 4.5927 | 2.013 | 4.6966 | 6.3102
Q03405 | Urokinase plasminogen activator surface receptor | PLAUR | 1:10 | Block_5 | -6.6 | 4.8897 | TRUE | 12.5931 | 8.3241 | -0.53207 | 0.44055 | 0.90861 | 0.20082 | 0.89664 | 1.9967
O60664 | Perilipin-3 | PLIN3 | 1:10 | Block_5 | -1.238 | 6.3685 | TRUE | 11.1738 | 13.4894 | -0.030748 | 0.82476 | 1.7691 | 1.1843 | 2.0633 | 2.8536
O94903 | Pyridoxal phosphate homeostasis protein | PLPBP | 1:10 | Block_5 | -1.9338 | 7.4223 | TRUE | 9.9437 | 7.5375 | -0.005526 | 1.279 | 3.488 | 1.4768 | 3.7795 | 5.1836
Q9ULL4 | Plexin-B3 | PLXNB3 | 1:10 | Block_5 | -1.3736 | 6.5722 | TRUE | 11.1821 | 8.7079 | -0.30328 | 0.48603 | 1.1219 | -0.027629 | 1.0953 | 1.9788
Q6GTS8 | N-fatty-acyl-amino acid synthase/hydrolase PM20D1 | PM20D1 | 1:10 | Block_5 | -1.8146 | 7.3817 | FALSE | 16.7839 | 13.9171 | NA | -0.055801 | 1.7948 | NA | -0.51356 | 1.6491
P54317 | Pancreatic lipase-related protein 2 | PNLIPRP2 | 1:10 | Block_5 | -6.2509 | 4.0751 | FALSE | 9.3701 | 7.7647 | -1.6771 | -0.59178 | 0.82348 | NA | -0.79981 | 0.73678
P00491 | Purine nucleoside phosphorylase | PNP | 1:10 | Block_5 | -1.9063 | 8.293 | FALSE | 21.9859 | 11.1357 | -1.8983 | 2.535 | 6.1021 | 2.8937 | 5.7539 | 7.892
Q9NZ53 | Podocalyxin-like protein 2 | PODXL2 | 1:10 | Block_5 | -1.7911 | 5.8193 | TRUE | 11.3429 | 6.3329 | -0.46606 | 0.091164 | 0.61513 | -0.22011 | 0.40502 | 1.1754
Q15166 | Serum paraoxonase/lactonase 3 | PON3 | 1:10 | Block_5 | -5 | 3.4514 | FALSE | 6.0986 | 4.2637 | -0.22912 | 0.54651 | 0.98895 | -0.35498 | 0.30632 | 0.79633
Q96CD2 | Phosphopantothenoylcysteine decarboxylase | PPCDC | 1:10 | Block_5 | -2.9484 | 7.3494 | TRUE | 9.0379 | 5.8847 | -0.14185 | 1.097 | 2.1347 | 0.95572 | 2.2082 | 3.1757
P49593 | Protein phosphatase 1F | PPM1F | 1:10 | Block_5 | -0.88337 | 7.8436 | FALSE | 9.4116 | 11.7506 | -0.10523 | 1.17 | 3.2537 | 1.3876 | 3.4516 | 4.34
O14974 | Protein phosphatase 1 regulatory subunit 12A | PPP1R12A | 1:10 | Block_5 | -0.1746 | 8.0785 | TRUE | 8.1817 | 3.5091 | NA | 0.74567 | 4.4681 | 1.4469 | 4.9431 | 6.4435
P41236 | Protein phosphatase inhibitor 2 | PPP1R2 | 1:10 | Block_5 | 0.49311 | 9.2484 | TRUE | NA | NA | NA | NA | 2.6892 | 0.86446 | 2.9847 | 4.417
P01298 | Pancreatic prohormone | PPY | 1:10 | Block_5 | -4.2366 | 7.3498 | TRUE | 10.8729 | 10.4665 | -2.5178 | 0.19403 | 2.0868 | -1.8956 | 0.054231 | 2.4745
P51888 | Prolargin | PRELP | 1:10 | Block_5 | -4 | 4.9684 | TRUE | 12.0303 | 8.2716 | -0.39475 | 0.48031 | 0.84603 | 0.2263 | 0.91076 | 1.7538
P10644 | cAMP-dependent protein kinase type I-alpha regulatory subunit | PRKAR1A | 1:10 | Block_5 | -2.2421 | 8.8529 | FALSE | 11.0337 | 10.2021 | -0.89096 | 0.70005 | 4.4314 | 1.8484 | 4.105 | 5.6399
Q13976 | cGMP-dependent protein kinase 1 | PRKG1 | 1:10 | Block_5 | 0.001 | 12.0118 | TRUE | 10.5077 | 10.5215 | NA | 1.5236 | 5.7965 | 2.1384 | 6.0781 | 8.299
P01236 | Prolactin | PRL | 1:10 | Block_5 | -5.5735 | 4.5769 | FALSE | 10.0886 | 8.5012 | -1.962 | -0.42656 | 1.676 | -0.96718 | 0.20547 | 1.3913
Q16378 | Proline-rich protein 4 | PRR4 | 1:10 | Block_5 | -2.71 | 4.1316 | TRUE | 21.571 | 32.8283 | -2.2635 | -0.39874 | 0.71248 | -1.786 | -0.34018 | 1.0255
Q9BQR3 | Serine protease 27 | PRSS27 | 1:10 | Block_5 | -6.97 | 3.0683 | FALSE | 9.1128 | 8.0875 | -1.2272 | -0.51225 | 0.26371 | -0.69869 | 0.073324 | 0.80577
Q16651 | Prostasin | PRSS8 | 1:10 | Block_5 | -6.25 | 5.2453 | TRUE | 10.5666 | 7.9716 | -0.95552 | 0.20676 | 0.866 | -0.15899 | 0.77024 | 1.7725
Q8TBF2 | Prostamide/prostaglandin F synthase | PRXL2B | 1:10 | Block_5 | 0.10844 | 9.4283 | FALSE | NA | NA | NA | 1.1348 | 4.2043 | 1.1889 | 3.9801 | 5.9179
Q9H939 | Proline-serine-threonine phosphatase-interacting protein 2 | PSTPIP2 | 1:10 | Block_5 | -2.2261 | 4.6611 | FALSE | 9.5975 | 8.5068 | -0.35239 | 2.927 | 6.8333 | 2.6406 | 6.5295 | 8.0539
Q14914 | Prostaglandin reductase 1 | PTGR1 | 1:10 | Block_5 | -2.6384 | 4.7488 | TRUE | 14.2128 | 5.3992 | -1.4402 | -0.21299 | 1.0742 | -0.84351 | 0.3156 | 1.6214
P29350 | Tyrosine-protein phosphatase non-receptor type 6 | PTPN6 | 1:10 | Block_5 | -0.6 | 6.8316 | FALSE | 17.302 | 16.3438 | -0.49774 | 1.7237 | 5.1543 | 1.4626 | 5.2039 | 5.9452
P23467 | Receptor-type tyrosine-protein phosphatase beta | PTPRB | 1:10 | Block_5 | -4.3 | 6.7132 | TRUE | 15.6634 | 6.3665 | -0.24437 | 0.47509 | 0.96873 | -0.03818 | 0.6694 | 1.2136
P23471 | Receptor-type tyrosine-protein phosphatase zeta | PTPRZ1 | 1:10 | Block_5 | -2.3204 | 7.3772 | TRUE | 12.2784 | 4.0003 | -0.24729 | 0.37457 | 1.0979 | 0.063609 | 0.82012 | 1.5999
Q6GMV3 | Putative peptidyl-tRNA hydrolase PTRHD1 | PTRHD1 | 1:10 | Block_5 | 1.02 | 10.1996 | TRUE | 2.1368 | NA | 1.0989 | 2.9943 | 5.4291 | 2.9592 | 5.3327 | 6.9617
P26022 | Pentraxin-related protein PTX3 | PTX3 | 1:10 | Block_5 | -1.2357 | 7.5151 | TRUE | 16.8666 | 8.0032 | -0.82208 | 0.01253 | 1.0564 | -0.26322 | 0.80936 | 2.014
Q08623 | Pseudouridine-5'-phosphatase | PUDP | 1:10 | Block_5 | 0.95762 | 4.433 | TRUE | NA | NA | NA | NA | 1.5351 | NA | 1.2716 | 2.4308
P15151 | Poliovirus receptor | PVR | 1:10 | Block_5 | -3.4889 | 4.4671 | FALSE | 7.376 | 5.6067 | -0.48365 | 0.2401 | 0.85619 | 0.0088892 | 0.74938 | 1.5558
Q9BRP8 | Partner of Y14 and mago | PYM1 | 1:10 | Block_5 | -1.2304 | 11.1739 | TRUE | 12.8005 | 7.4697 | -0.64605 | 1.426 | 6.3029 | 2.0217 | 6.6412 | 8.2569
P09417 | Dihydropteridine reductase | QDPR | 1:10 | Block_5 | -2.7093 | 6.3595 | TRUE | 8.2622 | 3.3815 | -0.43135 | 0.76293 | 1.8732 | 0.61341 | 1.7767 | 2.5347
P61026 | Ras-related protein Rab-10 | RAB10 | 1:10 | Block_5 | -0.5 | 2.1206 | FALSE | 16.355 | 20.3155 | 0.11073 | 1.1174 | 3.1446 | 0.64419 | 2.3425 | 4.1684
O75154 | Rab11 family-interacting protein 3 | RAB11FIP3 | 1:10 | Block_5 | 0.90489 | 11.2786 | TRUE | NA | NA | NA | NA | 4.4676 | 1.0404 | 4.3308 | 5.9354
O00194 | Ras-related protein Rab-27B | RAB27B | 1:10 | Block_5 | -2.5 | 9.2593 | TRUE | 34.9172 | 13.3637 | -1.0971 | 2.2611 | 5.337 | 2.1993 | 4.7581 | 5.66
Q15276 | Rab GTPase-binding effector protein 1 | RABEP1 | 1:10 | Block_5 | -0.028383 | 9.5987 | TRUE | 7.5149 | 4.3679 | NA | 0.75762 | 2.769 | 0.82745 | 2.5991 | 4.2495
Q9BYZ8 | Regenerating islet-derived protein 4 | REG4 | 1:10 | Block_5 | -5.5101 | 4.2464 | TRUE | 12.8621 | 10.0667 | -0.7218 | -0.020079 | 0.76219 | -0.42036 | 0.43025 | 1.7343
Q9P260 | RAB11-binding protein RELCH | RELCH | 1:10 | Block_5 | -0.16639 | 11.2917 | TRUE | 8.3222 | 14.2165 | NA | 1.0516 | 4.6401 | 1.4224 | 4.792 | 6.2688
P00797 | Renin | REN | 1:10 | Block_5 | -3 | 5.8817 | TRUE | 6.5387 | 10.094 | -1.3003 | -0.57719 | 0.57282 | -1.1283 | 0.69397 | 2.9012
Q96B86 | Repulsive guidance molecule A | RGMA | 1:10 | Block_5 | -5 | 4.8587 | TRUE | 11.2264 | 6.2033 | -0.52885 | -0.017139 | 0.71295 | -0.43892 | 0.24169 | 0.95293
Q6NW40 | Repulsive guidance molecule B | RGMB | 1:10 | Block_5 | -2.8884 | 7.4374 | TRUE | 9.5781 | 8.5611 | -0.44865 | 0.10522 | 0.58997 | -0.11209 | 0.53335 | 1.4429
P52758 | 2-iminobutanoate/2-iminopropanoate deaminase | RIDA | 1:10 | Block_5 | -1.0849 | 5.8075 | TRUE | 18.2868 | 11.2893 | -0.76883 | -0.1076 | 0.45499 | -0.35027 | 0.46224 | 1.5681
Q969X0 | RILP-like protein 2 | RILPL2 | 1:10 | Block_5 | -0.40658 | 10.2679 | FALSE | 6.6024 | 12.8126 | -0.40412 | 1.2163 | 4.802 | 1.7382 | 4.9608 | 6.2022
P12724 | Eosinophil cationic protein | RNASE3 | 1:10 | Block_5 | -5.0794 | 7.9692 | TRUE | 18.8124 | 19.3622 | -2.7205 | -0.98039 | 1.8752 | -1.6475 | 0.2097 | 4.0631
Q9Y6N7 | Roundabout homolog 1 | ROBO1 | 1:10 | Block_5 | -3.37 | 6.097 | TRUE | 14.4898 | 6.9355 | -0.46883 | 0.12453 | 0.63407 | -0.20772 | 0.44349 | 1.1471
Q9HCK4 | Roundabout homolog 2 | ROBO2 | 1:10 | Block_5 | -2.6561 | 7.4511 | TRUE | 11.3839 | 6.7963 | -0.37711 | 0.32122 | 0.61232 | -0.0041764 | 0.5706 | 1.1488
Q01973 | Inactive tyrosine-protein kinase transmembrane receptor ROR1 | ROR1 | 1:10 | Block_5 | -2.3416 | 4.1668 | TRUE | 12.8395 | 7.019 | -0.37014 | 0.30187 | 0.79517 | -0.031408 | 0.74314 | 1.478
Q9H446 | RWD domain-containing protein 1 | RWDD1 | 1:10 | Block_5 | -0.85701 | 9.3535 | TRUE | 8.1225 | 5.4269 | -0.011781 | 1.155 | 3.5796 | 1.783 | 3.8212 | 4.8641
P31949 | Protein S100-A11 | S100A11 | 1:10 | Block_5 | -1.161 | 6.5909 | TRUE | 6.8211 | 4.7364 | -0.68391 | -0.33385 | 0.35055 | -0.50221 | 0.14859 | 3.2537
P26447 | Protein S100-A4 | S100A4 | 1:10 | Block_5 | -1.3854 | 4.0515 | TRUE | 14.2518 | 6.1321 | 0.14868 | 1.1304 | 2.3179 | 0.41765 | 2.0416 | 3.2753
Q9BW04 | Specifically androgen-regulated gene protein | SARG | 1:10 | Block_5 | -0.99605 | 11.036 | TRUE | 7.7078 | 17.3823 | -0.47398 | 1.2808 | 4.6984 | 1.5301 | 4.4328 | 6.2051
Q6ZMJ2 | Scavenger receptor class A member 5 | SCARA5 | 1:10 | Block_5 | -3.5736 | 5.6491 | TRUE | 9.3526 | 6.9589 | -0.68061 | 0.19574 | 0.65784 | 0.12823 | 0.73582 | 1.7841
Q96GP6 | Scavenger receptor class F member 2 | SCARF2 | 1:10 | Block_5 | -4.1549 | 6.708 | FALSE | 7.5834 | 8.2052 | -0.36991 | 0.34974 | 0.95238 | 0.1744 | 1.0737 | 2.3908
Q8WXD2 | Secretogranin-3 | SCG3 | 1:10 | Block_5 | -4.88 | 4.9626 | TRUE | 14.2482 | 10.4737 | -0.55291 | 0.33646 | 0.89239 | -0.1926 | 0.66516 | 1.4557
P11684 | Uteroglobin | SCGB1A1 | 1:10 | Block_5 | -2.94 | 2.7582 | FALSE | 12.6024 | 7.7056 | -0.88476 | -0.1239 | 0.88195 | -0.20387 | 1.0006 | 2.2511
Q96QR1 | Secretoglobin family 3A member 1 | SCGB3A1 | 1:10 | Block_5 | -1.1867 | 2.4712 | FALSE | 6.3023 | 5.6703 | -0.62243 | 0.14387 | 0.59433 | -0.60865 | 0.051668 | 0.68218
Q96PL1 | Secretoglobin family 3A member 2 | SCGB3A2 | 1:10 | Block_5 | -7.49 | 4.8978 | FALSE | 12.1637 | 13.7338 | -1.8768 | 0.10701 | 1.3335 | -2.1103 | -0.38619 | 1.519
O75711 | Scrapie-responsive protein 1 | SCRG1 | 1:10 | Block_5 | -2.98 | 4.5638 | TRUE | 13.3822 | 7.6809 | -0.57336 | 0.38215 | 1.0219 | 0.06064 | 0.9159 | 2.0516
P18827 | Syndecan-1 | SDC1 | 1:10 | Block_5 | -3.6282 | 6.6789 | TRUE | 11.5518 | 7.3318 | -0.66315 | -0.14423 | 0.99944 | -0.14834 | 0.64478 | 2.1642
P31431 | Syndecan-4 | SDC4 | 1:10 | Block_5 | -1.3437 | 9.0852 | TRUE | 9.0725 | 8.007 | -0.11174 | 2.2247 | 3.0866 | 0.59481 | 2.9126 | 4.293
Q99470 | Stromal cell-derived factor 2 | SDF2 | 1:10 | Block_5 | 0.99626 | 4.7814 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q58EX2 | Protein sidekick-2 | SDK2 | 1:10 | Block_5 | -2.6128 | 6.7739 | TRUE | 20.1979 | 16.3991 | -0.7932 | 0.18903 | 0.73363 | -0.5667 | 0.1886 | 0.81414
Q13275 | Semaphorin-3F | SEMA3F | 1:10 | Block_5 | -2.4926 | 8.3713 | TRUE | 10.9708 | 4.2981 | -0.42812 | 0.21296 | 0.7429 | -0.14118 | 0.56676 | 1.5898
Q9NS98 | Semaphorin-3G | SEMA3G | 1:10 | Block_5 | -1.8307 | 4.0413 | FALSE | 7.6806 | 8.3651 | -0.70442 | 0.13246 | 1.3543 | -0.19832 | 0.48843 | 1.5153
Q92854 | Semaphorin-4D | SEMA4D | 1:10 | Block_5 | -1.25 | 4.7601 | TRUE | 8.5833 | 7.4812 | -0.24995 | 0.49418 | 0.81306 | 0.14244 | 0.81533 | 1.4798
Q8IW75 | Serpin A12 | SERPINA12 | 1:10 | Block_5 | -3 | 5.5955 | TRUE | 16.8559 | 13.6304 | NA | -2.0758 | -0.95509 | NA | -2.0821 | -0.74599
P35237 | Serpin B6 | SERPINB6 | 1:10 | Block_5 | -0.3319 | 5.0477 | TRUE | 26.9278 | 20.621 | -0.21118 | 0.61934 | 1.8604 | 0.89345 | 2.1047 | 3.0141
P07093 | Glia-derived nexin | SERPINE2 | 1:10 | Block_5 | 0.69543 | 6.4112 | FALSE | NA | 13.6963 | NA | 0.82881 | 2.3258 | NA | 1.8676 | 3.3684
Q99574 | Neuroserpin | SERPINI1 | 1:10 | Block_5 | -1.1423 | 4.1561 | TRUE | 19.3669 | 12.6808 | -1.0973 | -0.42023 | 0.3159 | -0.96634 | -0.30835 | 0.5561
Q9BYH1 | Seizure 6-like protein | SEZ6L | 1:10 | Block_5 | -3.2006 | 6.0982 | TRUE | 14.4564 | 3.5946 | -0.44915 | 0.19526 | 0.67703 | -0.4135 | 0.44158 | 1.2403
Q6FHJ7 | Secreted frizzled-related protein 4 | SFRP4 | 1:10 | Block_5 | -1.418 | 4.4823 | TRUE | 12.1712 | 8.7298 | -0.48813 | 0.23885 | 1.0607 | -0.4892 | 0.43442 | 1.3322
P35247 | Pulmonary surfactant-associated protein D | SFTPD | 1:10 | Block_5 | -1.8285 | 5.4953 | TRUE | 6.4484 | 1.8683 | -0.69062 | 0.1176 | 1.417 | -0.50658 | 0.55483 | 1.8624
P51688 | N-sulphoglucosamine sulphohydrolase | SGSH | 1:10 | Block_5 | -2.5 | 7.0918 | TRUE | 35.0244 | 17.3511 | -1.1463 | 0.15326 | 0.84211 | -0.5813 | 0.26802 | 1.2054
O43765 | Small glutamine-rich tetratricopeptide repeat-containing protein alpha | SGTA | 1:10 | Block_5 | -0.50156 | NA | NA | 8.528 | 8.0851 | 0.84329 | 1.9485 | 4.26 | 2.3187 | 4.0912 | 5.1454
P34896 | Serine hydroxymethyltransferase, cytosolic | SHMT1 | 1:10 | Block_5 | -1.4867 | 7.2456 | FALSE | 14.9986 | 7.1144 | -0.97591 | 1.3611 | 3.2215 | 1.0529 | 3.1015 | 4.8472
P34897 | Serine hydroxymethyltransferase, mitochondrial | SHMT2 | 1:10 | Block_5 | -0.28425 | 7.1889 | FALSE | 5.0763 | 9.8169 | -0.28089 | 0.45884 | 3.801 | 0.6287 | 2.7789 | 4.4119
Q9HAT2 | Sialate O-acetylesterase | SIAE | 1:10 | Block_5 | -1.5026 | 5.0135 | TRUE | 9.8713 | 8.0734 | -0.46702 | 0.7411 | 1.6693 | 0.61882 | 1.733 | 2.8826
Q9BZZ2 | Sialoadhesin | SIGLEC1 | 1:10 | Block_5 | -1.4909 | 5.4049 | TRUE | 9.6488 | 6.6467 | -0.58214 | 0.21384 | 0.67792 | -0.41883 | 0.48507 | 1.5093
Q9Y286 | Sialic acid-binding Ig-like lectin 7 | SIGLEC7 | 1:10 | Block_5 | -3.3427 | 5.763 | TRUE | 8.0553 | 5.168 | -0.52503 | 0.29243 | 0.84813 | 0.064752 | 0.59954 | 1.2051
O00241 | Signal-regulatory protein beta-1 | SIRPB1 | 1:10 | Block_5 | -7.2 | 4.8783 | TRUE | 9.4116 | 4.8892 | -0.41915 | 0.13703 | 0.90313 | -0.08716 | 0.91354 | 2.2015
O75563 | Src kinase-associated phosphoprotein 2 | SKAP2 | 1:10 | Block_5 | -1.2383 | 7.4593 | TRUE | 15.6436 | 14.9885 | -1.1051 | 1.2772 | 5.3471 | 1.6555 | 5.4937 | 6.7173
Q9HBR0 | Putative sodium-coupled neutral amino acid transporter 10 | SLC38A10 | 1:10 | Block_5 | -3.3647 | 5.9533 | TRUE | 11.5803 | 4.9532 | 0.24693 | 0.7442 | 1.4314 | 0.29234 | 1.4325 | 2.6072
Q9H3U7 | SPARC-related modular calcium-binding protein 2 | SMOC2 | 1:10 | Block_5 | -4.2294 | 5.8101 | TRUE | 13.4253 | 5.0862 | -0.62432 | 0.039056 | 0.77786 | -0.17669 | 0.51116 | 1.5971
P17405 | Sphingomyelin phosphodiesterase | SMPD1 | 1:10 | Block_5 | -2.7392 | 6.8883 | FALSE | 10.176 | 8.3769 | -0.90449 | 0.2524 | 0.98585 | -0.70125 | 0.12939 | 0.84166
Q92484 | Acid sphingomyelinase-like phosphodiesterase 3a | SMPDL3A | 1:10 | Block_5 | -3.8585 | 5.7112 | TRUE | 8.1867 | 5.1584 | -2.1116 | -0.23218 | 0.39293 | -1.955 | -0.86198 | 0.38613
O00161 | Synaptosomal-associated protein 23 | SNAP23 | 1:10 | Block_5 | -2.17 | 8.4152 | TRUE | 20.8889 | 14.4581 | -0.63422 | 1.3089 | 5.1385 | 1.2968 | 5.0498 | 6.2201
O95721 | Synaptosomal-associated protein 29 | SNAP29 | 1:10 | Block_5 | -0.7898 | 10.0245 | FALSE | 16.7216 | 11.6356 | -0.12886 | 1.6984 | 5.6957 | 2.4737 | 5.0583 | 6.5105
P37840 | Alpha-synuclein | SNCA | 1:10 | Block_5 | -0.66 | NA | NA | 23.5775 | 17.3407 | -0.069016 | 3.3375 | 6.6823 | 3.2337 | 6.2194 | 7.6799
Q9Y5X1 | Sorting nexin-9 | SNX9 | 1:10 | Block_5 | -0.59967 | 6.307 | FALSE | 6.51 | 6.8459 | 0.017441 | 1.224 | 3.6737 | 1.3502 | 3.4832 | 4.7581
Q99523 | Sortilin | SORT1 | 1:10 | Block_5 | -3.5283 | 5.3843 | TRUE | 11.0287 | 5.9799 | -0.5467 | 0.50017 | 1.0411 | 0.22464 | 0.94801 | 1.6398
Q9BQB4 | Sclerostin | SOST | 1:10 | Block_5 | -4.8456 | 9.3565 | TRUE | 9.3575 | 8.3014 | -0.21851 | 0.66277 | 1.372 | 0.0058415 | 0.92111 | 1.7993
O60271 | C-Jun-amino-terminal kinase-interacting protein 4 | SPAG9 | 1:10 | Block_5 | 0.32814 | 9.0281 | TRUE | NA | NA | NA | 0.6223 | 2.507 | 0.84406 | 2.5761 | 4.055
P09486 | SPARC | SPARC | 1:10 | Block_5 | -4.6 | 5.4853 | FALSE | 18.0864 | 9.712 | -0.47306 | 0.98027 | 2.0276 | 0.84565 | 2.2263 | 3.1347
Q8N0X7 | Spartin | SPART | 1:10 | Block_5 | 0.5 | 10.3769 | FALSE | NA | 0.28746 | NA | 0.8206 | 2.9132 | 0.98897 | 3.0249 | 4.7737
Q6UW49 | Sperm equatorial segment protein 1 | SPESP1 | 1:10 | Block_5 | -0.64544 | 7.3666 | TRUE | 8.7558 | 10.8081 | NA | 0.40502 | 1.2059 | -0.40996 | 0.62884 | 2.1178
P20155 | Serine protease inhibitor Kazal-type 2 | SPINK2 | 1:10 | Block_5 | -4.22 | 7.8847 | TRUE | 6.3756 | 9.0206 | -0.60393 | 0.050538 | 1.1575 | -0.51259 | 0.42489 | 1.738
P58062 | Serine protease inhibitor Kazal-type 7 | SPINK7 | 1:10 | Block_5 | -3.2969 | 5.7491 | TRUE | 10.097 | 10.0189 | -0.95827 | -0.2321 | 0.68738 | -1.0115 | 0.21335 | 1.7483
O43291 | Kunitz-type protease inhibitor 2 | SPINT2 | 1:10 | Block_5 | -3.1484 | 7.2077 | TRUE | 8.4325 | 8.1849 | 0.24513 | 0.86046 | 1.4826 | -0.16425 | 1.8009 | 2.6282
Q9HCB6 | Spondin-1 | SPON1 | 1:10 | Block_5 | -2.9726 | 7.2154 | TRUE | 13.4909 | 3.5117 | -0.46373 | 0.27806 | 0.91579 | 0.19332 | 0.93494 | 2.1995
Q9BUD6 | Spondin-2 | SPON2 | 1:10 | Block_5 | -3.9136 | 3.6823 | TRUE | 9.1842 | 4.5831 | -1.0345 | -0.14139 | 0.5781 | -0.18598 | 0.6382 | 1.5712
P12931 | Proto-oncogene tyrosine-protein kinase Src | SRC | 1:10 | Block_5 | -0.86762 | 8.1239 | TRUE | 8.8844 | 15.7636 | -0.59893 | 0.94504 | 4.9984 | 1.7623 | 4.7815 | 6.4921
P78539 | Sushi repeat-containing protein SRPX | SRPX | 1:10 | Block_5 | -1.376 | 7.2686 | TRUE | 10.786 | 1.4946 | -0.37313 | 0.081033 | 0.68236 | -0.078303 | 0.63753 | 1.9878
Q8WTU2 | Scavenger receptor cysteine-rich domain-containing group B protein | SSC4D | 1:10 | Block_5 | -5.3531 | 5.808 | TRUE | 8.8115 | 5.2848 | -2.4418 | -0.79933 | 1.2704 | -2.3631 | -0.88718 | 1.5759
P50502 | Hsc70-interacting protein | ST13 | 1:10 | Block_5 | -4.3944 | 7.366 | TRUE | 9.7742 | 7.3987 | 0.43223 | 2.146 | 3.6504 | 1.7929 | 3.7214 | 4.5981
P52823 | Stanniocalcin-1 | STC1 | 1:10 | Block_5 | -0.73 | 4.7512 | TRUE | 11.4899 | 9.1432 | -0.34268 | 0.35935 | 0.90345 | 0.065562 | 0.897 | 1.9967
Q13043 | Serine/threonine-protein kinase 4 | STK4 | 1:10 | Block_5 | 0.78517 | 6.4646 | TRUE | NA | NA | NA | 0.9519 | 4.0816 | 0.85331 | 4.1298 | 5.204
O43815 | Striatin | STRN | 1:10 | Block_5 | 0.23633 | 8.0361 | FALSE | NA | NA | NA | NA | 1.2634 | NA | 0.52595 | 3.0656
P50225 | Sulfotransferase 1A1 | SULT1A1 | 1:10 | Block_5 | -0.31199 | 8.4903 | FALSE | 10.3416 | 9.7409 | -0.16288 | 1.6961 | 5.6182 | 1.3393 | 5.7016 | 7.3286
Q8NBJ7 | Inactive C-alpha-formylglycine-generating enzyme 2 | SUMF2 | 1:10 | Block_5 | -2.0462 | 4.7013 | TRUE | 9.8238 | 7.1881 | -0.09665 | 0.98434 | 1.6601 | 0.80725 | 1.6739 | 2.3332
O60279 | Sushi domain-containing protein 5 | SUSD5 | 1:10 | Block_5 | -4.8693 | 3.1975 | TRUE | 9.193 | 4.4597 | -0.34472 | 0.33909 | 0.88062 | 0.035546 | 0.95741 | 1.6052
Q96A49 | Synapse-associated protein 1 | SYAP1 | 1:10 | Block_5 | 1.75 | 6.8294 | TRUE | NA | NA | NA | 1.9357 | 3.9438 | NA | 3.7194 | 5.319
Q96C24 | Synaptotagmin-like protein 4 | SYTL4 | 1:10 | Block_5 | 1.9154 | 5.5331 | FALSE | NA | NA | NA | NA | 5.3917 | NA | 4.2695 | 7.2896
O75347 | Tubulin-specific chaperone A | TBCA | 1:10 | Block_5 | -0.1591 | 8.7939 | TRUE | 6.9966 | 11.0449 | 0.10126 | 1.847 | 5.0044 | 2.6332 | 5.586 | 7.0211
Q13428 | Treacle protein | TCOF1 | 1:10 | Block_5 | -0.96578 | 8.5932 | TRUE | 11.2499 | 6.4803 | -0.61782 | -0.0074258 | 0.62958 | -0.029881 | 0.53243 | 2.0579
Q6NUS6 | Tectonic-3 | TCTN3 | 1:10 | Block_5 | -2.9814 | 4.58 | FALSE | 10.7084 | 3.9886 | -0.4848 | 0.13097 | 0.54641 | -0.43877 | 0.14226 | 0.84682
Q02763 | Angiopoietin-1 receptor | TEK | 1:10 | Block_5 | -1.2768 | 3.7921 | TRUE | 6.4554 | 8.4119 | -0.41752 | 0.22093 | 0.65756 | -0.27502 | 0.33224 | 1.048
P48307 | Tissue factor pathway inhibitor 2 | TFPI2 | 1:10 | Block_5 | -4.0196 | NA | NA | 8.8415 | 8.1289 | -0.46124 | -0.069195 | 0.51233 | -0.086022 | 0.67212 | 1.8621
P37173 | TGF-beta receptor type-2 | TGFBR2 | 1:10 | Block_5 | -3.2684 | 5.0883 | TRUE | 9.4466 | 5.3377 | -0.52229 | 0.052718 | 0.51756 | 0.021235 | 0.89388 | 1.8601
P21980 | Protein-glutamine gamma-glutamyltransferase 2 | TGM2 | 1:10 | Block_5 | -2.0782 | 9.3948 | FALSE | 14.4012 | 5.7559 | 0.78141 | 2.6753 | 4.5937 | 1.3779 | 2.7034 | 3.9609
P07204 | Thrombomodulin | THBD | 1:10 | Block_5 | -7.36 | 3.587 | FALSE | 9.568 | 8.4211 | -0.52617 | 0.28219 | 0.81411 | 0.044884 | 0.70202 | 1.6884
P52888 | Thimet oligopeptidase | THOP1 | 1:10 | Block_5 | -1.0898 | 5.8856 | TRUE | 13.4463 | 6.3055 | -0.62362 | 0.20786 | 0.95528 | -0.55023 | 0.64531 | 1.4314
P04216 | Thy-1 membrane glycoprotein | THY1 | 1:10 | Block_5 | -5.1135 | 4.6536 | TRUE | 9.8663 | 5.4926 | -0.39056 | 0.14777 | 0.48958 | -0.0048685 | 0.61446 | 1.7002
P62072 | Mitochondrial import inner membrane translocase subunit Tim10 | TIMM10 | 1:10 | Block_5 | 1.7647 | 9.3349 | TRUE | NA | NA | NA | NA | 2.4609 | NA | NA | 2.6695
P16035 | Metalloproteinase inhibitor 2 | TIMP2 | 1:10 | Block_5 | -1.1844 | 6.2558 | FALSE | 7.8354 | 8.8959 | -0.56275 | 0.35275 | 0.84749 | -0.12796 | 0.62683 | 1.2041
P35625 | Metalloproteinase inhibitor 3 | TIMP3 | 1:10 | Block_5 | -1.0239 | 7.7858 | FALSE | 16.1972 | 11.7586 | 0.41635 | 1.7526 | 4.0296 | 1.4347 | 3.7255 | 6.2525
Q9GZM7 | Tubulointerstitial nephritis antigen-like | TINAGL1 | 1:10 | Block_5 | -2.3536 | 7.6934 | TRUE | 8.7212 | 3.0458 | -0.19549 | 0.35994 | 0.97121 | 0.16084 | 0.64005 | 1.2012
Q5JTD0 | Tight junction-associated protein 1 | TJAP1 | 1:10 | Block_5 | 0.095279 | 9.9536 | FALSE | NA | NA | NA | 0.44745 | 2.595 | 0.58624 | 2.3684 | 4.3082
Q6P589 | Tumor necrosis factor alpha-induced protein 8-like protein 2 | TNFAIP8L2 | 1:10 | Block_5 | 0.55149 | 7.2091 | FALSE | NA | NA | NA | 0.71627 | 3.0974 | NA | 3.7161 | 5.5723
O00300 | Tumor necrosis factor receptor superfamily member 11B | TNFRSF11B | 1:10 | Block_5 | -3.4102 | 7.0943 | TRUE | 10.815 | 5.9405 | -0.77842 | 0.066938 | 0.78081 | -0.080742 | 0.49371 | 1.3637
Q9NP84 | Tumor necrosis factor receptor superfamily member 12A | TNFRSF12A | 1:10 | Block_5 | -3.047 | 6.2094 | TRUE | 11.4643 | 12.0248 | -0.77111 | -0.17814 | 0.65491 | -0.26624 | 0.81983 | 1.7524
O14836 | Tumor necrosis factor receptor superfamily member 13B | TNFRSF13B | 1:10 | Block_5 | -2.0031 | 7.6856 | TRUE | 12.9156 | 10.4416 | -0.39912 | 0.17525 | 0.75397 | -0.38725 | 0.21854 | 1.2651
Q92956 | Tumor necrosis factor receptor superfamily member 14 | TNFRSF14 | 1:10 | Block_5 | -4.9571 | 5.3281 | TRUE | 10.3332 | 9.1343 | -0.37367 | 0.32612 | 0.67351 | 0.77417 | 1.3011 | 2.2409
O75509 | Tumor necrosis factor receptor superfamily member 21 | TNFRSF21 | 1:10 | Block_5 | -6.4829 | 5.39 | TRUE | 8.749 | 7.5589 | -0.56034 | 0.26938 | 0.63895 | 0.04446 | 0.69716 | 1.5861
P28908 | Tumor necrosis factor receptor superfamily member 8 | TNFRSF8 | 1:10 | Block_5 | -4.5943 | 7.2892 | TRUE | 9.6395 | 5.9085 | -0.81307 | -0.010151 | 0.9107 | -0.14694 | 0.76657 | 2.3644
O75888 | Tumor necrosis factor ligand superfamily member 13 | TNFSF13 | 1:10 | Block_5 | -1.8066 | 7.5807 | FALSE | 11.4318 | 9.6078 | 0.41872 | 0.8675 | 1.2778 | 0.81341 | 1.3972 | 2.4217
Q15388 | Mitochondrial import receptor subunit TOM20 homolog | TOMM20 | 1:10 | Block_5 | 0.83 | 9.8875 | TRUE | NA | 9.6549 | NA | 1.1183 | 3.4203 | 0.94453 | 2.339 | 3.902
Q53FA7 | Quinone oxidoreductase PIG3 | TP53I3 | 1:10 | Block_5 | -4.43 | 6.7238 | TRUE | 9.1024 | 9.3036 | -0.76461 | 0.57898 | 1.8432 | 0.68506 | 1.9331 | 3.0914
P55327 | Tumor protein D52 | TPD52 | 1:10 | Block_5 | -2.1139 | 6.9934 | TRUE | 5.1199 | 4.2533 | -0.061958 | 1.1472 | 3.1227 | 0.82333 | 3.2257 | 5.1503
O43399 | Tumor protein D54 | TPD52L2 | 1:10 | Block_5 | 0.21492 | 6.5969 | TRUE | 13.7627 | 10.5813 | NA | 1.5814 | 5.4047 | 1.8797 | 5.5124 | 6.6525
Q9H3S4 | Thiamin pyrophosphokinase 1 | TPK1 | 1:10 | Block_5 | -3.9585 | 6.7963 | TRUE | 16.3382 | 11.2078 | -0.92046 | -0.050713 | 0.71536 | -0.184 | 0.57036 | 1.255
O14773 | Tripeptidyl-peptidase 1 | TPP1 | 1:10 | Block_5 | -6.67 | 3.1434 | TRUE | 7.4593 | 5.3834 | -1.1499 | -0.23484 | 0.32861 | -0.45432 | -0.0041788 | 0.71492
Q9NZC2 | Triggering receptor expressed on myeloid cells 2 | TREM2 | 1:10 | Block_5 | -4.0079 | 3.7273 | FALSE | 12.0925 | 7.9347 | -0.27833 | 0.42086 | 1.3105 | -0.36193 | 0.92822 | 2.1867
Q7L0Y3 | tRNA methyltransferase 10 homolog C | TRMT10C | 1:10 | Block_5 | 1.0169 | 11.6853 | TRUE | NA | NA | NA | 1.1251 | 3.5349 | 1.2306 | 2.8301 | 4.3064
P01222 | Thyrotropin subunit beta | TSHB | 1:10 | Block_5 | -2.4093 | 6.622 | TRUE | 8.6062 | 7.0693 | -0.8335 | 0.28175 | 1.0799 | -0.18729 | 0.84244 | 2.1142
Q96J42 | Thioredoxin domain-containing protein 15 | TXNDC15 | 1:10 | Block_5 | -1.5588 | 6.3948 | TRUE | 10.0796 | 8.9603 | -0.36985 | 0.26802 | 0.67091 | -0.27754 | 0.41519 | 1.188
Q8NBS9 | Thioredoxin domain-containing protein 5 | TXNDC5 | 1:10 | Block_5 | -0.16037 | 1.8027 | FALSE | 6.034 | 8.2297 | NA | 0.54328 | 1.7649 | 0.91004 | 2.3348 | 3.4648
P19971 | Thymidine phosphorylase | TYMP | 1:10 | Block_5 | -3.08 | 2.1571 | FALSE | 11.8258 | 11.1066 | -0.0045353 | 1.4569 | 3.3005 | 1.8534 | 3.4337 | 4.3572
Q06418 | Tyrosine-protein kinase receptor TYRO3 | TYRO3 | 1:10 | Block_5 | -4.5283 | 4.5251 | TRUE | 8.8164 | 6.7191 | -0.48193 | 0.1855 | 0.578 | -0.42594 | 0.37901 | 1.0881
Q9BZF9 | Uveal autoantigen with coiled-coil domains and ankyrin repeats | UACA | 1:10 | Block_5 | -0.2529 | 11.6619 | TRUE | 35.0227 | 16.0566 | NA | 0.93626 | 2.8246 | 0.52515 | 2.3041 | 5.599
O14933 | Ubiquitin/ISG15-conjugating enzyme E2 L6 | UBE2L6 | 1:10 | Block_5 | -1.5063 | 8.0328 | FALSE | 9.5174 | 10.8005 | -0.14325 | 0.80375 | 2.4234 | 1.4013 | 2.8931 | 4.0498
Q04323 | UBX domain-containing protein 1 | UBXN1 | 1:10 | Block_5 | 0.88 | 5.7051 | TRUE | NA | 9.1992 | NA | 1.1451 | 2.4082 | NA | 2.1659 | 3.1075
Q9UBR1 | Beta-ureidopropionase | UPB1 | 1:10 | Block_5 | 1.9592 | 5.677 | TRUE | NA | NA | NA | NA | 2.6149 | NA | NA | NA
Q9BZI7 | Regulator of nonsense transcripts 3B | UPF3B | 1:10 | Block_5 | -0.097044 | 9.9771 | TRUE | NA | 3.1896 | NA | 0.18105 | 1.0755 | 0.0061039 | 0.94558 | 2.6364
P06132 | Uroporphyrinogen decarboxylase | UROD | 1:10 | Block_5 | -0.4196 | 7.9854 | TRUE | 6.9267 | 6.8303 | 0.35997 | 1.7804 | 4.0155 | 1.833 | 3.7805 | 4.9243
P13611 | Versican core protein | VCAN | 1:10 | Block_5 | -2.129 | 8.6918 | TRUE | 13.5438 | 6.4195 | -0.29453 | 0.068207 | 0.50146 | -0.051109 | 0.65334 | 1.5106
P49767 | Vascular endothelial growth factor C | VEGFC | 1:10 | Block_5 | -1.0623 | 6.5341 | TRUE | 12.1229 | 11.0316 | -0.21699 | 0.86491 | 1.7028 | 0.76396 | 2.0907 | 3.1271
Q6UXI7 | Vitrin | VIT | 1:10 | Block_5 | -3.3135 | 6.7748 | TRUE | 10.7887 | 5.1502 | -1.2396 | -0.198 | 0.48422 | -0.43399 | 0.24713 | 1.1664
Q9H7M9 | V-type immunoglobulin domain-containing suppressor of T-cell activation | VSIR | 1:10 | Block_5 | -0.49868 | 7.2369 | TRUE | 13.8178 | 7.7761 | -0.38597 | 0.70764 | 2.1968 | 0.85416 | 1.6127 | 2.8714
Q9NP79 | Vacuolar protein sorting-associated protein VTA1 homolog | VTA1 | 1:10 | Block_5 | -2.3905 | 6.9886 | FALSE | 8.3369 | 5.4471 | 0.08283 | 1.9668 | 4.7194 | 2.2353 | 4.9708 | 6.2126
Q96AJ9 | Vesicle transport through interaction with t-SNAREs homolog 1A | VTI1A | 1:10 | Block_5 | 0.34473 | 6.4847 | TRUE | NA | NA | NA | NA | 1.1056 | NA | 0.82569 | 2.1901
Q14508 | WAP four-disulfide core domain protein 2 | WFDC2 | 1:10 | Block_5 | -3.45 | 5.0691 | TRUE | 10.9732 | 6.4721 | -0.78256 | 0.15965 | 0.85058 | -0.20538 | 0.8942 | 2.3174
Q8TEU8 | WAP, Kazal, immunoglobulin, Kunitz and NTR domain-containing protein 2 | WFIKKN2 | 1:10 | Block_5 | -6.953 | 4.4763 | TRUE | 9.9812 | 9.2761 | -0.50025 | -0.029526 | 0.69411 | -0.34307 | 0.35178 | 1.086
O43895 | Xaa-Pro aminopeptidase 2 | XPNPEP2 | 1:10 | Block_5 | -4.5711 | 3.9157 | TRUE | 7.7776 | 6.0011 | -2.0692 | -0.37846 | 1.3437 | -2.1277 | -0.58732 | 0.88649
P27348 | 14-3-3 protein theta (14-3-3 protein T-cell) (14-3-3 protein tau) (Protein HS1) | YWHAQ | 1:10 | Block_5 | 0.23328 | 7.4763 | TRUE | 5.146 | 3.1984 | NA | 0.6172 | 2.841 | 0.87535 | 2.8577 | 4.4167
Q9H7C9 | Mth938 domain-containing protein | AAMDC | 1:1 | Block_4 | 0.18913 | 5.7158 | FALSE | 2.1575 | 3.874 | NA | 0.28654 | 1.4063 | 0.28291 | 1.3552 | 2.4131
Q9BTE6 | Alanyl-tRNA editing protein Aarsd1 | AARSD1 | 1:1 | Block_4 | -0.14 | 9.3918 | TRUE | NA | 3.6042 | 0.23948 | 1.4675 | 2.3456 | 1.5542 | 2.8379 | 3.8868
Q96IU4 | Putative protein-lysine deacylase ABHD14B | ABHD14B | 1:1 | Block_4 | -3.4776 | 5.2198 | TRUE | 14.4884 | 7.6027 | -1.0631 | 0.014526 | 2.1143 | 0.96257 | 2.9704 | 4.5481
Q7Z7G0 | Target of Nesh-SH3 | ABI3BP | 1:1 | Block_4 | -2.1539 | 4.9512 | TRUE | 9.088 | 3.942 | -0.29194 | 1.0178 | 1.7287 | 0.64274 | 1.5643 | 2.3686
P00519 | Tyrosine-protein kinase ABL1 | ABL1 | 1:1 | Block_4 | -2.3504 | 10.4332 | TRUE | 8.3577 | 6.6249 | -0.77246 | 0.3859 | 2.4328 | 1.4008 | 2.8921 | 4.3691
O14639 | Actin-binding LIM protein 1 | ABLIM1 | 1:1 | Block_4 | -0.091023 | 9.5347 | FALSE | 14.2474 | 8.7768 | -0.032117 | 1.1048 | 4.9534 | 1.6183 | 4.0802 | 6.0685
Q9BYF1 | Angiotensin-converting enzyme 2 | ACE2 | 1:1 | Block_4 | -0.89159 | 5.9035 | TRUE | 9.2719 | 8.9491 | NA | -0.7225 | 0.29214 | -0.81648 | -0.033861 | 1.3065
Q9NPJ3 | Acyl-coenzyme A thioesterase 13 | ACOT13 | 1:1 | Block_4 | 0.8 | 11.1853 | FALSE | NA | 10.8009 | NA | 3.807 | 7.1057 | 3.055 | 6.613 | 8.8323
P24666 | Low molecular weight phosphotyrosine protein phosphatase | ACP1 | 1:1 | Block_4 | 0.22 | 6.8457 | FALSE | NA | NA | NA | 0.81133 | 2.5745 | 0.53895 | 2.1394 | 3.6418
Q9NPH0 | Lysophosphatidic acid phosphatase type 6 | ACP6 | 1:1 | Block_4 | -2.9204 | 4.9053 | TRUE | 11.4364 | 10.343 | -0.83197 | 0.25144 | 0.89132 | -0.10606 | 0.97493 | 1.8627
Q8NEB7 | Acrosin-binding protein | ACRBP | 1:1 | Block_4 | 0.48029 | 8.7581 | TRUE | NA | 23.4225 | NA | NA | 1.0227 | NA | 1.232 | 2.1516
P26436 | Acrosomal protein SP-10 | ACRV1 | 1:1 | Block_4 | 0.64997 | 8.7972 | TRUE | 5.1768 | 5.3743 | NA | NA | NA | NA | NA | 2.253
P62736 | Actin, aortic smooth muscle | ACTA2 | 1:1 | Block_4 | -1.6081 | 2.4215 | FALSE | 5.7376 | 6.6808 | -0.35182 | 0.32004 | 0.74903 | 0.15205 | 1.1046 | 2.3343
P35609 | Alpha-actinin-2 | ACTN2 | 1:1 | Block_4 | -1.1362 | 5.6867 | TRUE | 9.4719 | 4.3416 | -1.1104 | -0.33946 | 1.8895 | -0.71161 | 1.0078 | 3.0081
P37023 | Serine/threonine-protein kinase receptor R3 | ACVRL1 | 1:1 | Block_4 | -6.1217 | 7.3332 | TRUE | 9.209 | 3.1817 | -0.39304 | 0.26547 | 0.75791 | 0.081067 | 0.81932 | 1.9565
P00813 | Adenosine deaminase | ADA | 1:1 | Block_4 | -4.97 | -0.9224 | FALSE | 6.5877 | 4.4846 | -1.1792 | -0.06923 | 0.52695 | -0.30532 | 0.66778 | 1.6841
Q13444 | Disintegrin and metalloproteinase domain-containing protein 15 | ADAM15 | 1:1 | Block_4 | -4.56 | 4.959 | TRUE | 9.77 | 13.2051 | -0.39442 | 0.25521 | 0.82906 | -0.12393 | 0.56706 | 1.3216
Q9P0K1 | Disintegrin and metalloproteinase domain-containing protein 22 | ADAM22 | 1:1 | Block_4 | -4.24 | 4.6195 | FALSE | 14.7155 | 18.9547 | -0.70377 | 0.15905 | 0.86319 | 0.0059581 | 0.65598 | 1.614
O75077 | Disintegrin and metalloproteinase domain-containing protein 23 | ADAM23 | 1:1 | Block_4 | -5.19 | 3.551 | TRUE | 6.2847 | 6.0272 | -0.2859 | 0.36948 | 0.931 | 0.067277 | 0.8524 | 1.4568
Q13443 | Disintegrin and metalloproteinase domain-containing protein 9 | ADAM9 | 1:1 | Block_4 | -1.52 | 2.8791 | TRUE | 10.3593 | 9.7427 | -0.77273 | 0.043458 | 0.6682 | -0.19937 | 0.49205 | 1.2911
Q9UHI8 | A disintegrin and metalloproteinase with thrombospondin motifs 1 | ADAMTS1 | 1:1 | Block_4 | 0.39517 | 9.9481 | TRUE | 7.402 | 9.9519 | NA | NA | 0.83549 | NA | 0.84877 | 1.5859
Q8TE58 | A disintegrin and metalloproteinase with thrombospondin motifs 15 | ADAMTS15 | 1:1 | Block_4 | -2.3842 | 10.373 | TRUE | 5.5911 | 4.1737 | 0.69873 | 1.8787 | 2.3135 | 0.40542 | 2.1184 | 3.0705
Q6UY14 | ADAMTS-like protein 4 | ADAMTSL4 | 1:1 | Block_4 | -1.91 | 4.8673 | TRUE | 10.3472 | 7.011 | 0.49419 | 1.0129 | 1.6437 | 0.36268 | 1.124 | 1.6945
O14514 | Adhesion G protein-coupled receptor B1 | ADGRB1 | 1:1 | Block_4 | -2.5243 | 9.6935 | TRUE | 12.8913 | 4.4161 | -0.45359 | 0.064597 | 0.77098 | -0.58367 | 0.33473 | 1.2901
O60242 | Adhesion G protein-coupled receptor B3 | ADGRB3 | 1:1 | Block_4 | -5.453 | 5.6064 | TRUE | 7.5049 | 5.8522 | -0.14474 | 0.47859 | 0.97326 | 0.12304 | 0.76361 | 1.634
Q14246 | Adhesion G protein-coupled receptor E1 | ADGRE1 | 1:1 | Block_4 | -4.0338 | 3.6145 | FALSE | 19.6592 | 11.9872 | -2.9324 | -1.9336 | -1.1251 | -2.4254 | -1.3581 | -0.17642
Q9BY15 | Adhesion G protein-coupled receptor E3 | ADGRE3 | 1:1 | Block_4 | -1.7232 | 6.1669 | TRUE | 14.5807 | 12.3645 | -0.56252 | -0.0047025 | 0.69673 | -0.55881 | 0.24726 | 1.4895
Q9Y653 | Adhesion G-protein coupled receptor G1 | ADGRG1 | 1:1 | Block_4 | 0.026935 | 4.2133 | TRUE | 6.0361 | 6.4398 | NA | 0.13506 | 0.47267 | NA | 0.56037 | 2.3515
P35318 | Pro-adrenomedullin | ADM | 1:1 | Block_4 | -4.3497 | 4.5199 | FALSE | 6.0535 | 4.2305 | 0.71794 | 2.4048 | 2.9563 | 1.5468 | 2.8675 | 3.6235
P30566 | Adenylosuccinate lyase | ADSL | 1:1 | Block_4 | -1.8931 | 4.9095 | FALSE | 9.3166 | 7.2129 | 0.42872 | 2.686 | 3.3699 | 1.8215 | 3.1346 | 4.2765
Q8TED9 | Actin filament-associated protein 1-like 1 | AFAP1L1 | 1:1 | Block_4 | -1.2309 | 11.2303 | TRUE | 10.7196 | 5.225 | -0.30776 | 0.11431 | 0.50507 | 0.40868 | 1.1866 | 2.5383
Q5U5Z8 | Cytosolic carboxypeptidase 2 | AGBL2 | 1:1 | Block_4 | 1.06 | 9.5271 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O95994 | Anterior gradient protein 2 homolog | AGR2 | 1:1 | Block_4 | -0.96905 | 7.0737 | FALSE | 13.4035 | 6.237 | NA | -0.19418 | 1.1911 | -0.29211 | 1.044 | 2.9051
O00253 | Agouti-related protein | AGRP | 1:1 | Block_4 | -4.74 | 5.7185 | TRUE | 12.4936 | 11.7636 | -0.37121 | 0.59697 | 1.2699 | 0.11517 | 1.2653 | 2.5676
Q09666 | Neuroblast differentiation-associated protein AHNAK | AHNAK | 1:1 | Block_4 | -2.7405 | NA | NA | 5.0299 | 8.6181 | -0.52492 | 0.20631 | 0.63199 | -0.045538 | 0.91081 | 1.8942
Q8IVF2 | Protein AHNAK2 | AHNAK2 | 1:1 | Block_4 | -1.5996 | 7.6497 | TRUE | 5.8618 | 3.9526 | -0.00042436 | 0.51444 | 1.1227 | 0.39872 | 1.108 | 1.9485
O95433 | Activator of 90 kDa heat shock protein ATPase homolog 1 | AHSA1 | 1:1 | Block_4 | 0.56 | 6.9734 | FALSE | 9.9939 | 6.3497 | NA | NA | 2.3238 | NA | 0.84774 | 1.7725
Q9BQI0 | Allograft inflammatory factor 1-like | AIF1L | 1:1 | Block_4 | -0.25349 | 7.2277 | TRUE | NA | NA | NA | NA | 0.65063 | NA | NA | 0.6107
O95831 | Apoptosis-inducing factor 1, mitochondrial | AIFM1 | 1:1 | Block_4 | -0.057363 | 8.2165 | TRUE | 4.5506 | 4.7533 | NA | 0.45469 | 3.6485 | 0.28634 | 2.6752 | 4.9433
P54819 | Adenylate kinase 2, mitochondrial | AK2 | 1:1 | Block_4 | 2.3 | 10.1059 | FALSE | NA | NA | NA | 2.4097 | 5.2688 | NA | 5.4464 | 7.1624
Q9Y2D5 | A-kinase anchor protein 2 | AKAP2 | 1:1 | Block_4 | -0.95952 | 8.4146 | TRUE | 10.0608 | 5.9726 | -0.14223 | 0.40004 | 0.81803 | 0.59511 | 1.5354 | 2.6882
Q96B36 | Proline-rich AKT1 substrate 1 | AKT1S1 | 1:1 | Block_4 | -1.1887 | 6.9194 | FALSE | 10.1309 | 8.9314 | 0.29476 | 1.9249 | 3.5948 | 1.6846 | 3.5716 | 4.8923
P31751 | RAC-beta serine/threonine-protein kinase | AKT2 | 1:1 | Block_4 | 1.7 | 13.0111 | FALSE | NA | NA | NA | 2.6254 | 7.1095 | NA | 6.4233 | 9.5831
P05091 | Aldehyde dehydrogenase, mitochondrial | ALDH2 | 1:1 | Block_4 | -0.18254 | 6.436 | TRUE | 8.6258 | 12.0419 | NA | NA | 1.4326 | NA | 0.16786 | 2.3507
P30838 | Aldehyde dehydrogenase, dimeric NADP-preferring | ALDH3A1 | 1:1 | Block_4 | 0.64094 | 7.9176 | TRUE | NA | NA | NA | NA | 0.83651 | NA | NA | 1.0852
P51649 | Succinate-semialdehyde dehydrogenase, mitochondrial | ALDH5A1 | 1:1 | Block_4 | -0.02202 | 5.4827 | TRUE | 1.0692 | 0.25119 | NA | NA | 0.96111 | NA | 0.93705 | 2.4518
P05187 | Alkaline phosphatase, placental type | ALPP | 1:1 | Block_4 | -3.9861 | 4.1836 | TRUE | 6.2677 | 5.1872 | -2.6539 | -0.52778 | 1.5823 | -3.5345 | -2.368 | -0.71024
P02760 | Protein AMBP | AMBP | 1:1 | Block_4 | -3.62 | -2.7502 | FALSE | 3.923 | 2.2084 | -0.21936 | 0.095351 | 0.34506 | -0.10808 | 0.27818 | 0.62603
Q96NU7 | Probable imidazolonepropionase | AMDHD1 | 1:1 | Block_4 | -4.4733 | 3.8182 | FALSE | 10.5376 | 9.6972 | -2.2362 | -0.04492 | 1.5394 | -1.0438 | 0.46621 | 2.0967
Q86SJ2 | Amphoterin-induced protein 2 | AMIGO2 | 1:1 | Block_4 | -4.83 | 2.2245 | TRUE | 10.3711 | 2.9117 | -0.18573 | 0.26912 | 0.60208 | 0.077762 | 0.4557 | 0.75202
Q01432 | AMP deaminase 3 | AMPD3 | 1:1 | Block_4 | 0.17166 | 1.5028 | FALSE | NA | 4.3842 | NA | 1.1087 | 2.4807 | NA | 1.6536 | 3.9838
O15123 | Angiopoietin-2 | ANGPT2 | 1:1 | Block_4 | -2.8578 | 4.9629 | FALSE | 10.466 | 5.5512 | -0.33166 | 0.46468 | 0.92832 | -0.093466 | 1.0001 | 1.8845
O43827 | Angiopoietin-related protein 7 | ANGPTL7 | 1:1 | Block_4 | 1.5 | NA | NA | NA | NA | NA | NA | NA | NA | NA | 2.1225
Q8IV38 | Ankyrin repeat and MYND domain-containing protein 2 | ANKMY2 | 1:1 | Block_4 | 0.54112 | 8.3819 | TRUE | NA | NA | NA | 0.73422 | 2.6981 | NA | 2.7243 | 3.9135
Q9H8Y5 | Ankyrin repeat and zinc finger domain-containing protein 1 | ANKZF1 | 1:1 | Block_4 | -1.48 | 11.2911 | FALSE | 12.8897 | 5.7416 | -1.4688 | -0.011486 | 2.1175 | 0.61448 | 2.3434 | 4.208
O43423 | Acidic leucine-rich nuclear phosphoprotein 32 family member C | ANP32C | 1:1 | Block_4 | 1.073 | 9.4413 | TRUE | 14.2762 | 27.9334 | NA | NA | NA | NA | NA | NA
Q9UJ72 | Annexin A10 | ANXA10 | 1:1 | Block_4 | -1.3525 | 8.1706 | TRUE | 12.5513 | 8.7371 | -1.2641 | 0.016578 | 1.3926 | -0.91327 | 0.41879 | 2.4436
P50995 | Annexin A11 | ANXA11 | 1:1 | Block_4 | -3.2802 | 1.7344 | FALSE | 8.0536 | 4.6762 | -0.19067 | 0.54662 | 1.2196 | 0.55673 | 1.3397 | 2.0436
P07355 | Annexin A2 | ANXA2 | 1:1 | Block_4 | -1.11 | 2.833 | FALSE | 24.8473 | 18.6959 | -0.89947 | 0.044334 | 0.69683 | -0.21302 | 0.79561 | 3.2118
O00203 | AP-3 complex subunit beta-1 | AP3B1 | 1:1 | Block_4 | 0.04 | 10.1007 | FALSE | 17.5976 | 18.0887 | NA | 0.67267 | 2.8928 | 1.1652 | 3.3536 | 5.1124
P59780 | AP-3 complex subunit sigma-2 | AP3S2 | 1:1 | Block_4 | 1.79 | 5.4248 | FALSE | NA | NA | NA | NA | NA | NA | 2.3075 | 3.7602
Q7Z5R6 | Amyloid beta A4 precursor protein-binding family B member 1-interacting protein | APBB1IP | 1:1 | Block_4 | -0.32735 | 6.7243 | FALSE | 16.8769 | 11.3482 | NA | 0.055213 | 0.82106 | 0.35348 | 1.0905 | 3.9566
P27695 | DNA-(apurinic or apyrimidinic site) endonuclease | APEX1 | 1:1 | Block_4 | -5.4435 | 4.7976 | TRUE | 7.6243 | 5.0741 | -3.3566 | -1.6608 | -0.15194 | -1.4229 | 0.34547 | 3.075
Q96GX9 | Methylthioribulose-1-phosphate dehydratase | APIP | 1:1 | Block_4 | -3.6878 | 5.0775 | FALSE | 6.6208 | 8.5205 | -1.0556 | -0.040466 | 0.96761 | -0.090038 | 1.2124 | 2.2692
P07741 | Adenine phosphoribosyltransferase | APRT | 1:1 | Block_4 | -0.49486 | 4.6036 | FALSE | 9.1299 | 10.08 | 0.54116 | 2.2536 | 4.1784 | 1.8064 | 3.9615 | 4.6838
P15514 | Amphiregulin | AREG | 1:1 | Block_4 | -2.21 | 7.6038 | TRUE | 18.7793 | 4.035 | -1.0279 | -0.45209 | 0.48787 | -0.60168 | 0.46676 | 2.5985
P62330 | ADP-ribosylation factor 6 | ARF6 | 1:1 | Block_4 | -0.66661 | 9.3903 | FALSE | 12.2635 | 9.1098 | NA | 2.3791 | 6.3186 | 1.9517 | 6.6514 | 8.5305
P53367 | Arfaptin-1 | ARFIP1 | 1:1 | Block_4 | -0.52 | 8.2303 | FALSE | 13.8774 | 14.1838 | -0.43919 | 1.0973 | 4.5496 | 1.528 | 4.4515 | 5.6919
P05089 | Arginase-1 | ARG1 | 1:1 | Block_4 | -0.76 | 5.4288 | TRUE | 14.1929 | 5.8542 | -0.43404 | 0.55733 | 1.7889 | -0.18704 | 0.74968 | 3.809
Q6ZUM4 | Rho GTPase-activating protein 27 | ARHGAP27 | 1:1 | Block_4 | -1.51 | 7.4667 | TRUE | 7.7599 | 4.0217 | -1.1681 | -0.31083 | 0.20937 | -0.14762 | 0.67583 | 4.2659
Q7Z6I6 | Rho GTPase-activating protein 30 | ARHGAP30 | 1:1 | Block_4 | 0.4594 | 10.8955 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.3488
Q92888 | Rho guanine nucleotide exchange factor 1 | ARHGEF1 | 1:1 | Block_4 | -2.52 | 2.5309 | FALSE | 14.172 | 19.2695 | -0.73489 | 1.9042 | 4.7979 | 1.3371 | 5.1265 | 6.6533
O15013 | Rho guanine nucleotide exchange factor 10 | ARHGEF10 | 1:1 | Block_4 | 1.5034 | 9.003 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.7872
Q9NZN5 | Rho guanine nucleotide exchange factor 12 | ARHGEF12 | 1:1 | Block_4 | 1.5791 | 6.4432 | FALSE | NA | NA | NA | NA | 3.307 | NA | 2.7762 | 4.9456
P15848 | Arylsulfatase B | ARSB | 1:1 | Block_4 | -0.26961 | 5.6526 | TRUE | 9.715 | 11.0866 | NA | 0.68092 | 1.5902 | 0.2339 | 1.8618 | 3.295
Q9NR71 | Neutral ceramidase | ASAH2 | 1:1 | Block_4 | -4.4165 | 2.4145 | TRUE | 7.5925 | 7.417 | -0.6978 | 0.070073 | 0.71935 | -1.0344 | -0.10192 | 0.69998
Q8N9N2 | Activating signal cointegrator 1 complex subunit 1 | ASCC1 | 1:1 | Block_4 | -1.91 | 8.9552 | FALSE | 8.8271 | 6.0156 | -1.6823 | 0.53156 | 3.9412 | 1.2426 | 4.2937 | 6.0327
P07306 | Asialoglycoprotein receptor 1 | ASGR1 | 1:1 | Block_4 | -5.5938 | 4.4243 | TRUE | 6.3539 | 3.4119 | -0.53554 | 0.16888 | 0.59526 | 0.034992 | 0.75072 | 1.709
Q9BZE9 | Tether containing UBX domain for GLUT4 | ASPSCR1 | 1:1 | Block_4 | 0.09746 | 9.5258 | FALSE | 19.0593 | 9.6209 | 0.35014 | 1.2178 | 2.5443 | 1.1985 | 2.4296 | 3.663
Q7L266 | Isoaspartyl peptidase/L-asparaginase | ASRGL1 | 1:1 | Block_4 | -0.22864 | 8.4333 | TRUE | 28.4788 | 5.9919 | NA | 0.98376 | 2.9358 | 0.66694 | 3.3343 | 5.6495
P00966 | Argininosuccinate synthase | ASS1 | 1:1 | Block_4 | -3.1355 | 3.8207 | FALSE | 7.5245 | 5.0164 | -2.1395 | -1.1968 | -0.13667 | -1.9361 | -0.69138 | 0.80216
Q676U5 | Autophagy-related protein 16-1 | ATG16L1 | 1:1 | Block_4 | -0.69853 | 9.3661 | FALSE | 12.7412 | 10.7438 | -0.52665 | 0.78819 | 2.8304 | 1.5079 | 3.0006 | 4.9055
Q9Y5K8 | V-type proton ATPase subunit D | ATP6V1D | 1:1 | Block_4 | 0.816 | 3.4253 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O75348 | V-type proton ATPase subunit G 1 | ATP6V1G1 | 1:1 | Block_4 | 0.18173 | 8.3141 | FALSE | 13.3759 | 8.2797 | NA | 1.8657 | 4.4757 | 2.097 | 4.6739 | 6.1322
O95670 | V-type proton ATPase subunit G 2 | ATP6V1G2 | 1:1 | Block_4 | 0.89681 | 9.0701 | TRUE | NA | NA | NA | 0.99507 | 2.2635 | NA | 2.2389 | 3.4205
Q6UW56 | All-trans retinoic acid-induced differentiation factor | ATRAID | 1:1 | Block_4 | -1.447 | 7.4543 | TRUE | 12.5185 | 17.9499 | -0.5436 | 0.012462 | 0.6822 | 0.04817 | 0.68658 | 1.891
P54252 | Ataxin-3 | ATXN3 | 1:1 | Block_4 | -2 | 6.9972 | TRUE | 32.2581 | 15.4896 | 0.17893 | 2.115 | 3.411 | 1.0731 | 3.5065 | 5.1009
O15169 | Axin-1 | AXIN1 | 1:1 | Block_4 | 0.16276 | 8.5942 | FALSE | NA | 3.5726 | NA | 0.65489 | 3.9324 | 1.2688 | 3.8397 | 5.4732
Q9H6S1 | 5-azacytidine-induced protein 2 | AZI2 | 1:1 | Block_4 | 0.27413 | 8.7472 | FALSE | NA | NA | NA | NA | 1.1649 | NA | 1.0336 | 3.2554
Q8NFL0 | UDP-GlcNAc:betaGal beta-1,3-N-acetylglucosaminyltransferase 7 | B3GNT7 | 1:1 | Block_4 | -4.5 | 4.8134 | FALSE | 10.4483 | 6.1716 | 0.032449 | 1.104 | 1.6143 | 0.79887 | 1.4936 | 2.4206
O14867 | Transcription regulator protein BACH1 | BACH1 | 1:1 | Block_4 | 0.47255 | 9.7579 | TRUE | NA | NA | NA | NA | 1.5972 | NA | 1.4078 | 3.4124
O95817 | BAG family molecular chaperone regulator 3 | BAG3 | 1:1 | Block_4 | -3.031 | 4.2249 | TRUE | 6.7267 | 4.8264 | -1.3065 | 0.73948 | 1.6417 | -0.2204 | 1.0592 | 1.8903
Q9UL15 | BAG family molecular chaperone regulator 5 | BAG5 | 1:1 | Block_4 | -0.3457 | 8.2585 | TRUE | 15.9173 | 5.6134 | -0.25947 | 1.4342 | 3.5602 | 1.8931 | 3.8632 | 5.3359
P46379 | Large proline-rich protein BAG6 | BAG6 | 1:1 | Block_4 | -4 | 5.374 | FALSE | 9.0048 | 8.2187 | -0.064439 | 0.79506 | 1.2245 | 0.73997 | 1.4952 | 2.1625
Q8NDB2 | B-cell scaffold protein with ankyrin repeats | BANK1 | 1:1 | Block_4 | 0.22433 | 6.7845 | FALSE | 13.3616 | 7.6761 | NA | 1.9613 | 4.7537 | 1.2247 | 4.3981 | 6.4293
Q8IXM2 | Chromatin complexes subunit BAP18 | BAP18 | 1:1 | Block_4 | -1.0308 | 6.4896 | TRUE | 13.0655 | 15.2808 | NA | -0.28006 | 0.73886 | -0.26797 | 1.029 | 3.0659
Q96GW7 | Brevican core protein | BCAN | 1:1 | Block_4 | -4.52 | 4.0228 | FALSE | 9.4497 | 12.5075 | -0.37795 | 0.20087 | 0.94111 | -0.32097 | 0.54029 | 1.2401
P10415 | Apoptosis regulator Bcl-2 | BCL2 | 1:1 | Block_4 | -2 | 13.0979 | TRUE | 14.1953 | 8.8763 | 0.15411 | 0.99991 | 1.6093 | 0.64938 | 1.8878 | 3.0975
Q07817 | Bcl-2-like protein 1 | BCL2L1 | 1:1 | Block_4 | 1.8366 | 8.5574 | TRUE | NA | NA | NA | NA | 4.9544 | NA | 4.472 | 6.2533
Q9BQE9 | B-cell CLL/lymphoma 7 protein family member B | BCL7B | 1:1 | Block_4 | 0.52238 | 6.1186 | TRUE | NA | NA | NA | NA | 1.3138 | NA | NA | 1.1721
P11274 | Breakpoint cluster region protein | BCR | 1:1 | Block_4 | -0.013207 | 8.177 | FALSE | 9.3906 | 5.7345 | NA | 0.54377 | 2.6694 | 0.93526 | 3.379 | 5.1123
Q14457 | Beclin-1 | BECN1 | 1:1 | Block_4 | 3.9218 | 13.3988 | FALSE | NA | NA | NA | NA | NA | NA | NA | 4.9971
Q8WV28 | B-cell linker protein | BLNK | 1:1 | Block_4 | -0.93931 | 7.4643 | FALSE | 14.0025 | 4.1508 | -0.65894 | 0.090273 | 0.84368 | -0.22785 | 0.65229 | 1.7529
Q6QNY0 | Biogenesis of lysosome-related organelles complex 1 subunit 3 | BLOC1S3 | 1:1 | Block_4 | 0.62661 | 6.3716 | TRUE | NA | NA | NA | NA | 1.6198 | NA | 2.0461 | 3.4329
O00327 | Aryl hydrocarbon receptor nuclear translocator-like protein 1 | BMAL1 | 1:1 | Block_4 | 0.86 | 4.5891 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P12644 | Bone morphogenetic protein 4 | BMP4 | 1:1 | Block_4 | -3.165 | 6.3471 | TRUE | 15.9177 | 16.1858 | -2.497 | -1.5237 | 0.20318 | -2.1201 | -0.71263 | 0.62404
P22004 | Bone morphogenetic protein 6 | BMP6 | 1:1 | Block_4 | 2.1092 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
P18075 | Bone morphogenetic protein 7 | BMP7 | 1:1 | Block_4 | -4 | 4.875 | FALSE | 15.3286 | 5.2406 | 0.18602 | 0.93239 | 1.4303 | -0.22389 | 0.96299 | 1.5975
O60238 | BCL2/adenovirus E1B 19 kDa protein-interacting protein 3-like | BNIP3L | 1:1 | Block_4 | -0.30425 | 10.5752 | TRUE | 11.1824 | 7.3199 | NA | 0.29246 | 1.1105 | 0.35194 | 1.5797 | 3.3437
Q9BWV1 | Brother of CDO | BOC | 1:1 | Block_4 | -5.8496 | 4.3281 | TRUE | 6.3318 | 4.7714 | -0.39736 | 0.18702 | 0.53055 | -0.19103 | 0.31963 | 0.89354
Q9H3K6 | BolA-like protein 2 | BOLA2 | 1:1 | Block_4 | -1.574 | 7.7114 | TRUE | 16.3894 | 5.8654 | 1.1799 | 2.6868 | 4.0286 | 2.6488 | 4.2094 | 4.9573
Q96DR5 | BPI fold-containing family A member 2 | BPIFA2 | 1:1 | Block_4 | -5 | 5.803 | TRUE | 11.9756 | 13.8864 | -3.1419 | -1.501 | 0.029714 | -3.0903 | -1.4235 | 0.38837
Q7Z569 | BRCA1-associated protein | BRAP | 1:1 | Block_4 | -0.54652 | 10.2779 | TRUE | 0.08604 | 11.4134 | NA | 0.40879 | 3.452 | 0.76796 | 3.4914 | 5.2048
Q58F21 | Bromodomain testis-specific protein | BRDT | 1:1 | Block_4 | 0.87198 | 12.3881 | TRUE | NA | NA | NA | NA | 1.7963 | NA | 1.467 | 2.5267
Q8WUW1 | Protein BRICK1 | BRK1 | 1:1 | Block_4 | -1.2724 | NA | NA | 3.988 | 3.558 | 0.0589 | 0.58516 | 1.0717 | 0.82125 | 1.5691 | 2.7
Q8IWQ3 | Serine/threonine-protein kinase BRSK2 | BRSK2 | 1:1 | Block_4 | 0.57116 | NA | NA | NA | NA | NA | NA | 2.1356 | NA | 1.4142 | 2.7421
P35613 | Basigin | BSG | 1:1 | Block_4 | -4.2665 | 6.7725 | TRUE | 4.9444 | 2.9068 | -0.25365 | 0.34518 | 0.78332 | -0.044021 | 0.55571 | 1.3549
Q13410 | Butyrophilin subfamily 1 member A1 | BTN1A1 | 1:1 | Block_4 | 1.3419 | 12.023 | TRUE | 3.7679 | 20.6028 | NA | NA | 2.6315 | NA | NA | 1.7275
P78410 | Butyrophilin subfamily 3 member A2 | BTN3A2 | 1:1 | Block_4 | -4 | 5.3216 | FALSE | 9.5389 | 3.9797 | 0.1588 | 0.61353 | 1.3698 | 0.6881 | 1.3427 | 2.2926
Q6UX73 | UPF0764 protein C16orf89 | C16orf89 | 1:1 | Block_4 | -2.3363 | 6.9741 | FALSE | 6.6616 | 5.3288 | -0.37703 | 0.54106 | 1.0584 | -0.25882 | 0.40792 | 1.0724
Q96EU7 | C1GALT1-specific chaperone 1 | C1GALT1C1 | 1:1 | Block_4 | 0.001 | 5.9876 | FALSE | 9.6111 | 10.0432 | NA | 0.47972 | 1.0729 | NA | 0.38869 | 0.97512
P02745 | Complement C1q subcomponent subunit A | C1QA | 1:1 | Block_4 | -1.7444 | 3.6054 | TRUE | 5.0376 | 2.9578 | -0.29034 | 0.20406 | 0.63287 | 0.05936 | 0.52952 | 1.1645
Q07021 | Complement component 1 Q subcomponent-binding protein, mitochondrial | C1QBP | 1:1 | Block_4 | 0.71182 | 10.642 | TRUE | NA | NA | NA | NA | 2.4112 | NA | 1.1062 | 2.2807
Q7Z5L3 | Complement C1q-like protein 2 | C1QL2 | 1:1 | Block_4 | -2.4119 | 7.1932 | FALSE | 10.252 | 6.8017 | -0.31717 | 0.43691 | 1.3944 | -0.78525 | 0.25716 | 1.4158
Q9BXI9 | Complement C1q tumor necrosis factor-related protein 6 | C1QTNF6 | 1:1 | Block_4 | -1 | 7.0929 | FALSE | 18.0066 | 10.0868 | -0.60571 | 0.34451 | 2.1598 | -0.15234 | 0.72933 | 1.858
Q9BRJ6 | Uncharacterized protein C7orf50 | C7orf50 | 1:1 | Block_4 | 0.001 | 6.9645 | FALSE | NA | NA | NA | NA | 0.072143 | NA | 0.23037 | 2.0567
Q8IXQ3 | Uncharacterized protein C9orf40 | C9orf40 | 1:1 | Block_4 | -3 | 7.0983 | TRUE | 18.8327 | 11.9506 | 1.0264 | 3.2458 | 4.0803 | 1.6664 | 3.0111 | 4.5932
O43570 | Carbonic anhydrase 12 | CA12 | 1:1 | Block_4 | -2.194 | NA | NA | 7.2019 | 7.7594 | -0.59757 | -0.1255 | 0.57907 | -0.038265 | 0.60626 | 1.6324
Q9ULX7 | Carbonic anhydrase 14 | CA14 | 1:1 | Block_4 | -2.7659 | 6.7184 | TRUE | 7.6615 | 12.9624 | -0.93302 | -0.23326 | 0.54965 | -0.72108 | -0.097808 | 0.73089
Q16790 | Carbonic anhydrase 9 | CA9 | 1:1 | Block_4 | -1.0129 | 7.8567 | TRUE | 25.2776 | 7.1726 | -0.82065 | 0.038501 | 1.3462 | -0.71856 | 0.51243 | 1.7891
P54284 | Voltage-dependent L-type calcium channel subunit beta-3 | CACNB3 | 1:1 | Block_4 | 0.5 | 13.6461 | FALSE | NA | NA | NA | 0.88747 | 3.8195 | 1.061 | 3.9593 | 6.0297
P05937 | Calbindin | CALB1 | 1:1 | Block_4 | -4.5261 | 9.7398 | TRUE | 5.7045 | 6.5405 | -0.54106 | 0.089032 | 0.51869 | 0.11286 | 1.0653 | 2.2296
P01258 | Calcitonin | CALCA | 1:1 | Block_4 | -4.6347 | 7.8033 | TRUE | 8.554 | 3.0388 | -1.5099 | -0.60415 | 0.47093 | -0.40127 | 0.86276 | 2.9394
P10092 | Calcitonin gene-related peptide 2 | CALCB | 1:1 | Block_4 | -0.69954 | 6.4684 | FALSE | 10.7225 | 10.6165 | -0.55592 | -0.054831 | 0.76784 | -0.13794 | 0.592 | 2.0649
Q9P1Z2 | Calcium-binding and coiled-coil domain-containing protein 1 | CALCOCO1 | 1:1 | Block_4 | -1.8327 | 7.9229 | TRUE | 9.8399 | 5.5749 | -0.86668 | 1.4141 | 4.007 | 1.2298 | 3.4469 | 6.27
Q13137 | Calcium-binding and coiled-coil domain-containing protein 2 | CALCOCO2 | 1:1 | Block_4 | 0.001 | 7.7904 | TRUE | NA | 7.4453 | NA | 0.43403 | 2.252 | 0.29442 | 1.986 | 3.7961
Q96GE6 | Calmodulin-like protein 4 | CALML4 | 1:1 | Block_4 | -1.5424 | 8.5725 | TRUE | 12.4667 | 7.2963 | -0.62177 | 0.012873 | 0.7078 | 0.37739 | 1.0676 | 2.4922
Q5T5Y3 | Calmodulin-regulated spectrin-associated protein 1 | CAMSAP1 | 1:1 | Block_4 | -1.392 | 10.9558 | TRUE | 14.6227 | 8.0744 | -1.3552 | 0.46498 | 4.3179 | 0.8802 | 3.5564 | 6.3028
P40121 | Macrophage-capping protein | CAPG | 1:1 | Block_4 | -3.7927 | 3.0895 | FALSE | 7.5091 | 10.6875 | -0.7406 | 0.016208 | 0.69252 | -0.11641 | 0.97966 | 3.282
P20807 | Calpain-3 | CAPN3 | 1:1 | Block_4 | -1.22 | 10.9488 | FALSE | 13.7048 | 10.9304 | -0.61003 | 0.072818 | 0.97298 | -0.43508 | 0.28898 | 1.2163
Q13938 | Calcyphosin | CAPS | 1:1 | Block_4 | -2 | 6.0746 | TRUE | 17.6279 | 15.0826 | -0.52711 | 0.22359 | 1.2023 | 0.14061 | 0.8955 | 2.0204
Q92851 | Caspase-10 | CASP10 | 1:1 | Block_4 | -0.1592 | 7.7321 | TRUE | 6.5873 | 3.1298 | NA | 0.14776 | 0.89638 | 0.0006526 | 1.1958 | 3.1336
P42575 | Caspase-2 | CASP2 | 1:1 | Block_4 | -0.67637 | 8.5563 | FALSE | NA | NA | NA | NA | 0.53991 | NA | 0.59707 | 1.9583
P55210 | Caspase-7 | CASP7 | 1:1 | Block_4 | 0.5 | 6.4442 | FALSE | 24.7797 | 11.9576 | NA | 1.2975 | 3.1687 | 0.97936 | 3.1025 | 5.1541
Q14790 | Caspase-8 | CASP8 | 1:1 | Block_4 | -2 | 8.7565 | TRUE | 20.1997 | 20.5365 | -0.34238 | 0.97272 | 3.2869 | 1.7223 | 3.4235 | 5.8598
P55211 | Caspase-9 | CASP9 | 1:1 | Block_4 | 1.08 | 5.8859 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.2787
Q9NTU7 | Cerebellin-4 | CBLN4 | 1:1 | Block_4 | -2 | 5.1041 | TRUE | 12.9514 | 9.4408 | -0.14033 | 0.48506 | 0.92625 | -0.49825 | 0.58277 | 1.468
P35520 | Cystathionine beta-synthase | CBS | 1:1 | Block_4 | -1.4424 | 8.3783 | FALSE | 12.9592 | 8.5871 | -1.3885 | -0.61877 | 0.37375 | NA | -0.22084 | 1.5155
Q6P1N0 | Coiled-coil and C2 domain-containing protein 1A | CC2D1A | 1:1 | Block_4 | -2.7316 | 6.874 | FALSE | 8.3963 | 4.8601 | -1.375 | 0.39002 | 2.833 | 0.97246 | 3.0512 | 4.5804
Q8IX12 | Cell division cycle and apoptosis regulator protein 1 | CCAR1 | 1:1 | Block_4 | -1.5734 | 8.8877 | TRUE | 6.9121 | 4.4448 | -1.0272 | -0.37664 | 0.33445 | -0.26808 | 0.36816 | 2.2191
Q8N163 | Cell cycle and apoptosis regulator protein 2 | CCAR2 | 1:1 | Block_4 | -0.63591 | 7.929 | TRUE | 11.0913 | 7.6747 | -0.49402 | 0.60462 | 1.6109 | 0.56238 | 1.4852 | 3.1233
Q9H6E4 | Coiled-coil domain-containing protein 134 | CCDC134 | 1:1 | Block_4 | 0.5 | 9.7752 | FALSE | NA | 2.5566 | NA | 0.98145 | 3.3437 | 0.93008 | 3.3004 | 5.264
Q8IWP9 | Coiled-coil domain-containing protein 28A | CCDC28A | 1:1 | Block_4 | 0.56471 | 4.4558 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.60093
Q8IVM0 | Coiled-coil domain-containing protein 50 | CCDC50 | 1:1 | Block_4 | 0.15636 | 7.8659 | FALSE | 6.6251 | 7.7437 | NA | 0.40452 | 0.82835 | NA | 0.9749 | 1.9557
I3L3R5 | Coiled-coil domain-containing glutamate-rich protein 2 | CCER2 | 1:1 | Block_4 | -2.8233 | 6.2247 | TRUE | 9.4208 | 5.5074 | -0.88904 | -0.20242 | 0.50045 | -0.022542 | 0.87539 | 2.1136
P06307 | Cholecystokinin | CCK | 1:1 | Block_4 | 0.38284 | 9.6817 | FALSE | 13.0797 | 20.8745 | 1.2621 | 3.5755 | 5.4271 | 1.038 | 3.2959 | 5.4397
P51671 | Eotaxin | CCL11 | 1:1 | Block_4 | -4.8468 | 7.8547 | FALSE | 6.9768 | 3.6932 | -0.10867 | 0.6182 | 1.2494 | 0.39029 | 1.0857 | 1.8781
P78556 | C-C motif chemokine 20 | CCL20 | 1:1 | Block_4 | -7 | 5.7777 | TRUE | 5.9326 | 6.9666 | -2.0576 | -1.1228 | -0.13299 | -1.5798 | 0.015332 | 2.444
O15444 | C-C motif chemokine 25 | CCL25 | 1:1 | Block_4 | -4.5188 | 5.2366 | TRUE | 9.661 | 4.9412 | -0.59202 | 0.41289 | 1.1846 | -0.15658 | 0.82984 | 1.7377
Q9Y258 | C-C motif chemokine 26 | CCL26 | 1:1 | Block_4 | -2 | 11.1208 | FALSE | 13.6071 | 14.5768 | -0.22707 | 0.69927 | 3.1237 | 0.25284 | 1.3308 | 3.0677
Q9NRJ3 | C-C motif chemokine 28 | CCL28 | 1:1 | Block_4 | -1.5864 | 6.1714 | TRUE | 11.1757 | 9.5319 | 0.08072 | 1.4549 | 2.9246 | 1.1373 | 2.5163 | 4.34
P10147 | C-C motif chemokine 3 | CCL3 | 1:1 | Block_4 | -3.5957 | 9.1564 | TRUE | 6.5478 | 5.4071 | -1.5923 | -0.71275 | -0.24713 | -0.73117 | 0.24421 | 1.2953
P80098 | C-C motif chemokine 7 | CCL7 | 1:1 | Block_4 | -1.5565 | 9.9651 | TRUE | 22.6179 | 9.2151 | NA | -0.89452 | 0.49784 | -0.62057 | 0.27062 | 1.57
P80075 | C-C motif chemokine 8 | CCL8 | 1:1 | Block_4 | -5.6567 | 6.7182 | TRUE | 7.0266 | 7.2998 | -1.1459 | 0.039633 | 0.90632 | -0.0032453 | 1.2017 | 2.2626
O00622 | CCN family member 1 | CCN1 | 1:1 | Block_4 | -5 | 4.261 | TRUE | 7.6866 | 6.1662 | -0.46115 | 1.036 | 1.869 | 0.29199 | 1.3462 | 2.257
O95388 | CCN family member 4 | CCN4 | 1:1 | Block_4 | -7 | 5.0715 | TRUE | 4.5718 | 5.2016 | -0.34819 | 0.54752 | 1.0161 | 0.64574 | 1.3553 | 2.2533
O14618 | Copper chaperone for superoxide dismutase | CCS | 1:1 | Block_4 | 0.85951 | 6.9906 | FALSE | 16.8623 | 16.6649 | 0.92702 | 2.4996 | 3.8855 | 2.1938 | 3.3318 | 4.5767
Q93033 | Immunoglobulin superfamily member 2 | CD101 | 1:1 | Block_4 | -4 | 6.4666 | FALSE | 11.7536 | 9.4863 | -0.32474 | 0.29353 | 0.79039 | -0.63157 | 0.28826 | 0.90951
O95971 | CD160 antigen | CD160 | 1:1 | Block_4 | -5 | 5.0589 | TRUE | 9.0774 | 8.6825 | -0.073825 | 0.34989 | 1.2331 | -0.17891 | 0.76952 | 2.3858
P41217 | OX-2 membrane glycoprotein | CD200 | 1:1 | Block_4 | -3 | 3.2007 | TRUE | 15.2087 | 9.2267 | -0.27529 | 0.35315 | 0.69214 | -0.043388 | 0.52132 | 1.1607
Q8TD46 | Cell surface glycoprotein CD200 receptor 1 | CD200R1 | 1:1 | Block_4 | -5.2576 | 4.8777 | TRUE | 5.3316 | 5.4353 | -0.67007 | 0.20742 | 0.65368 | -0.18324 | 0.48116 | 1.1885
Q9UJ71 | C-type lectin domain family 4 member K | CD207 | 1:1 | Block_4 | -1.5 | 6.1763 | TRUE | 14.6074 | 13.8433 | -0.80602 | -0.14919 | 0.68672 | -1.0214 | -0.030972 | 0.85602
P20273 | B-cell receptor CD22 | CD22 | 1:1 | Block_4 | -2.7703 | 4.8714 | FALSE | 4.654 | 6.6557 | -0.63446 | 0.40616 | 0.89402 | -0.87099 | 0.39581 | 1.5342
Q15762 | CD226 antigen | CD226 | 1:1 | Block_4 | -0.63443 | 3.7695 | FALSE | 16.7014 | 7.7829 | -0.058739 | 0.37653 | 1.1193 | 0.22546 | 1.1103 | 1.9904
Q9BZW8 | Natural killer cell receptor 2B4 | CD244 | 1:1 | Block_4 | -3 | 5.5879 | TRUE | 11.3361 | 7.7463 | -0.21313 | 0.44985 | 0.76101 | 0.1139 | 0.78616 | 1.4227
Q9NZQ7 | Programmed cell death 1 ligand 1 | CD274 | 1:1 | Block_4 | -4 | 5.8694 | TRUE | 14.1171 | 6.867 | -0.63489 | 0.033367 | 0.62765 | 0.030027 | 0.80485 | 1.9194
Q496F6 | CMRF35-like molecule 2 | CD300E | 1:1 | Block_4 | -5 | 5.5475 | FALSE | 8.46 | 5.8878 | -0.78176 | 0.40127 | 1.0601 | 0.17772 | 1.2326 | 2.6665
Q8TDQ1 | CMRF35-like molecule 1 | CD300LF | 1:1 | Block_4 | -5 | 4.5005 | TRUE | 9.9051 | 10.5232 | -1.4878 | 0.25743 | 1.2169 | 0.11072 | 1.6098 | 2.978
Q8IX05 | CD302 antigen | CD302 | 1:1 | Block_4 | -1.9218 | 4.5154 | FALSE | 11.0995 | 10.9302 | -0.089548 | 0.37499 | 0.92566 | 0.37306 | 1.2364 | 2.6748
P28907 | ADP-ribosyl cyclase/cyclic ADP-ribose hydrolase 1 | CD38 | 1:1 | Block_4 | -4.5 | 7.1305 | TRUE | 8.1273 | 8.84 | 0.055236 | 0.73818 | 1.2797 | 0.4558 | 1.1177 | 2.3221
P01730 | T-cell surface glycoprotein CD4 | CD4 | 1:1 | Block_4 | -0.35436 | 7.7949 | TRUE | 12.2146 | 10.3259 | 0.11303 | 0.47254 | 0.97138 | 0.33858 | 0.87343 | 1.78
P29965 | CD40 ligand | CD40LG | 1:1 | Block_4 | -0.5 | 7.2654 | TRUE | 33.2528 | 22.3837 | -0.015532 | 1.8479 | 3.5385 | 1.1474 | 3.6538 | 5.2667
P06127 | T-cell surface glycoprotein CD5 | CD5 | 1:1 | Block_4 | -2.5837 | 6.6545 | TRUE | 8.0871 | 7.7371 | -0.32958 | 0.41974 | 0.98723 | -0.07174 | 0.8215 | 1.5212
P30203 | T-cell differentiation antigen CD6 | CD6 | 1:1 | Block_4 | -4.5 | 5.2312 | TRUE | 8.6115 | 3.763 | -0.15106 | 0.65855 | 1.7802 | -0.2263 | 0.88543 | 1.7524
P08962 | CD63 antigen | CD63 | 1:1 | Block_4 | -4.258 | 6.0121 | TRUE | 7.9423 | 4.5813 | -0.40456 | 0.5406 | 1.6226 | 0.40579 | 1.3718 | 2.199
P09564 | T-cell antigen CD7 | CD7 | 1:1 | Block_4 | -0.048505 | 6.5943 | TRUE | 22.8011 | 15.0606 | NA | 0.20028 | 0.70649 | NA | 0.6439 | 1.6678
P32970 | CD70 antigen | CD70 | 1:1 | Block_4 | -2.8646 | 6.1721 | TRUE | 9.396 | 5.914 | -0.67982 | -0.15768 | 0.47408 | -0.38482 | 0.35341 | 1.098
P40259 | B-cell antigen receptor complex-associated protein beta chain | CD79B | 1:1 | Block_4 | -7 | 8.7487 | TRUE | 10.4305 | 4.1548 | -0.2682 | 0.80678 | 1.4112 | 0.46726 | 1.5021 | 3.1126
Q01151 | CD83 antigen | CD83 | 1:1 | Block_4 | -1.6531 | 7.5404 | TRUE | 9.5033 | 6.4925 | -0.65528 | 0.33513 | 0.71359 | -0.18649 | 0.64499 | 1.5678
Q9UIB8 | SLAM family member 5 | CD84 | 1:1 | Block_4 | -2.4126 | 3.1799 | FALSE | 15.9519 | 7.373 | -0.13972 | 0.5306 | 1.1601 | 0.1171 | 1.1064 | 1.8772
P42081 | T-lymphocyte activation antigen CD86 | CD86 | 1:1 | Block_4 | -4 | 5.4382 | TRUE | 7.5251 | 9.7296 | -0.5206 | 0.17061 | 0.64446 | -0.64913 | 0.19786 | 0.9265
Q8IWY9 | Codanin-1 | CDAN1 | 1:1 | Block_4 | 1.1521 | 3.3171 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NHZ8 | Anaphase-promoting complex subunit CDC26 | CDC26 | 1:1 | Block_4 | 0.001 | 7.7167 | TRUE | 22.0355 | 16.0018 | NA | 0.98571 | 3.2752 | 1.6096 | 3.4609 | 5.0688
Q9Y5S2 | Serine/threonine-protein kinase MRCK beta | CDC42BPB | 1:1 | Block_4 | 0.1456 | 8.0603 | FALSE | NA | NA | NA | 1.261 | 3.8796 | 0.86414 | 3.7491 | 6.4098
Q9H5V8 | CUB domain-containing protein 1 | CDCP1 | 1:1 | Block_4 | -6 | 4.4636 | TRUE | 10.9003 | 6.1412 | -0.8153 | 0.43508 | 0.9024 | -0.34646 | 0.92208 | 2.211
Q9BYE9 | Cadherin-related family member 2 | CDHR2 | 1:1 | Block_4 | -1.3941 | 4.7439 | FALSE | 8.9599 | 5.0833 | NA | -0.461 | 0.31942 | -1.2245 | 0.10828 | 1.718
P38936 | Cyclin-dependent kinase inhibitor 1 | CDKN1A | 1:1 | Block_4 | -1 | 11.0856 | FALSE | 16.8922 | 25.8382 | -0.1965 | 1.7665 | 4.9693 | 2.5333 | 4.8981 | 6.9224
P55273 | Cyclin-dependent kinase 4 inhibitor D | CDKN2D | 1:1 | Block_4 | 0.70159 | 3.7422 | FALSE | 10.3043 | 6.217 | NA | 0.99148 | 4.0421 | 1.1162 | 4.2941 | 5.3172
Q49AH0 | Cerebral dopamine neurotrophic factor | CDNF | 1:1 | Block_4 | -1.5451 | 8.6418 | TRUE | 9.473 | 11.9637 | -0.23516 | 0.4688 | 1.0349 | 0.53791 | 1.3133 | 2.1524
Q15517 | Corneodesmosin | CDSN | 1:1 | Block_4 | -3.5 | 3.828 | FALSE | 13.7247 | 10.6848 | -0.69997 | -0.026269 | 0.75876 | -0.13201 | 0.72346 | 1.7652
Q2WEN9 | Carcinoembryonic antigen-related cell adhesion molecule 16 | CEACAM16 | 1:1 | Block_4 | -5 | 6.6063 | TRUE | 7.6255 | 10.5343 | -0.87086 | 0.045466 | 2.0611 | -1.4061 | 0.26519 | 1.7546
Q7Z692 | Carcinoembryonic antigen-related cell adhesion molecule 19 | CEACAM19 | 1:1 | Block_4 | -0.94626 | 7.7229 | TRUE | 12.2124 | 6.6969 | -0.44597 | 0.003118 | 0.5118 | -0.093193 | 0.39792 | 1.272
Q3KPI0 | Carcinoembryonic antigen-related cell adhesion molecule 21 | CEACAM21 | 1:1 | Block_4 | -2.9876 | 9.097 | TRUE | 10.0426 | 9.0597 | -2.1436 | -0.17477 | 1.1843 | -1.3157 | -0.018587 | 2.1953
P31997 | Carcinoembryonic antigen-related cell adhesion molecule 8 | CEACAM8 | 1:1 | Block_4 | -3.9679 | 5.6812 | FALSE | 9.5869 | 12.1567 | -0.53971 | 0.17194 | 0.81648 | 0.14503 | 1.1922 | 2.9712
Q66GS9 | Centrosomal protein of 135 kDa | CEP135 | 1:1 | Block_4 | 0.001 | 8.2329 | TRUE | 13.9531 | 11.178 | NA | 0.62998 | 3.1497 | 0.27687 | 2.8851 | 4.745
Q5SW79 | Centrosomal protein of 170 kDa | CEP170 | 1:1 | Block_4 | 0.49251 | 4.6744 | TRUE | 5.3128 | 5.7292 | NA | 0.74359 | 2.7007 | 0.7912 | 2.5213 | 4.7467
Q9BV73 | Centrosome-associated protein CEP250 | CEP250 | 1:1 | Block_4 | -0.4267 | 11.1559 | TRUE | 22.6716 | 8.0724 | NA | 0.79703 | 3.7627 | 0.89189 | 3.4791 | 5.8354
O95684 | Centrosomal protein 43 | CEP43 | 1:1 | Block_4 | 0.56828 | 6.3311 | FALSE | NA | NA | NA | NA | 1.5853 | 0.59389 | 2.1651 | 3.5427
O00748 | Cocaine esterase | CES2 | 1:1 | Block_4 | -3.7961 | 5.3181 | FALSE | 9.5633 | 3.1372 | -0.7235 | 0.56175 | 1.6744 | -0.94364 | 0.23881 | 3.5742
Q6UWW8 | Carboxylesterase 3 | CES3 | 1:1 | Block_4 | -4 | 3.0963 | FALSE | 15.5909 | 11.7416 | -2.6621 | -1.0267 | 0.54412 | -2.7312 | -0.84833 | 0.92632
O15182 | Centrin-3 | CETN3 | 1:1 | Block_4 | -0.14906 | 8.3051 | TRUE | 7.609 | 4.8506 | -0.14298 | 0.6377 | 2.6495 | 0.98039 | 2.465 | 4.2047
Q96G28 | Cilia- and flagella-associated protein 36 | CFAP36 | 1:1 | Block_4 | -0.98061 | 9.5822 | TRUE | 17.1042 | 8.7991 | -0.71172 | 0.58548 | 3.6407 | 1.5508 | 3.6172 | 5.0957
Q8WYQ3 | Coiled-coil-helix-coiled-coil-helix domain-containing protein 10, mitochondrial | CHCHD10 | 1:1 | Block_4 | -2.3962 | 5.4554 | TRUE | 9.4994 | 8.6161 | -0.64408 | 0.075695 | 0.63982 | 0.37077 | 1.2818 | 3.2051
Q9BRQ6 | MICOS complex subunit MIC25 | CHCHD6 | 1:1 | Block_4 | -0.8817 | 10.0347 | TRUE | 13.2005 | 13.7195 | -0.51133 | 0.16758 | 0.86924 | 0.29823 | 1.0087 | 2.1872
P24386 | Rab proteins geranylgeranyltransferase component A 1 | CHM | 1:1 | Block_4 | 1 | 11.2383 | TRUE | NA | NA | NA | NA | 1.6909 | NA | 1.8822 | 3.3916
Q9HD42 | Charged multivesicular body protein 1a | CHMP1A | 1:1 | Block_4 | 0.89761 | 5.3449 | FALSE | NA | NA | NA | 1.2737 | 4.3833 | 1.6694 | 4.106 | 5.6988
Q14011 | Cold-inducible RNA-binding protein | CIRBP | 1:1 | Block_4 | 0.001 | 11.2694 | TRUE | 13.5968 | 18.8541 | 0.023693 | 1.7346 | 5.1647 | 2.8071 | 5.5212 | 6.9068
O14578 | Citron Rho-interacting kinase | CIT | 1:1 | Block_4 | 0.71252 | 11.5654 | TRUE | NA | NA | NA | NA | 0.76483 | NA | 0.75562 | 1.9253
Q07065 | Cytoskeleton-associated protein 4 | CKAP4 | 1:1 | Block_4 | -3.5 | 5.057 | FALSE | 11.659 | 10.4621 | -0.39484 | 0.35126 | 0.86694 | 0.10398 | 0.908 | 1.9597
Q8IUN9 | C-type lectin domain family 10 member A | CLEC10A | 1:1 | Block_4 | -2.8317 | 5.4879 | TRUE | 8.7865 | 13.6977 | -0.16089 | 0.54194 | 1.0397 | -0.17791 | 0.54242 | 1.2997
Q5QGZ9 | C-type lectin domain family 12 member A | CLEC12A | 1:1 | Block_4 | 0.32489 | 8.3561 | FALSE | NA | NA | NA | NA | 0.65977 | NA | 0.45613 | 1.7274
Q8NC01 | C-type lectin domain family 1 member A | CLEC1A | 1:1 | Block_4 | -2.6965 | 8.3592 | TRUE | 9.2861 | 5.0759 | -0.77229 | 0.08272 | 0.67191 | -0.21752 | 0.68946 | 1.7579
Q9UMR7 | C-type lectin domain family 4 member A | CLEC4A | 1:1 | Block_4 | -3.0671 | 6.0275 | TRUE | 8.387 | 5.1706 | -0.76886 | -0.21065 | 0.51496 | -0.77966 | -0.042869 | 0.79589
Q8WTT0 | C-type lectin domain family 4 member C | CLEC4C | 1:1 | Block_4 | -1.5597 | 8.904 | FALSE | 8.6899 | 3.9207 | -0.21421 | 0.32744 | 1.075 | -0.64237 | 0.45951 | 1.2529
Q8WXI8 | C-type lectin domain family 4 member D | CLEC4D | 1:1 | Block_4 | -1.4684 | 8.5449 | TRUE | 12.0799 | 9.0272 | 0.092347 | 0.886 | 1.3387 | 0.53287 | 1.7352 | 3.9034
Q6UXB4 | C-type lectin domain family 4 member G | CLEC4G | 1:1 | Block_4 | -0.86423 | 8.135 | TRUE | 12.2666 | 11.973 | -0.57577 | 0.1128 | 0.6336 | -0.38022 | 0.38912 | 1.0358
Q9BXN2 | C-type lectin domain family 7 member A | CLEC7A | 1:1 | Block_4 | -4 | 5.3083 | TRUE | 17.7772 | 12.1764 | -0.20763 | 0.56678 | 1.2092 | -0.10276 | 1.1488 | 2.597
Q9UDT6 | CAP-Gly domain-containing linker protein 2 | CLIP2 | 1:1 | Block_4 | -1.8402 | 6.7791 | TRUE | 12.9694 | 10.5954 | -1.5836 | 1.7203 | 4.7121 | 1.5553 | 4.6946 | 5.8235
Q9H6B4 | CXADR-like membrane protein | CLMP | 1:1 | Block_4 | -3 | 2.6916 | TRUE | 14.505 | 5.6543 | -0.38732 | 0.1173 | 0.46874 | 0.06091 | 0.65189 | 1.3622
Q16740 | ATP-dependent Clp protease proteolytic subunit, mitochondrial | CLPP | 1:1 | Block_4 | 0.46319 | 9.174 | TRUE | NA | 7.94 | NA | 1.741 | 4.7506 | 1.6498 | 4.0643 | 5.9095
Q9H4D0 | Calsyntenin-2 | CLSTN2 | 1:1 | Block_4 | -4 | 5.9295 | TRUE | 11.6783 | 4.2808 | -0.043971 | 0.44262 | 1.0862 | 0.13674 | 1.042 | 2.3245
Q9BQT9 | Calsyntenin-3 | CLSTN3 | 1:1 | Block_4 | -1.2693 | 5.7204 | TRUE | 6.7061 | 7.321 | -0.27175 | 0.56513 | 1.2247 | 0.34044 | 1.0265 | 1.8536
Q8IY22 | C-Maf-inducing protein | CMIP | 1:1 | Block_4 | 0.54038 | 5.1503 | FALSE | NA | NA | NA | 0.72729 | 3.5804 | NA | 3.2077 | 4.9974
Q8N3K9 | Cardiomyopathy-associated protein 5 | CMYA5 | 1:1 | Block_4 | -0.10994 | 10.397 | TRUE | 8.7425 | 7.3919 | NA | NA | 1.6661 | NA | 0.43763 | 2.3808
P09543 | 2',3'-cyclic-nucleotide 3'-phosphodiesterase | CNP | 1:1 | Block_4 | -3 | 9.5887 | FALSE | 17.647 | 11.6304 | 0.12447 | 2.2096 | 4.1699 | 2.0274 | 4.1339 | 5.5311
Q8N129 | Protein canopy homolog 4 | CNPY4 | 1:1 | Block_4 | -1 | 6.2843 | TRUE | 20.8821 | 11.4616 | -0.54917 | 0.84406 | 2.2835 | 1.0131 | 2.4441 | 3.8014
Q02246 | Contactin-2 | CNTN2 | 1:1 | Block_4 | -1.0146 | 3.4502 | TRUE | 16.3955 | 8.2611 | -0.40593 | 0.44776 | 1.0668 | -0.086003 | 0.48665 | 1.2836
Q9UHC6 | Contactin-associated protein-like 2 | CNTNAP2 | 1:1 | Block_4 | -0.88327 | 7.2513 | TRUE | 8.0124 | 4.5854 | -0.61473 | -0.12692 | 0.34557 | -0.44508 | 0.15064 | 0.78816
Q17RW2 | Collagen alpha-1(XXIV) chain | COL24A1 | 1:1 | Block_4 | -3.1103 | 6.384 | TRUE | 8.5931 | 10.2021 | -0.12153 | 0.46879 | 1.2279 | 0.081775 | 0.89572 | 1.7098
Q2UY09 | Collagen alpha-1(XXVIII) chain | COL28A1 | 1:1 | Block_4 | -1.6562 | 11.0446 | TRUE | 11.5218 | 5.4778 | -0.73235 | 0.87919 | 2.0358 | -0.45137 | 0.59651 | 1.7493
P02458 | Collagen alpha-1(II) chain | COL2A1 | 1:1 | Block_4 | -2.9033 | 6.2141 | TRUE | 8.3025 | 5.5282 | 0.56988 | 0.91567 | 1.6056 | 0.51749 | 1.4612 | 3.186
P20849 | Collagen alpha-1(IX) chain | COL9A1 | 1:1 | Block_4 | -2.1461 | 10.6449 | TRUE | 12.6163 | 3.2107 | -0.60999 | 0.39705 | 1.2594 | -0.022956 | 0.85445 | 2.4065
Q8N668 | COMM domain-containing protein 1 | COMMD1 | 1:1 | Block_4 | -0.65362 | 6.9992 | TRUE | 9.1965 | 6.2084 | NA | 0.082422 | 1.3183 | 0.56704 | 1.882 | 3.1779
Q9P000 | COMM domain-containing protein 9 | COMMD9 | 1:1 | Block_4 | -0.91017 | 9.4136 | TRUE | 13.8274 | 12.5794 | -0.84791 | 0.26442 | 1.1837 | 0.48768 | 1.3804 | 2.4627
P21964 | Catechol O-methyltransferase | COMT | 1:1 | Block_4 | -1.3581 | 7.0533 | FALSE | 16.6869 | 2.7433 | -0.44623 | 1.394 | 2.8875 | 0.64666 | 3.1106 | 4.8612
P35606 | Coatomer subunit beta' | COPB2 | 1:1 | Block_4 | 1.749 | 6.0369 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q99807 | 5-demethoxyubiquinone hydroxylase, mitochondrial | COQ7 | 1:1 | Block_4 | -0.77339 | 8.6686 | TRUE | 17.7361 | 6.8444 | -0.66807 | -0.0039445 | 0.70316 | -0.54071 | 0.18617 | 0.99209
P10606 | Cytochrome c oxidase subunit 5B, mitochondrial | COX5B | 1:1 | Block_4 | 0.40826 | 6.8601 | TRUE | NA | 0.36374 | NA | 0.90251 | 4.3567 | 1.3153 | 3.1088 | 5.2839
P16870 | Carboxypeptidase E | CPE | 1:1 | Block_4 | -1.1678 | 4.7077 | FALSE | 7.8788 | 5.5553 | -0.26361 | 0.29055 | 0.68264 | -0.028689 | 0.35654 | 0.74122
Q9H3G5 | Probable serine carboxypeptidase CPVL | CPVL | 1:1 | Block_4 | -3.5231 | 3.9455 | TRUE | 10.6812 | 7.5339 | -0.45627 | 0.81047 | 1.7837 | -0.99142 | 0.58139 | 1.5348
P78560 | Death domain-containing protein CRADD | CRADD | 1:1 | Block_4 | -2.4852 | 7.9703 | FALSE | 12.2452 | 4.6855 | -0.57286 | 1.8206 | 4.3195 | 1.3106 | 4.6929 | 5.9814
P52943 | Cysteine-rich protein 2 | CRIP2 | 1:1 | Block_4 | -3.0668 | 7.0212 | TRUE | 5.825 | 7.8915 | -0.46504 | 0.20314 | 0.87162 | -0.5332 | 0.51136 | 1.3093
P16562 | Cysteine-rich secretory protein 2 | CRISP2 | 1:1 | Block_4 | -3.8486 | 5.4715 | TRUE | 8.5758 | 3.6867 | -1.112 | -0.0033617 | 1.0492 | -0.83608 | 0.21561 | 1.0141
Q9UBG3 | Cornulin | CRNN | 1:1 | Block_4 | -4.5 | 5.4028 | TRUE | 7.3665 | 5.4664 | 0.062015 | 0.67267 | 1.6346 | -0.56865 | 0.46999 | 1.689
O95727 | Cytotoxic and regulatory T-cell molecule | CRTAM | 1:1 | Block_4 | -3.6709 | 5.1216 | TRUE | 8.8041 | 8.1678 | -0.12028 | 0.51729 | 1.6059 | 0.070967 | 0.97025 | 1.7812
P43320 | Beta-crystallin B2 | CRYBB2 | 1:1 | Block_4 | 0.2422 | 11.3025 | TRUE | 6.3746 | 5.6657 | NA | NA | 1.5033 | NA | 0.48767 | 5.7758
P07320 | Gamma-crystallin D | CRYGD | 1:1 | Block_4 | -1.2172 | 8.0105 | TRUE | 10.1654 | 11.7186 | 0.056062 | 1.5954 | 3.71 | 1.9298 | 4.1635 | 6.2176
O95825 | Quinone oxidoreductase-like protein 1 | CRYZL1 | 1:1 | Block_4 | -4 | 8.9664 | TRUE | 12.4929 | 13.1135 | 0.38683 | 2.6122 | 5.3232 | 2.397 | 5.3946 | 7.0026
O75534 | Cold shock domain-containing protein E1 | CSDE1 | 1:1 | Block_4 | 0.6818 | 7.5516 | FALSE | NA | NA | NA | 0.93386 | 2.7367 | NA | 2.6621 | 5.1758
P15509 | Granulocyte-macrophage colony-stimulating factor receptor subunit alpha | CSF2RA | 1:1 | Block_4 | -1.3311 | 3.4481 | FALSE | 9.8456 | 11.7993 | 0.086853 | 0.70958 | 1.5357 | 0.030541 | 0.85873 | 1.679
P32927 | Cytokine receptor common subunit beta | CSF2RB | 1:1 | Block_4 | -4 | 7.9431 | TRUE | 9.9847 | 13.1856 | -0.66569 | 0.24874 | 0.98944 | -0.82478 | 0.32517 | 1.2928
Q99062 | Granulocyte colony-stimulating factor receptor | CSF3R | 1:1 | Block_4 | -0.99764 | 6.0448 | TRUE | 11.7117 | 3.3741 | -0.48812 | 0.067166 | 0.93408 | -0.11833 | 0.41143 | 1.2125
P68400 | Casein kinase II subunit alpha | CSNK2A1 | 1:1 | Block_4 | 1.0397 | 4.8617 | FALSE | NA | NA | NA | NA | 2.5197 | NA | 1.8932 | 3.1618
P50461 | Cysteine and glycine-rich protein 3 | CSRP3 | 1:1 | Block_4 | -0.11359 | 8.253 | TRUE | 13.8293 | 8.1436 | NA | 0.041915 | 2.0575 | NA | 1.0747 | 3.0504
P28325 | Cystatin-D | CST5 | 1:1 | Block_4 | -6 | 3.5565 | TRUE | 6.384 | 5.4035 | -1.1319 | -0.26856 | 0.49807 | -0.59241 | 0.070537 | 1.2276
P32929 | Cystathionine gamma-lyase | CTH | 1:1 | Block_4 | -1.026 | 7.0811 | TRUE | 13.6831 | 12.4427 | -0.82749 | 1.1423 | 2.9305 | 0.52809 | 2.3744 | 3.756
P40313 | Chymotrypsin-like protease CTRL-1 | CTRL | 1:1 | Block_4 | -6 | 4.2252 | TRUE | 12.668 | 8.856 | -1.8371 | -0.30663 | 0.96723 | -1.282 | -0.13306 | 1.2739
P09668 | Pro-cathepsin H | CTSH | 1:1 | Block_4 | -2.6968 | 3.9074 | FALSE | 8.3135 | 10.6509 | -1.6715 | -0.6145 | 0.17284 | -1.5574 | -0.31246 | 0.85622
P43234 | Cathepsin O | CTSO | 1:1 | Block_4 | -3.5916 | 3.5948 | FALSE | 4.6621 | 6.2916 | -0.24578 | 0.33232 | 0.90986 | -0.014292 | 0.48295 | 1.2031
O60911 | Cathepsin L2 | CTSV | 1:1 | Block_4 | -5 | 3.0397 | TRUE | 7.6387 | 8.1142 | -0.91347 | -0.28614 | 0.70697 | -1.4251 | -0.40641 | 0.54338
Q9P013 | Spliceosome-associated protein CWC15 homolog | CWC15 | 1:1 | Block_4 | 0.86782 | 8.9168 | TRUE | NA | NA | NA | NA | NA | NA | 1.0596 | 2.6868
P78423 | Fractalkine | CX3CL1 | 1:1 | Block_4 | -3 | 3.7288 | FALSE | 13.9332 | 14.6395 | -0.63647 | -0.1505 | 0.54643 | -0.44259 | 0.52804 | 1.611
P78310 | Coxsackievirus and adenovirus receptor | CXADR | 1:1 | Block_4 | -3.3664 | 5.1638 | TRUE | 10.6837 | 9.0454 | -0.85559 | -0.1027 | 0.71912 | -0.33788 | 0.57578 | 1.9498
Q6UXB2 | C-X-C motif chemokine 17 | CXCL17 | 1:1 | Block_4 | -1.5879 | 9.6243 | TRUE | 8.3362 | 6.3652 | -0.067985 | 2.0213 | 3.0547 | 0.79144 | 2.0213 | 3.1922
P10145 | Interleukin-8 | CXCL8 | 1:1 | Block_4 | -2.0843 | 8.4236 | TRUE | 12.3976 | 5.6841 | -1.3195 | -0.40403 | 0.98215 | 0.0086402 | 0.94789 | 2.6536
Q6BCY4 | NADH-cytochrome b5 reductase 2 | CYB5R2 | 1:1 | Block_4 | -2 | 7.406 | TRUE | 18.0003 | 7.876 | -0.4114 | 0.32895 | 1.6064 | 0.92256 | 1.8774 | 3.419
P98082 | Disabled homolog 2 | DAB2 | 1:1 | Block_4 | -0.29402 | 8.9844 | TRUE | 20.5354 | 12.4108 | -0.28631 | 1.092 | 6.3796 | 2.9827 | 6.1306 | 7.563
Q9UIK4 | Death-associated protein kinase 2 | DAPK2 | 1:1 | Block_4 | -0.6496 | 7.9506 | FALSE | 11.0505 | 7.4028 | NA | -0.15169 | 1.2345 | -0.22748 | 1.4062 | 6.1236
Q9UN19 | Dual adapter for phosphotyrosine and 3-phosphotyrosine and 3-phosphoinositide | DAPP1 | 1:1 | Block_4 | 0.61351 | 12.4058 | FALSE | NA | NA | NA | 0.66659 | 4.6391 | NA | 4.3515 | 5.98
P14868 | Aspartate--tRNA ligase, cytoplasmic | DARS1 | 1:1 | Block_4 | 0.53433 | 6.9228 | FALSE | NA | NA | NA | 0.79936 | 2.7251 | 0.56085 | 2.8097 | 4.1273
Q16643 | Drebrin | DBN1 | 1:1 | Block_4 | -1.7597 | 6.1846 | TRUE | 15.0091 | 10.9209 | -0.95587 | 0.25739 | 1.6247 | 0.6413 | 1.7532 | 2.5994
P07585 | Decorin | DCN | 1:1 | Block_4 | -3.5885 | 3.2679 | TRUE | 6.8669 | 5.61 | -0.076808 | 0.32646 | 0.68536 | 0.17666 | 0.62216 | 1.1553
Q96C86 | m7GpppX diphosphatase | DCPS | 1:1 | Block_4 | -4 | 5.4991 | FALSE | 17.3022 | 17.3056 | -0.61836 | 0.16336 | 0.63038 | -0.15299 | 0.5924 | 1.3392
Q14203 | Dynactin subunit 1 | DCTN1 | 1:1 | Block_4 | -1.6503 | 6.8879 | FALSE | 13.3978 | 6.5041 | 0.15306 | 2.2205 | 4.0783 | 2.4386 | 4.0261 | 5.3493
Q13561 | Dynactin subunit 2 | DCTN2 | 1:1 | Block_4 | 0.22192 | 6.7844 | TRUE | NA | 2.9871 | NA | 0.69487 | 2.259 | 0.34789 | 1.5185 | 3.1697
O00399 | Dynactin subunit 6 | DCTN6 | 1:1 | Block_4 | 0.60304 | 8.3376 | FALSE | NA | NA | NA | NA | 1.7992 | NA | 2.2148 | 4.0769
O94760 | N(G),N(G)-dimethylarginine dimethylaminohydrolase 1 | DDAH1 | 1:1 | Block_4 | 0.093368 | 5.3337 | TRUE | 14.3224 | 4.0216 | NA | 0.23427 | 1.1127 | 0.45126 | 1.1655 | 2.3047
O94830 | Phospholipase DDHD2 | DDHD2 | 1:1 | Block_4 | -0.91832 | 7.9274 | TRUE | 13.6213 | 12.586 | -0.61542 | 0.6721 | 2.0548 | 0.44786 | 2.1602 | 3.8672
Q92499 | ATP-dependent RNA helicase DDX1 | DDX1 | 1:1 | Block_4 | 1 | NA | NA | NA | NA | NA | NA | 1.6593 | NA | 1.987 | 3.8867
Q16698 | 2,4-dienoyl-CoA reductase [(3E)-enoyl-CoA-producing], mitochondrial | DECR1 | 1:1 | Block_4 | -0.6979 | 7.4321 | FALSE | 12.7295 | 5.5425 | -0.39276 | 1.479 | 4.3657 | 1.322 | 3.6076 | 5.4987
P81534 | Beta-defensin 103 | DEFB103A_DEFB103B | 1:1 | Block_4 | 0.5 | 8.3503 | FALSE | NA | NA | NA | NA | 1.3634 | NA | NA | 1.8141
O15263 | Defensin beta 4A | DEFB4A_DEFB4B | 1:1 | Block_4 | -5.7224 | 1.0795 | TRUE | 11.7632 | 17.5575 | -2.6532 | -0.73895 | 1.3664 | -1.9142 | 0.58797 | 2.0887
O00273 | DNA fragmentation factor subunit alpha | DFFA | 1:1 | Block_4 | -1.1896 | 7.1805 | TRUE | 14.7557 | 9.0669 | -0.3392 | 1.184 | 3.2669 | 1.6789 | 3.5683 | 5.1144
P23743 | Diacylglycerol kinase alpha | DGKA | 1:1 | Block_4 | 0.87547 | 5.3338 | FALSE | NA | NA | NA | NA | 1.5371 | NA | 1.3615 | 2.7483
O75912 | Diacylglycerol kinase iota | DGKI | 1:1 | Block_4 | -2.5447 | 9.4787 | TRUE | 12.7755 | 13.1817 | NA | -1.3375 | 2.9521 | -0.80349 | 2.3587 | 4.6281
P49366 | Deoxyhypusine synthase | DHPS | 1:1 | Block_4 | -2.644 | 6.6579 | TRUE | 15.9899 | 22.0576 | -2.4964 | -1.6583 | -1.0692 | -1.5483 | -0.73575 | 1.039
Q6PKH6 | Dehydrogenase/reductase SDR family member 4-like 2 | DHRS4L2 | 1:1 | Block_4 | -2.4663 | 9.4231 | FALSE | 12.7164 | 9.779 | -2.0468 | -0.14502 | 2.1139 | -0.33463 | 2.0946 | 4.0802
Q9UBT3 | Dickkopf-related protein 4 | DKK4 | 1:1 | Block_4 | -3.617 | 7.2694 | TRUE | 7.1541 | 6.9292 | -0.26387 | 0.41257 | 1.3887 | 0.10487 | 0.96921 | 2.2268
Q9UK85 | Dickkopf-like protein 1 | DKKL1 | 1:1 | Block_4 | -1.2715 | 6.2273 | FALSE | 7.9472 | 12.3182 | -0.61072 | 0.68473 | 1.7997 | -0.65659 | 0.5953 | 1.8332
P78352 | Disks large homolog 4 | DLG4 | 1:1 | Block_4 | 0.71459 | 10.9519 | TRUE | NA | NA | NA | NA | 2.9267 | NA | 2.3582 | 4.695
Q9NR61 | Delta-like protein 4 | DLL4 | 1:1 | Block_4 | 0.001 | 6.4669 | FALSE | 33.9742 | 55.3707 | NA | 0.021586 | 2.0489 | NA | 0.58851 | 1.9354
P11532 | Dystrophin | DMD | 1:1 | Block_4 | 0.35047 | 8.1377 | FALSE | 4.8318 | 5.5243 | NA | NA | 1.7484 | 0.519 | 1.6899 | 3.8308
Q13316 | Dentin matrix acidic phosphoprotein 1 | DMP1 | 1:1 | Block_4 | -2.2972 | 4.1277 | FALSE | 9.8721 | 5.3601 | -0.051119 | 1.5062 | 2.4467 | -1.2865 | 1.424 | 2.7213
O60884 | DnaJ homolog subfamily A member 2 | DNAJA2 | 1:1 | Block_4 | -0.58918 | 8.3937 | FALSE | 21.9819 | 16.0761 | -0.46052 | 1.397 | 4.5923 | 1.7573 | 4.5594 | 5.7126
Q8WW22 | DnaJ homolog subfamily A member 4 | DNAJA4 | 1:1 | Block_4 | -1.5294 | 6.295 | FALSE | 11.4529 | 5.9403 | 0.90421 | 1.733 | 2.7608 | 1.1922 | 2.1528 | 3.6385
P25685 | DnaJ homolog subfamily B member 1 | DNAJB1 | 1:1 | Block_4 | 0.7875 | 6.1355 | FALSE | NA | NA | NA | 2.8722 | 4.9005 | 1.6192 | 4.6332 | 6.2096
Q8TBM8 | DnaJ homolog subfamily B member 14 | DNAJB14 | 1:1 | Block_4 | 0.74473 | 11.4519 | TRUE | NA | NA | NA | NA | 1.8353 | NA | 2.2585 | 4.1791
Q9UDY4 | DnaJ homolog subfamily B member 4 | DNAJB4 | 1:1 | Block_4 | -0.13452 | 8.5656 | FALSE | 11.4556 | 7.1937 | -0.046475 | 0.51409 | 1.0481 | 0.70712 | 1.3643 | 2.3307
Q8NHS0 | DnaJ homolog subfamily B member 8 | DNAJB8 | 1:1 | Block_4 | 0.10972 | 9.3664 | FALSE | 0.49819 | 3.2551 | NA | NA | 0.83281 | NA | 0.50969 | 1.2244
Q9NVM6 | DnaJ homolog subfamily C member 17 | DNAJC17 | 1:1 | Block_4 | 0.49139 | 8.8217 | TRUE | NA | NA | NA | 0.99828 | 2.0433 | 0.75885 | 2.1916 | 3.696
Q5F1R6 | DnaJ homolog subfamily C member 21 | DNAJC21 | 1:1 | Block_4 | 0.60794 | 6.412 | FALSE | NA | NA | NA | NA | 1.09 | NA | 1.0094 | 2.2064
O75061 | Putative tyrosine-protein phosphatase auxilin | DNAJC6 | 1:1 | Block_4 | -1.2589 | 10.6324 | TRUE | 19.2528 | 8.143 | 0.10116 | 2.2807 | 5.4075 | 2.8684 | 5.1855 | 6.9371
Q8WXX5 | DnaJ homolog subfamily C member 9 | DNAJC9 | 1:1 | Block_4 | 0.54301 | 7.2561 | TRUE | NA | 9.7591 | NA | 1.3077 | 2.8146 | 1.3271 | 2.884 | 4.5258
Q05193 | Dynamin-1 | DNM1 | 1:1 | Block_4 | 0.18 | 7.9238 | FALSE | 15.6212 | 11.8739 | 0.29334 | 2.145 | 6.4281 | 1.2859 | 5.9683 | 7.912
Q9UQ16 | Dynamin-3 | DNM3 | 1:1 | Block_4 | -0.53702 | 9.3166 | TRUE | 8.7705 | 10.9144 | -0.19318 | 0.59323 | 2.0463 | 0.88045 | 1.8409 | 3.4317
Q6XZF7 | Dynamin-binding protein | DNMBP | 1:1 | Block_4 | -3.3382 | 7.8568 | TRUE | 7.9827 | 5.835 | -0.80289 | 0.37091 | 1.627 | 0.76226 | 1.958 | 4.0949
Q9ULA0 | Aspartyl aminopeptidase | DNPEP | 1:1 | Block_4 | -1.1543 | 7.3704 | FALSE | 13.8844 | 9.1767 | 0.29275 | 1.38 | 2.1326 | 1.1389 | 2.1015 | 3.0567
Q99704 | Docking protein 1 | DOK1 | 1:1 | Block_4 | 0.59441 | 8.7658 | TRUE | NA | NA | NA | 1.0253 | 5.9755 | 0.74413 | 5.7526 | 7.3538
P16444 | Dipeptidase 1 | DPEP1 | 1:1 | Block_4 | -4.0161 | 2.8119 | TRUE | 7.6841 | 7.6883 | -1.1229 | -0.013748 | 0.45135 | -0.86032 | -0.030783 | 0.83243
Q9H4A9 | Dipeptidase 2 | DPEP2 | 1:1 | Block_4 | -6 | 1.7051 | TRUE | 10.3513 | 7.7671 | -1.101 | -0.57136 | -0.10077 | -0.98098 | -0.36208 | 0.55309
Q8N608 | Inactive dipeptidyl peptidase 10 | DPP10 | 1:1 | Block_4 | -1.5367 | 8.552 | TRUE | 10.5744 | 3.8648 | -0.7105 | -0.21686 | 0.32455 | -0.67317 | -0.17939 | 0.83367
P42658 | Dipeptidyl aminopeptidase-like protein 6 | DPP6 | 1:1 | Block_4 | -2.3501 | 6.9269 | FALSE | 6.5975 | 8.8273 | -0.50827 | 0.16134 | 0.72112 | -0.51263 | 0.27951 | 1.0956
Q9C005 | Protein dpy-30 homolog | DPY30 | 1:1 | Block_4 | -3.2582 | 5.9621 | TRUE | 7.8344 | 4.6822 | -0.77809 | 0.1991 | 0.87589 | 0.90155 | 1.8455 | 3.2048
Q8NBI3 | Draxin | DRAXIN | 1:1 | Block_4 | -2.1073 | 7.6764 | TRUE | 20.7979 | 9.8601 | 0.030982 | 1.0542 | 1.8186 | 0.39366 | 1.6757 | 2.7897
P55039 | Developmentally-regulated GTP-binding protein 2 | DRG2 | 1:1 | Block_4 | 0.20169 | 7.7895 | TRUE | 2.0605 | 2.5841 | NA | 0.39466 | 2.1165 | 0.49429 | 2.3401 | 3.4588
P32926 | Desmoglein-3 | DSG3 | 1:1 | Block_4 | -6 | 5.5146 | TRUE | 8.541 | 5.7467 | -0.065898 | 0.33982 | 1.32 | 0.18483 | 0.99477 | 1.7027
Q86SJ6 | Desmoglein-4 | DSG4 | 1:1 | Block_4 | -1.8657 | 7.7119 | TRUE | 11.4418 | 10.236 | -0.69669 | 0.084834 | 1.0205 | -1.1616 | 0.012668 | 1.0474
O60941 | Dystrobrevin beta | DTNB | 1:1 | Block_4 | -2.2 | 10.2217 | TRUE | 14.6199 | 6.2295 | -0.68846 | 0.19482 | 1.2255 | 0.55456 | 1.5523 | 2.9178
Q8N9I9 | Probable E3 ubiquitin-protein ligase DTX3 | DTX3 | 1:1 | Block_4 | -0.88303 | 5.4321 | FALSE | 10.7833 | 3.0492 | -0.37939 | 0.22523 | 0.59923 | 0.3223 | 0.895 | 1.7256
P33316 | Deoxyuridine 5'-triphosphate nucleotidohydrolase, mitochondrial | DUT | 1:1 | Block_4 | 0.5 | 10.1538 | FALSE | 22.0653 | 17.5014 | NA | 2.6916 | 5.7693 | 1.9452 | 4.7339 | 6.7641
O77932 | Decapping and exoribonuclease protein | DXO | 1:1 | Block_4 | 0.57857 | 6.5873 | TRUE | NA | NA | NA | 1.0262 | 2.3011 | 1.09 | 2.7969 | 4.4046
P51808 | Dynein light chain Tctex-type 3 | DYNLT3 | 1:1 | Block_4 | 1 | 4.6019 | TRUE | NA | 3.1847 | NA | NA | NA | NA | NA | 1.8961
O00559 | Receptor-binding cancer antigen expressed on SiSo cells | EBAG9 | 1:1 | Block_4 | -0.56813 | 8.7579 | TRUE | 11.0331 | 7.9432 | -0.40812 | 0.8994 | 4.5029 | 1.7633 | 4.4111 | 5.8589
Q14213_Q8NEV9 | Interleukin-27 | EBI3_IL27 | 1:1 | Block_4 | -3.6576 | 4.7731 | FALSE | 7.0803 | 5.5932 | -0.55598 | -0.10336 | 0.42337 | -0.31965 | 0.34643 | 1.3067
P42892 | Endothelin-converting enzyme 1 | ECE1 | 1:1 | Block_4 | -2.94 | 2.0061 | TRUE | 15.766 | 9.4512 | -1.2102 | -0.48636 | -0.041195 | -0.97782 | -0.34089 | 0.42851
P30084 | Enoyl-CoA hydratase, mitochondrial | ECHS1 | 1:1 | Block_4 | -2.1956 | 6.7448 | TRUE | 21.088 | 7.4994 | -1.1455 | 0.54608 | 4.0097 | 0.083113 | 2.2441 | 3.7453
Q9HAV5 | Tumor necrosis factor receptor superfamily member 27 | EDA2R | 1:1 | Block_4 | -3.4861 | 7.1034 | TRUE | 7.2628 | 3.4048 | -0.50562 | 0.13432 | 0.86779 | 0.12644 | 1.5421 | 3.0148
Q9UNE0 | Tumor necrosis factor receptor superfamily member EDAR | EDAR | 1:1 | Block_4 | 1.4405 | 8.2739 | TRUE | NA | 10.095 | NA | NA | 2.0283 | NA | 1.8794 | 3.2739
P56851 | Epididymal secretory protein E3-beta | EDDM3B | 1:1 | Block_4 | -0.52653 | 6.3888 | TRUE | 8.5623 | 8.8207 | NA | NA | 0.59146 | NA | NA | 1.057
O60869 | Endothelial differentiation-related factor 1 | EDF1 | 1:1 | Block_4 | 0.93896 | 9.2324 | FALSE | NA | NA | NA | NA | 2.7207 | 1.3001 | 3.8805 | 6.1509
O43854 | EGF-like repeat and discoidin I-like domain-containing protein 3 | EDIL3 | 1:1 | Block_4 | -2.0494 | 7.5266 | FALSE | 7.2442 | 10.5154 | -0.45987 | 0.3101 | 1.9672 | -1.2458 | -0.21753 | 1.0304
P05305 | Endothelin-1 | EDN1 | 1:1 | Block_4 | -1.3584 | NA | NA | 11.7043 | 6.641 | 0.29739 | 1.0739 | 1.7954 | 0.52361 | 1.4989 | 2.9867
Q9BUP0 | EF-hand domain-containing protein D1 | EFHD1 | 1:1 | Block_4 | -1.254 | 8.3704 | TRUE | 11.69 | 10.0265 | -0.5197 | -0.050909 | 0.57231 | -0.58606 | 0.55324 | 1.8779
Q96C19 | EF-hand domain-containing protein D2 | EFHD2 | 1:1 | Block_4 | -1.8657 | 7.14 | TRUE | 10.8899 | 9.1 | -1.0245 | 0.10102 | 0.95916 | 0.23547 | 1.2384 | 2.3909
P52798 | Ephrin-A4 | EFNA4 | 1:1 | Block_4 | -3.6447 | 5.6089 | TRUE | 9.3234 | 4.9828 | -0.46746 | 0.22045 | 0.67084 | 0.21222 | 1.0424 | 2.2559
Q63HQ2 | Pikachurin | EGFLAM | 1:1 | Block_4 | -2 | 8.0301 | FALSE | 14.3411 | 10.5055 | -0.17157 | 0.57899 | 1.1246 | 0.23434 | 0.85914 | 1.6881
Q9GZT9 | Egl nine homolog 1 | EGLN1 | 1:1 | Block_4 | -1.1416 | 9.682 | FALSE | 16.8067 | 8.2228 | -0.95254 | 0.056048 | 1.0937 | -0.30552 | 1.1123 | 5.3474
Q9NZN3 | EH domain-containing protein 3 | EHD3 | 1:1 | Block_4 | 1 | 7.5663 | FALSE | 22.5585 | 5.4916 | NA | 2.3435 | 5.9914 | 1.9422 | 5.7003 | 7.4139
P47813 | Eukaryotic translation initiation factor 1A, X-chromosomal | EIF1AX | 1:1 | Block_4 | 2.5 | 9.6352 | TRUE | NA | NA | NA | NA | 3.086 | NA | 3.2738 | 5.545
P19525 | Interferon-induced, double-stranded RNA-activated protein kinase | EIF2AK2 | 1:1 | Block_4 | -2.0159 | NA | NA | 17.0701 | 10.8485 | NA | 0.1848 | 2.3726 | -0.39054 | 3.2704 | 6.1436
Q9UI10 | Translation initiation factor eIF-2B subunit delta | EIF2B4 | 1:1 | Block_4 | -1.2007 | 8.2394 | TRUE | 11.2354 | 10.0577 | -0.83063 | 0.63072 | 2.2634 | 1.0743 | 2.8833 | 4.4659
P20042 | Eukaryotic translation initiation factor 2 subunit 2 | EIF2S2 | 1:1 | Block_4 | 1.5242 | 6.4896 | FALSE | NA | NA | NA | NA | 3.3421 | 1.9571 | 3.6636 | 4.6567
P23588 | Eukaryotic translation initiation factor 4B | EIF4B | 1:1 | Block_4 | 1.6931 | 10.7167 | TRUE | NA | NA | NA | 2.4211 | 5.324 | 2.3237 | 5.2526 | 6.7785
P06730 | Eukaryotic translation initiation factor 4E | EIF4E | 1:1 | Block_4 | 3.1647 | 12.6952 | TRUE | NA | NA | NA | NA | 6.6449 | NA | 6.7432 | 8.7681
Q04637 | Eukaryotic translation initiation factor 4 gamma 1 | EIF4G1 | 1:1 | Block_4 | -1.4288 | 6.7128 | TRUE | 16.7376 | 19.0201 | -1.3173 | 1.5258 | 4.5401 | 2.658 | 5.0299 | 6.1101
O43432 | Eukaryotic translation initiation factor 4 gamma 3 | EIF4G3 | 1:1 | Block_4 | 1.5131 | 5.7388 | FALSE | NA | NA | NA | NA | 2.8753 | NA | 2.2495 | 5.0922
Q15056 | Eukaryotic translation initiation factor 4H | EIF4H | 1:1 | Block_4 | 1.0405 | 8.3516 | FALSE | 7.9903 | 1.5291 | NA | 2.2528 | 4.8593 | 2.0205 | 5.072 | 6.1104
Q9H777 | Zinc phosphodiesterase ELAC protein 1 | ELAC1 | 1:1 | Block_4 | -0.19547 | 9.0038 | TRUE | 16.3059 | 9.3674 | NA | 0.4284 | 3.0345 | 0.13086 | 2.8872 | 4.6449
Q14241 | Elongin-A | ELOA | 1:1 | Block_4 | -1.8698 | 8.8666 | TRUE | 16.2014 | 7.5467 | NA | -1.2643 | -0.35146 | -1.2603 | -0.38446 | 2.1817
Q9HC35 | Echinoderm microtubule-associated protein-like 4 | EML4 | 1:1 | Block_4 | -2.3079 | 8.9862 | FALSE | 9.7543 | 6.4905 | -1.5897 | 0.74998 | 3.096 | 0.90127 | 3.367 | 5.5549
P21128 | Uridylate-specific endoribonuclease | ENDOU | 1:1 | Block_4 | -1.7207 | 8.4958 | TRUE | 10.7632 | 4.6314 | -0.86332 | -0.22474 | 0.91628 | -0.56546 | 0.33825 | 1.4652
P13929 | Beta-enolase | ENO3 | 1:1 | Block_4 | -0.44822 | 3.1038 | FALSE | 10.2449 | 9.5182 | NA | -0.29565 | 1.165 | NA | -0.18291 | 0.81784
Q6UWR7 | Glycerophosphocholine cholinephosphodiesterase ENPP6 | ENPP6 | 1:1 | Block_4 | -4.5 | 5.7398 | TRUE | 9.9845 | 9.7327 | -0.31257 | 0.34779 | 0.78301 | -0.88189 | 0.18911 | 1.0777
O75356 | Nucleoside diphosphate phosphatase ENTPD5 | ENTPD5 | 1:1 | Block_4 | -3.7529 | 6.697 | FALSE | 7.5657 | 6.7749 | -0.46277 | -0.055255 | 0.23783 | -0.56009 | -0.162 | 0.34235
O75354 | Ectonucleoside triphosphate diphosphohydrolase 6 | ENTPD6 | 1:1 | Block_4 | -1.9974 | 4.7132 | FALSE | 26.9702 | 10.5692 | -0.29619 | 0.30376 | 0.82479 | -0.19946 | 0.39004 | 0.86879
Q96C92 | Endosome-associated-trafficking regulator 1 | ENTR1 | 1:1 | Block_4 | -1.1704 | 8.3143 | FALSE | 15.187 | 10.8904 | NA | -0.53279 | 0.35428 | -0.3683 | 0.69006 | 2.3316
O43491 | Band 4.1-like protein 2 | EPB41L2 | 1:1 | Block_4 | 0.27598 | 7.0928 | FALSE | 15.4201 | 11.5076 | NA | 0.41122 | 0.9575 | 0.33297 | 0.89668 | 1.6162
Q6UW88 | Epigen | EPGN | 1:1 | Block_4 | 0.50296 | 8.7337 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.85373
P29317 | Ephrin type-A receptor 2 | EPHA2 | 1:1 | Block_4 | -1.9853 | 7.4609 | TRUE | 10.2589 | 6.8327 | -0.44745 | 0.17025 | 0.66669 | 0.11486 | 0.86284 | 2.0485
O15197 | Ephrin type-B receptor 6 | EPHB6 | 1:1 | Block_4 | -4.8053 | 4.6979 | TRUE | 7.3299 | 6.8096 | -0.51476 | 0.18166 | 0.82192 | 0.074747 | 0.84709 | 1.7685
P01588 | Erythropoietin | EPO | 1:1 | Block_4 | -3.3757 | 6.1611 | FALSE | 12.629 | 10.1135 | -2.396 | -1.1009 | 0.35151 | -1.8851 | -0.39669 | 1.6953
P58107 | Epiplakin | EPPK1 | 1:1 | Block_4 | -0.9712 | 6.943 | FALSE | 24.0372 | 16.1935 | -0.36003 | 0.58824 | 1.5917 | 0.056743 | 1.1691 | 3.0485
Q9H6S3 | Epidermal growth factor receptor kinase substrate 8-like protein 2 | EPS8L2 | 1:1 | Block_4 | -2.7315 | 5.0359 | FALSE | 9.7903 | 3.3011 | -0.32394 | 0.39041 | 1.0746 | 0.079798 | 0.97873 | 2.0962
Q96RT1 | Erbin | ERBIN | 1:1 | Block_4 | 0.48881 | 9.4721 | TRUE | 9.4596 | 4.1414 | NA | 0.70062 | 3.7964 | 0.92583 | 3.9134 | 5.4425
O14944 | Proepiregulin | EREG | 1:1 | Block_4 | -0.92555 | 12.0842 | TRUE | 9.0886 | 4.377 | -0.48337 | 0.77084 | 2.3759 | 0.26668 | 2.9822 | 5.6503
O75460 | Serine/threonine-protein kinase/endoribonuclease IRE1 | ERN1 | 1:1 | Block_4 | -0.93234 | 9.6702 | FALSE | 10.8842 | 9.646 | -0.66983 | 0.066441 | 0.57132 | -0.39296 | 0.24106 | 0.87361
Q9BS26 | Endoplasmic reticulum resident protein 44 | ERP44 | 1:1 | Block_4 | -2.4936 | 4.4444 | TRUE | 10.4345 | 5.9615 | -0.17334 | 0.60976 | 1.1592 | 0.45297 | 0.99608 | 1.523
A0FGR8 | Extended synaptotagmin-2 | ESYT2 | 1:1 | Block_4 | 1 | 7.24 | FALSE | 2.6984 | 5.7075 | NA | NA | 2.5039 | NA | 2.6077 | 5.0607
P58658 | Protein eva-1 homolog C | EVA1C | 1:1 | Block_4 | -3.9797 | 7.4337 | FALSE | 6.2034 | 7.1461 | -0.45985 | 0.24905 | 0.75983 | 0.71045 | 1.3509 | 2.1323
O60447 | Ecotropic viral integration site 5 protein homolog | EVI5 | 1:1 | Block_4 | 0.28576 | 9.644 | TRUE | 7.2728 | 2.5546 | NA | 0.31128 | 1.9952 | NA | 2.2319 | 3.3075
P13726 | Tissue factor | F3 | 1:1 | Block_4 | -2.8673 | 5.3097 | TRUE | 17.2638 | 8.2484 | -0.72254 | 0.13136 | 0.8894 | -0.29756 | 0.46643 | 1.1976
P51161 | Gastrotropin | FABP6 | 1:1 | Block_4 | -0.38906 | 8.7132 | FALSE | 8.651 | 14.1972 | -0.25053 | 0.59963 | 1.901 | 0.10877 | 0.93104 | 2.496
Q9NRY5 | Protein FAM114A2 | FAM114A2 | 1:1 | Block_4 | 0.5 | 12.9409 | FALSE | NA | NA | NA | 1.1266 | 3.0533 | 1.0056 | 3.2082 | 4.5966
O94988 | Protein FAM13A | FAM13A | 1:1 | Block_4 | 0.5203 | 11.6585 | TRUE | NA | NA | NA | NA | 2.2022 | NA | 1.9311 | 3.6907
A8MVW0 | Protein FAM171A2 | FAM171A2 | 1:1 | Block_4 | -2.72 | 7.5684 | FALSE | 5.8288 | 4.4521 | -2.5206 | -0.79563 | 2.0638 | -2.6014 | -0.85835 | 2.8968
Q8WUF8 | Cotranscriptional regulator FAM172A | FAM172A | 1:1 | Block_4 | 0.001 | 9.1351 | FALSE | 22.3727 | 24.6916 | NA | 0.76989 | 2.7601 | 1.1674 | 2.9813 | 4.7123
P58499 | Protein FAM3B | FAM3B | 1:1 | Block_4 | -4.5616 | 5.1655 | TRUE | 6.7725 | 3.2678 | -0.50324 | 0.41036 | 0.91561 | 0.063804 | 0.84168 | 1.5929
Q96BQ1 | Protein FAM3D | FAM3D | 1:1 | Block_4 | -4 | 3.7614 | TRUE | 10.6998 | 7.4458 | -0.44012 | 0.32245 | 1.1502 | -0.17164 | 1.0782 | 2.1905
P48023 | Tumor necrosis factor ligand superfamily member 6 | FASLG | 1:1 | Block_4 | -4.1059 | 7.0761 | TRUE | 7.2694 | 8.156 | -0.28684 | 0.63526 | 1.2042 | -0.19743 | 0.55422 | 1.6323
P35556 | Fibrillin-2 | FBN2 | 1:1 | Block_4 | 0.74546 | 8.2007 | TRUE | NA | NA | NA | NA | 1.4253 | NA | NA | 1.6592
P24071 | Immunoglobulin alpha Fc receptor | FCAR | 1:1 | Block_4 | -1.7846 | 9.3668 | TRUE | 9.9885 | 6.3086 | -0.62236 | 0.25795 | 0.65596 | -0.3942 | 0.71287 | 1.9999
Q96LA5 | Fc receptor-like protein 2 | FCRL2 | 1:1 | Block_4 | -3.5112 | 5.8879 | TRUE | 10.0037 | 6.574 | -1.2245 | 0.44056 | 0.87619 | -0.30556 | 0.61468 | 1.4718
Q6DN72 | Fc receptor-like protein 6 | FCRL6 | 1:1 | Block_4 | -3.105 | 7.45 | FALSE | 9.7253 | 11.6504 | -0.53178 | 0.81462 | 1.8981 | -0.94769 | 0.69019 | 1.943
Q5JSP0 | FYVE, RhoGEF and PH domain-containing protein 3 | FGD3 | 1:1 | Block_4 | -0.19487 | 9.6381 | TRUE | 12.3657 | 12.0263 | NA | 0.56409 | 3.5569 | 0.34152 | 3.606 | 5.4861
O43320 | Fibroblast growth factor 16 | FGF16 | 1:1 | Block_4 | 1.2164 | 8.4484 | FALSE | 9.6789 | 20.4875 | NA | NA | NA | NA | NA | NA
P09038 | Fibroblast growth factor 2 | FGF2 | 1:1 | Block_4 | 1.3795 | 8.7204 | TRUE | NA | NA | NA | NA | 1.7066 | NA | NA | 2.5585
Q9NSA1 | Fibroblast growth factor 21 | FGF21 | 1:1 | Block_4 | -4.5568 | 5.0134 | TRUE | 6.9494 | 8.7592 | -3.0187 | -1.4527 | 0.24873 | -1.6852 | 0.54843 | 3.0258
Q9GZV9 | Fibroblast growth factor 23 | FGF23 | 1:1 | Block_4 | -5.061 | 6.8466 | TRUE | 10.2939 | 8.2915 | -2.5486 | -1.7102 | -0.096641 | -2.4934 | -1.3826 | 1.2754
P11487 | Fibroblast growth factor 3 | FGF3 | 1:1 | Block_4 | 1.0376 | 3.9654 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TAT2 | Fibroblast growth factor-binding protein 3 | FGFBP3 | 1:1 | Block_4 | -1.88 | 7.1616 | TRUE | 11.0518 | 8.1636 | -0.050809 | 0.65097 | 1.0819 | 0.52201 | 1.0999 | 1.8702
Q5W0V3 | FHF complex subunit HOOK interacting protein 2A | FHIP2A | 1:1 | Block_4 | -1.245 | 3.7596 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NWM8 | Peptidyl-prolyl cis-trans isomerase FKBP14 | FKBP14 | 1:1 | Block_4 | -2.13 | 6.6035 | TRUE | 24.9225 | 20.2889 | -1.3019 | 0.022854 | 2.8846 | 0.55827 | 2.8791 | 4.4421
P68106 | Peptidyl-prolyl cis-trans isomerase FKBP1B | FKBP1B | 1:1 | Block_4 | 2.43 | 9.7157 | FALSE | NA | NA | NA | NA | 5.657 | NA | 5.9955 | 8.3027
Q02790 | Peptidyl-prolyl cis-trans isomerase FKBP4 | FKBP4 | 1:1 | Block_4 | -2.0757 | 4.6018 | TRUE | 13.2934 | 6.6168 | -0.32753 | 0.86583 | 2.4325 | 1.5414 | 2.6738 | 3.3393
Q13451 | Peptidyl-prolyl cis-trans isomerase FKBP5 | FKBP5 | 1:1 | Block_4 | -0.57157 | 7.0238 | TRUE | 9.6406 | 7.2312 | NA | 1.0584 | 3.4686 | 1.6259 | 3.6156 | 4.9869
Q9UIM3 | FK506-binding protein-like | FKBPL | 1:1 | Block_4 | -0.51713 | 9.9692 | FALSE | 8.5024 | 12.5108 | -0.40003 | 1.3092 | 3.0015 | 1.6697 | 3.511 | 4.6656
O43155 | Leucine-rich repeat transmembrane protein FLRT2 | FLRT2 | 1:1 | Block_4 | -3 | 6.2687 | TRUE | 11.5178 | 11.5362 | -0.46503 | 0.022765 | 0.46842 | -0.35039 | 0.3343 | 1.14
P17948 | Vascular endothelial growth factor receptor 1 | FLT1 | 1:1 | Block_4 | -2.1551 | 4.3744 | TRUE | 9.6724 | 9.7181 | -0.30498 | 0.21692 | 0.90441 | 0.0048114 | 0.60039 | 1.1902
P49771 | Fms-related tyrosine kinase 3 ligand | FLT3LG | 1:1 | Block_4 | -5.872 | 4.7688 | TRUE | 6.4652 | 5.5511 | -0.36345 | 0.66207 | 1.3663 | -0.049872 | 0.94427 | 1.9671
O95466 | Formin-like protein 1 | FMNL1 | 1:1 | Block_4 | -3.006 | 1.5234 | FALSE | 12.2826 | 9.0104 | -1.8244 | -1.0155 | -0.34314 | -1.2055 | -0.29623 | 1.4658
Q8N3X1 | Formin-binding protein 4 | FNBP4 | 1:1 | Block_4 | -0.10565 | 8.6499 | TRUE | 8.2565 | 11.6435 | NA | 0.058404 | 0.95402 | 0.20267 | 1.0087 | 3.0428
Q4ZHG4 | Fibronectin type III domain-containing protein 1 | FNDC1 | 1:1 | Block_4 | -1.9064 | 5.5581 | TRUE | 14.5729 | 6.282 | -0.51419 | -0.0082611 | 0.64186 | -0.30811 | 0.52772 | 1.4394
P49354 | Protein farnesyltransferase/geranylgeranyltransferase type-1 subunit alpha | FNTA | 1:1 | Block_4 | 0.082902 | 10.783 | TRUE | 24.2021 | 13.0884 | 0.24692 | 1.2109 | 2.1022 | 0.90926 | 2.0711 | 3.2957
Q12778 | Forkhead box protein O1 | FOXO1 | 1:1 | Block_4 | -0.39366 | 7.9947 | FALSE | 10.3128 | 12.2879 | NA | 0.058629 | 2.859 | 0.099397 | 3.2331 | 4.9228
Q92765 | Secreted frizzled-related protein 3 | FRZB | 1:1 | Block_4 | -3.6532 | 5.7696 | TRUE | 5.3838 | 5.0771 | 0.14905 | 0.86929 | 1.4304 | 0.43198 | 1.2223 | 1.6639
P01225 | Follitropin subunit beta | FSHB | 1:1 | Block_4 | 0.61312 | 6.1834 | TRUE | NA | 1.1886 | NA | NA | 1.2902 | NA | NA | 1.7363
Q12841 | Follistatin-related protein 1 | FSTL1 | 1:1 | Block_4 | -1.663 | 2.326 | FALSE | 15.5298 | 4.7347 | -0.13634 | 0.24907 | 0.50473 | 0.19883 | 0.57648 | 0.96523
O95954 | Formimidoyltransferase-cyclodeaminase | FTCD | 1:1 | Block_4 | -1.4955 | 6.6682 | TRUE | 15.5273 | 14.6601 | NA | -0.77357 | 0.47607 | -1.4669 | -0.13844 | 2.1589
Q9BYC5 | Alpha-(1,6)-fucosyltransferase | FUT8 | 1:1 | Block_4 | -6.6151 | 4.2496 | FALSE | 14.8501 | 4.3983 | -0.61846 | 0.23568 | 1.3545 | -0.093637 | 1.2185 | 2.8108
Q16595 | Frataxin, mitochondrial | FXN | 1:1 | Block_4 | -2.8547 | 6.5862 | FALSE | 9.0508 | 7.98 | 0.23926 | 2.2308 | 4.6715 | 3.2098 | 4.7765 | 5.2967
O95166 | Gamma-aminobutyric acid receptor-associated protein | GABARAP | 1:1 | Block_4 | -0.13333 | 9.5939 | TRUE | 13.7413 | 3.644 | NA | NA | 0.39884 | NA | 0.013417 | 1.101
Q8TAE8 | Growth arrest and DNA damage-inducible proteins-interacting protein 1 | GADD45GIP1 | 1:1 | Block_4 | 0.43084 | 6.2606 | FALSE | NA | 4.3196 | NA | 0.84731 | 3.4892 | 1.1101 | 2.6619 | 4.3473
Q86SR1 | Polypeptide N-acetylgalactosaminyltransferase 10 | GALNT10 | 1:1 | Block_4 | -3.3904 | 3.8381 | FALSE | 7.2153 | 6.5989 | -0.3526 | 0.29571 | 0.90695 | 0.16138 | 0.68749 | 1.31
Q10471 | Polypeptide N-acetylgalactosaminyltransferase 2 | GALNT2 | 1:1 | Block_4 | -4.3316 | 3.3596 | TRUE | 6.9114 | 6.4066 | -0.16525 | 0.30261 | 0.57788 | 0.0029669 | 0.45175 | 0.96298
Q86SF2 | N-acetylgalactosaminyltransferase 7 | GALNT7 | 1:1 | Block_4 | -3.5024 | 2.0984 | TRUE | 11.2833 | 11.0602 | -0.27421 | 0.26191 | 0.64473 | -0.4056 | 0.20839 | 0.78514
P22102 | Trifunctional purine biosynthetic protein adenosine-3 | GART | 1:1 | Block_4 | 1.4211 | 7.1888 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P01350 | Gastrin | GAST | 1:1 | Block_4 | -0.37521 | 6.1661 | TRUE | 9.5501 | 2.8767 | -0.36644 | 1.1075 | 3.597 | NA | 1.5242 | 5.172
P0DPI2 | Glutamine amidotransferase-like class 1 domain-containing protein 3, mitochondrial | GATD3 | 1:1 | Block_4 | -0.56292 | 6.4593 | FALSE | 2.1256 | 7.0926 | NA | 1.3058 | 2.8714 | 1.4269 | 3.0753 | 4.8759
P04062 | Lysosomal acid glucosylceramidase | GBA1 | 1:1 | Block_4 | 0.3286 | 4.9303 | FALSE | 12.3724 | 6.6035 | NA | 0.33308 | 1.3084 | NA | 0.76942 | 1.9947
P32455 | Guanylate-binding protein 1 | GBP1 | 1:1 | Block_4 | -5 | NA | NA | 15.025 | 5.966 | -1.4626 | -0.487 | 0.41472 | -0.73675 | 0.57672 | 2.7503
Q96CN9 | GRIP and coiled-coil domain-containing protein 1 | GCC1 | 1:1 | Block_4 | -0.5 | 8.7258 | FALSE | 19.497 | 17.0674 | NA | 1.3221 | 4.9416 | 1.7688 | 4.9987 | 6.7278
Q02742 | Beta-1,3-galactosyl-O-glycosyl-glycoprotein beta-1,6-N-acetylglucosaminyltransferase | GCNT1 | 1:1 | Block_4 | -3.4368 | 6.3372 | FALSE | 12.529 | 6.891 | -0.48866 | 0.22843 | 0.80257 | 0.30025 | 0.84671 | 1.5821
P56159 | GDNF family receptor alpha-1 | GFRA1 | 1:1 | Block_4 | -4.8905 | 5.5222 | TRUE | 7.5758 | 7.7175 | -0.75961 | 0.11845 | 0.74451 | 0.21465 | 1.0817 | 1.9561
O00451 | GDNF family receptor alpha-2 | GFRA2 | 1:1 | Block_4 | -1.0581 | 2.7268 | TRUE | 17.5677 | 8.4389 | -0.34915 | -0.037482 | 0.41088 | -0.24514 | 0.28301 | 0.72647
O60609 | GDNF family receptor alpha-3 | GFRA3 | 1:1 | Block_4 | -3.2875 | 3.2661 | TRUE | 7.0185 | 10.1513 | -0.53678 | 0.23064 | 0.68784 | -0.28747 | 0.29543 | 1.005
Q6UXV0 | GDNF family receptor alpha-like | GFRAL | 1:1 | Block_4 | -0.55962 | 9.8125 | TRUE | 10.6243 | 5.7426 | NA | 0.38081 | 1.0832 | -0.45336 | 0.42037 | 1.3781
P36269 | Glutathione hydrolase 5 proenzyme | GGT5 | 1:1 | Block_4 | -1.7161 | 3.1682 | FALSE | 12.4595 | 5.9509 | 0.0058669 | 0.38529 | 0.88587 | 0.2233 | 0.80402 | 1.6903
Q9NWU2 | Glucose-induced degradation protein 8 homolog | GID8 | 1:1 | Block_4 | -1.8114 | 6.5741 | TRUE | NA | NA | NA | -1.8093 | -0.70484 | NA | -1.2704 | -0.22044
Q6Y7W6 | GRB10-interacting GYF protein 2 | GIGYF2 | 1:1 | Block_4 | -0.88271 | 8.87 | TRUE | 9.9442 | 9.7593 | NA | -0.0041459 | 2.0846 | 0.51382 | 1.9095 | 3.8518
Q8TF65 | PDZ domain-containing protein GIPC2 | GIPC2 | 1:1 | Block_4 | 0.40028 | 7.2466 | FALSE | NA | NA | NA | NA | 0.84652 | NA | 0.75466 | 1.7031
Q8TF64 | PDZ domain-containing protein GIPC3 | GIPC3 | 1:1 | Block_4 | 0.001 | 8.9703 | TRUE | 15.435 | 7.3018 | NA | 1.6263 | 5.3763 | 1.9044 | 5.0651 | 6.5012
P48546 | Gastric inhibitory polypeptide receptor | GIPR | 1:1 | Block_4 | 1.1514 | 9.6379 | TRUE | NA | NA | NA | NA | 1.8684 | NA | NA | 1.4563
Q9Y2X7 | ARF GTPase-activating protein GIT1 | GIT1 | 1:1 | Block_4 | 0.41333 | 8.4064 | TRUE | 19.2341 | 6.549 | NA | 0.91515 | 3.7762 | 1.078 | 3.4325 | 5.1738
Q9NS71 | Gastrokine-1 | GKN1 | 1:1 | Block_4 | -1.6341 | 1.1964 | FALSE | 20.0189 | 6.8985 | -0.56378 | -0.0013604 | 0.36144 | -0.2317 | 0.11356 | 0.41773
P16278 | Beta-galactosidase | GLB1 | 1:1 | Block_4 | -0.43589 | 2.73 | FALSE | 13.8346 | 11.1273 | -0.034558 | 0.68701 | 1.9575 | 0.48802 | 1.4942 | 2.594
P48060 | Glioma pathogenesis-related protein 1 | GLIPR1 | 1:1 | Block_4 | -5 | 5.2692 | TRUE | 6.3237 | 6.8607 | -0.22389 | 0.23492 | 0.65897 | 0.02133 | 0.47064 | 0.87153
Q04760 | Lactoylglutathione lyase | GLO1 | 1:1 | Block_4 | -2.5079 | 5.9686 | FALSE | 8.1221 | 7.5702 | -0.033301 | 1.5269 | 2.5459 | 1.2112 | 2.6245 | 3.7383
Q9NS18 | Glutaredoxin-2, mitochondrial | GLRX2 | 1:1 | Block_4 | -0.99602 | 8.9643 | TRUE | 10.4426 | 8.7458 | -0.30382 | 0.79594 | 2.6664 | 1.3211 | 2.8846 | 4.2255
Q49A26 | Cytokine-like nuclear factor N-PAC | GLYR1 | 1:1 | Block_4 | -0.47208 | 0.67841 | FALSE | NA | NA | NA | -0.32856 | 0.14524 | NA | 0.39328 | 2.3135
P36959 | GMP reductase 1 | GMPR | 1:1 | Block_4 | 0.34486 | 5.8837 | FALSE | NA | NA | 0.52793 | 2.5187 | 4.9491 | 2.0001 | 4.6135 | 5.7065
O95467 | Neuroendocrine secretory protein 55 | GNAS | 1:1 | Block_4 | 0.41667 | 6.4992 | TRUE | NA | NA | NA | NA | 0.81764 | NA | 0.59792 | 2.7496
P46926 | Glucosamine-6-phosphate isomerase 1 | GNPDA1 | 1:1 | Block_4 | -0.13079 | 8.9535 | TRUE | 10.3321 | 4.5796 | -0.11634 | 0.43209 | 1.0027 | 0.19133 | 0.96132 | 2.0947
Q08378 | Golgin subfamily A member 3 | GOLGA3 | 1:1 | Block_4 | 1.4278 | 4.7891 | FALSE | NA | NA | NA | NA | 2.2131 | NA | 2.3769 | 4.1874
Q9HD26 | Golgi-associated PDZ and coiled-coil motif-containing protein | GOPC | 1:1 | Block_4 | 0.0189 | 8.9443 | TRUE | 15.6171 | 6.4259 | NA | 1.7232 | 5.5552 | 1.1514 | 5.2385 | 6.9644
Q9H8Y8 | Golgi reassembly-stacking protein 2 | GORASP2 | 1:1 | Block_4 | 0.66785 | 4.3108 | TRUE | NA | 3.8346 | NA | 0.9984 | 2.7828 | 1.17 | 2.8875 | 3.68
P13224 | Platelet glycoprotein Ib beta chain | GP1BB | 1:1 | Block_4 | 0.57545 | 5.6638 | FALSE | 2.4672 | 6.7479 | NA | 0.70278 | 1.9992 | 0.63228 | 2.2561 | 3.6141
P55259 | Pancreatic secretory granule membrane major glycoprotein GP2 | GP2 | 1:1 | Block_4 | -2.514 | 7.7477 | TRUE | 8.4472 | 6.1326 | -1.173 | 0.028482 | 0.78487 | -0.53948 | 0.59913 | 2.7123
Q99795 | Cell surface A33 antigen | GPA33 | 1:1 | Block_4 | -4.7382 | 3.3787 | TRUE | 8.1796 | 9.1766 | -2.6173 | -0.51956 | 1.1292 | -2.3535 | -0.83513 | 1.3161
P35052 | Glypican-1 | GPC1 | 1:1 | Block_4 | -4.0914 | 3.0998 | TRUE | 8.0572 | 4.1003 | -0.56475 | 0.045639 | 0.60446 | -0.065995 | 0.4358 | 1.1489
P78333 | Glypican-5 | GPC5 | 1:1 | Block_4 | -1.7855 | 7.1153 | TRUE | 13.3094 | 6.842 | -0.31327 | 0.64456 | 1.6587 | -0.41625 | 0.93505 | 2.0033
P21695 | Glycerol-3-phosphate dehydrogenase [NAD(+)], cytoplasmic | GPD1 | 1:1 | Block_4 | -3.2525 | 5.3503 | FALSE | 17.9761 | 4.4016 | -1.6572 | -0.98118 | -0.23811 | -1.7774 | -0.54694 | 0.72911
Q96T91 | Glycoprotein hormone alpha-2 | GPHA2 | 1:1 | Block_4 | -0.13994 | 10.573 | TRUE | 6.6028 | 8.0302 | NA | 0.3856 | 0.84262 | 0.038133 | 0.64679 | 1.3422
Q8IV16 | Glycosylphosphatidylinositol-anchored high density lipoprotein-binding protein 1 | GPIHBP1 | 1:1 | Block_4 | -2.3284 | 4.5617 | TRUE | 9.3266 | 3.8808 | -0.29618 | 0.23538 | 0.7015 | 0.097769 | 0.6746 | 1.3637
Q92917 | G-patch domain and KOW motifs-containing protein | GPKOW | 1:1 | Block_4 | 0.012727 | 5.5419 | FALSE | NA | 10.3157 | NA | NA | 0.075746 | NA | 0.18087 | 1.0268
Q5T848 | Probable G-protein coupled receptor 158 | GPR158 | 1:1 | Block_4 | 0.001 | 9.0391 | TRUE | 15.3776 | 18.4304 | 0.018836 | 0.86019 | 1.4502 | 0.2208 | 1.3283 | 2.5496
Q6UWK7 | Protein GPR15LG | GPR15LG | 1:1 | Block_4 | -3.0907 | 5.153 | TRUE | 7.7333 | 7.9239 | -0.45861 | 0.27678 | 0.92684 | 0.24568 | 1.2861 | 2.3828
O15354 | Prosaposin receptor GPR37 | GPR37 | 1:1 | Block_4 | -1.9918 | 6.5715 | TRUE | 12.0378 | 12.5045 | 0.29964 | 1.2225 | 2.2094 | 0.84533 | 2.1837 | 3.6693
P07492 | Gastrin-releasing peptide | GRP | 1:1 | Block_4 | -2.7328 | 0.53602 | FALSE | 8.4954 | 7.9801 | -0.59684 | 0.48747 | 1.1497 | -0.26105 | 0.69614 | 1.6723
Q9HAV7 | GrpE protein homolog 1, mitochondrial | GRPEL1 | 1:1 | Block_4 | -2.11 | 7.2783 | FALSE | 15.0076 | 3.9516 | -0.70176 | 0.39957 | 2.9535 | 0.78532 | 1.9784 | 3.7143
P0CG30 | Glutathione S-transferase theta-2B | GSTT2B | 1:1 | Block_4 | -3.7263 | 5.6942 | TRUE | 9.1806 | 6.6581 | -2.5368 | -0.042224 | 3.4907 | -2.1808 | 1.1148 | 5.1075
P13807 | Glycogen [starch] synthase, muscle | GYS1 | 1:1 | Block_4 | -1.0076 | 9.0382 | FALSE | 7.6622 | 9.0833 | NA | 0.047343 | 2.5421 | -0.65197 | 2.892 | 5.0229
P12544 | Granzyme A | GZMA | 1:1 | Block_4 | -3.6855 | 7.7722 | FALSE | 8.9166 | 6.8259 | -0.44853 | 0.42711 | 1.0655 | 0.030595 | 0.92077 | 1.9887
P10144 | Granzyme B | GZMB | 1:1 | Block_4 | 1.0418 | 8.6738 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.5002
P20718 | Granzyme H | GZMH | 1:1 | Block_4 | -4.6168 | 8.805 | TRUE | 10.8965 | 9.9357 | -2.1671 | -0.47193 | 0.80888 | -1.3505 | 0.13951 | 2.8886
Q16775 | Hydroxyacylglutathione hydrolase, mitochondrial | HAGH | 1:1 | Block_4 | -0.92562 | 4.424 | TRUE | 7.2129 | 5.763 | -0.70941 | 0.87995 | 2.2348 | 0.24511 | 1.6035 | 3.2067
Q9UJM8 | 2-Hydroxyacid oxidase 1 | HAO1 | 1:1 | Block_4 | -4.7293 | 4.9476 | FALSE | 6.5569 | 5.1091 | -2.7283 | -1.6544 | 0.32061 | -2.776 | -0.71182 | 1.7675
Q96D42 | Hepatitis A virus cellular receptor 1 | HAVCR1 | 1:1 | Block_4 | -1.7092 | 5.5544 | TRUE | 27.8512 | 10.2421 | -0.93874 | -0.31898 | 2.3165 | -0.26181 | 1.5381 | 3.2377
Q99075 | Proheparin-binding EGF-like growth factor | HBEGF | 1:1 | Block_4 | -1.4782 | 8.8399 | TRUE | 10.5786 | 3.9851 | 0.20737 | 1.3379 | 2.5465 | 0.8407 | 2.6803 | 4.3555
P09105 | Hemoglobin subunit theta-1 | HBQ1 | 1:1 | Block_4 | -0.50785 | 3.6662 | FALSE | 14.3228 | 10.9484 | 0.34924 | 1.6653 | 2.2604 | 0.47799 | 1.476 | 2.6642
P14317 | Hematopoietic lineage cell-specific protein | HCLS1 | 1:1 | Block_4 | -0.049374 | 6.3035 | TRUE | 21.3492 | 16.9846 | NA | 1.1307 | 3.6494 | 1.334 | 3.7296 | 4.954
Q7Z4H3 | 5'-deoxynucleotidase HDDC2 | HDDC2 | 1:1 | Block_4 | 2.0262 | 8.018 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P51858 | Hepatoma-derived growth factor | HDGF | 1:1 | Block_4 | -2.0029 | 6.0969 | TRUE | 14.7853 | 10.1802 | -0.90587 | 0.82453 | 3.4044 | 1.2629 | 3.0155 | 5.4338
Q7Z4V5 | Hepatoma-derived growth factor-related protein 2 | HDGFL2 | 1:1 | Block_4 | -1.0468 | 7.3479 | TRUE | 21.7949 | 12.1088 | -0.722 | 0.45077 | 1.9697 | 0.51376 | 2.3606 | 4.1467
A8MVW5 | HEPACAM family member 2 | HEPACAM2 | 1:1 | Block_4 | -5 | 4.4034 | TRUE | 22.1584 | 11.6334 | -0.94797 | 0.045456 | 0.61971 | -0.78072 | 0.17609 | 1.1303
O94992 | Protein HEXIM1 | HEXIM1 | 1:1 | Block_4 | 0.57424 | 4.3057 | TRUE | NA | NA | NA | NA | 1.7716 | 0.6636 | 1.6981 | 3.2636
P14210 | Hepatocyte growth factor | HGF | 1:1 | Block_4 | -5.235 | 6.4463 | TRUE | 7.878 | 4.0872 | -0.25386 | 0.40173 | 1.4602 | 0.59255 | 1.545 | 3.1132
Q03014 | Hematopoietically-expressed homeobox protein HHEX | HHEX | 1:1 | Block_4 | -1.6735 | 9.2806 | TRUE | 23.0412 | 13.0791 | -1.1019 | -0.035936 | 1.9524 | 0.83948 | 2.5103 | 4.0659
O75146 | Huntingtin-interacting protein 1-related protein | HIP1R | 1:1 | Block_4 | -3 | 2.4498 | FALSE | 33.2727 | 25.8189 | -1.1711 | -0.070755 | 0.92422 | -0.043241 | 1.0506 | 2.9708
Q6ZVN8 | Hemojuvelin | HJV | 1:1 | Block_4 | -0.02103 | 4.7228 | TRUE | 4.2223 | 8.0001 | NA | 0.091965 | 0.61475 | NA | 0.17306 | 0.64549
P04439 | HLA class I histocompatibility antigen, A alpha chain | HLA-A | 1:1 | Block_4 | -0.84601 | 3.7605 | FALSE | 12.7285 | 7.6918 | -0.51041 | 0.5336 | 1.2772 | 0.041572 | 0.97075 | 2.1379
P01903 | HLA class II histocompatibility antigen, DR alpha chain | HLA-DRA | 1:1 | Block_4 | 0.64939 | 7.9607 | FALSE | 8.4377 | 8.4603 | 0.68536 | 1.7371 | 2.7291 | NA | 1.6797 | 3.0905
P13747 | HLA class I histocompatibility antigen, alpha chain E | HLA-E | 1:1 | Block_4 | 1.8589 | 5.1233 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P08397 | Porphobilinogen deaminase | HMBS | 1:1 | Block_4 | -0.19894 | 4.1479 | FALSE | 0.84795 | 6.8401 | -0.18505 | 0.98614 | 1.5473 | 0.086716 | 1.0284 | 2.043
P35914 | Hydroxymethylglutaryl-CoA lyase, mitochondrial | HMGCL | 1:1 | Block_4 | 0.091718 | 7.933 | FALSE | NA | 2.1847 | NA | 0.18034 | 1.8717 | 0.3009 | 1.8397 | 3.1406
P30519 | Heme oxygenase 2 | HMOX2 | 1:1 | Block_4 | -1.5107 | 8.8261 | FALSE | 6.8769 | 6.5608 | -0.6317 | 0.072626 | 0.86166 | 0.13852 | 1.0207 | 2.3642
P50135 | Histamine N-methyltransferase | HNMT | 1:1 | Block_4 | -5.1706 | 6.3446 | TRUE | 9.6659 | 7.1363 | -0.56325 | 0.1464 | 0.9101 | 0.27106 | 1.1834 | 2.528
Q9BUJ2 | Heterogeneous nuclear ribonucleoprotein U-like protein 1 | HNRNPUL1 | 1:1 | Block_4 | -1.5507 | 8.8499 | TRUE | 7.9631 | 11.5717 | -0.92248 | 0.02079 | 1.1022 | 0.74087 | 1.8167 | 3.4954
Q9NSC5 | Homer protein homolog 3 | HOMER3 | 1:1 | Block_4 | -2.0681 | 9.7358 | TRUE | 14.1507 | 10.5288 | -0.64041 | 0.2552 | 1.9621 | 0.10513 | 1.0121 | 2.8735
Q96ED9 | Protein Hook homolog 2 | HOOK2 | 1:1 | Block_4 | 0.27369 | 10.1626 | TRUE | NA | NA | NA | NA | 1.1399 | NA | 0.46826 | 1.6659
P37235 | Hippocalcin-like protein 1 | HPCAL1 | 1:1 | Block_4 | -1.93 | 5.6656 | TRUE | 14.3995 | 12.5924 | -0.26139 | 1.515 | 3.7057 | 1.193 | 3.7967 | 4.9376
P23327 | Sarcoplasmic reticulum histidine-rich calcium-binding protein | HRC | 1:1 | Block_4 | -1.8223 | 5.471 | FALSE | 9.786 | 9.6898 | -0.66088 | -0.061989 | 0.31034 | -0.2012 | 0.35266 | 1.0696
Q53T59 | HCLS1-binding protein 3 | HS1BP3 | 1:1 | Block_4 | -2 | 6.1194 | TRUE | 13.8771 | 8.8906 | -0.38166 | 1.1889 | 3.5894 | 1.7139 | 3.9523 | 5.258
Q9Y662 | Heparan sulfate glucosamine 3-O-sulfotransferase 3B1 | HS3ST3B1 | 1:1 | Block_4 | -2.0797 | 5.5598 | FALSE | 8.501 | 7.1031 | -0.30156 | 0.75713 | 1.313 | -0.10439 | 1.0987 | 1.9267
O60243 | Heparan-sulfate 6-O-sulfotransferase 1 | HS6ST1 | 1:1 | Block_4 | -0.40724 | 4.9762 | FALSE | 23.4901 | 20.0117 | -0.26876 | 0.41664 | 1.1197 | 0.058248 | 0.95262 | 1.8104
O75506 | Heat shock factor-binding protein 1 | HSBP1 | 1:1 | Block_4 | 0.58879 | 7.0201 | TRUE | NA | 0.25014 | NA | 0.76057 | 2.6687 | 1.1226 | 2.8793 | 4.6431
P28845 | 11-beta-hydroxysteroid dehydrogenase 1 | HSD11B1 | 1:1 | Block_4 | -3.1416 | 3.247 | FALSE | 7.5454 | 8.3753 | -0.6165 | 0.23509 | 1.2067 | -0.40968 | 0.33777 | 1.1304
Q9BPX1 | 17-beta-hydroxysteroid dehydrogenase 14 | HSD17B14 | 1:1 | Block_4 | -0.90879 | 7.3079 | TRUE | 8.5951 | 6.3705 | -0.59816 | 0.76049 | 1.8602 | -0.37336 | 0.92754 | 2.2018
Q6YN16 | Hydroxysteroid dehydrogenase-like protein 2 | HSDL2 | 1:1 | Block_4 | 0.069164 | 5.6745 | FALSE | 12.1759 | 6.9021 | NA | NA | 0.74178 | NA | 0.72934 | 1.8534
P48723 | Heat shock 70 kDa protein 13 | HSPA13 | 1:1 | Block_4 | -4.9544 | 2.0468 | FALSE | 20.0365 | 9.5397 | -0.79024 | 0.17897 | 0.69565 | -0.30494 | 0.2301 | 0.75241
P54652 | Heat shock-related 70 kDa protein 2 | HSPA2 | 1:1 | Block_4 | 0.93696 | 8.8067 | FALSE | NA | NA | NA | NA | 0.98863 | NA | 1.4757 | 2.7122
O14558 | Heat shock protein beta-6 | HSPB6 | 1:1 | Block_4 | -3.9522 | 4.2947 | TRUE | 7.9377 | 4.8702 | -0.52948 | 0.21214 | 1.0237 | -0.10214 | 0.97563 | 2.4648
P86397 | Hydroxyacyl-thioester dehydratase type 2, mitochondrial | HTD2 | 1:1 | Block_4 | 0.91929 | 13.1729 | TRUE | NA | NA | NA | 2.1319 | 5.5704 | NA | 3.9762 | 6.4646
Q14773 | Intercellular adhesion molecule 4 | ICAM4 | 1:1 | Block_4 | -3.5446 | 4.4822 | TRUE | 7.6689 | 7.0701 | -0.86613 | 0.070858 | 0.58422 | -0.7182 | 0.25748 | 1.2015
O75144 | ICOS ligand | ICOSLG | 1:1 | Block_4 | -3.2113 | 2.7291 | TRUE | 9.4122 | 9.0012 | -0.42006 | -0.005173 | 0.51737 | -0.5964 | -0.021654 | 0.62094
P14902 | Indoleamine 2,3-dioxygenase 1 | IDO1 | 1:1 | Block_4 | -1.9063 | 9.5595 | TRUE | 10.0719 | 8.728 | 0.079681 | 1.2478 | 2.2492 | 0.092195 | 1.4588 | 3.9477
P22304 | Iduronate 2-sulfatase | IDS | 1:1 | Block_4 | -2.0323 | 4.8056 | FALSE | 5.0736 | 2.1716 | 0.045637 | 0.33511 | 0.70388 | -0.025077 | 0.28167 | 0.62673
P35475 | Alpha-L-iduronidase | IDUA | 1:1 | Block_4 | -4.0099 | 2.1979 | TRUE | 10.6313 | 5.965 | -2.5435 | -1.0197 | 0.15967 | -1.9392 | -0.58944 | 0.26274
P13284 | Gamma-interferon-inducible lysosomal thiol reductase | IFI30 | 1:1 | Block_4 | -2.5 | 5.1613 | FALSE | 15.9868 | 11.1688 | -1.0906 | -0.30862 | 0.84895 | -0.31223 | 0.81979 | 2.148
P17181 | Interferon alpha/beta receptor 1 | IFNAR1 | 1:1 | Block_4 | -4 | 4.1574 | TRUE | 15.7204 | 11.592 | -0.27579 | 0.1893 | 0.60222 | -0.1995 | 0.42694 | 0.93502
P38484 | Interferon gamma receptor 2 | IFNGR2 | 1:1 | Block_4 | -4.1469 | 5.6936 | TRUE | 8.9743 | 1.6448 | -1.3061 | -0.33055 | 1.7682 | -0.94386 | 0.058549 | 1.1684
Q8IU57 | Interferon lambda receptor 1 | IFNLR1 | 1:1 | Block_4 | -2.422 | 9.8246 | FALSE | 15.5446 | 5.3745 | -0.79612 | 0.028449 | 0.56113 | -0.46029 | 0.22788 | 0.8161
P78318 | Immunoglobulin-binding protein 1 | IGBP1 | 1:1 | Block_4 | -1.5071 | 8.6788 | FALSE | 17.3471 | 6.8758 | 0.21676 | 2.0586 | 4.4075 | 2.2056 | 4.3951 | 5.7236
P08069 | Insulin-like growth factor 1 receptor | IGF1R | 1:1 | Block_4 | -2.4061 | 6.8791 | FALSE | 18.0776 | 10.2598 | -0.11632 | 0.23636 | 0.74565 | 0.13804 | 0.85479 | 1.6619
Q6B9Z1 | Insulin growth factor-like family member 4 | IGFL4 | 1:1 | Block_4 | -0.57271 | 11.8678 | TRUE | 13.9156 | 21.1657 | NA | -0.4398 | 1.3187 | NA | -0.065513 | 1.9461
Q96ID5 | Immunoglobulin superfamily member 21 | IGSF21 | 1:1 | Block_4 | -1.2518 | 6.003 | TRUE | 13.1669 | 8.4941 | -0.061282 | 0.47607 | 0.9434 | -0.21817 | 0.52656 | 1.2923
O75054 | Immunoglobulin superfamily member 3 | IGSF3 | 1:1 | Block_4 | -3.3324 | 7.9032 | TRUE | 6.924 | 7.1255 | -0.8154 | 0.05975 | 1.3323 | -0.46941 | 0.37689 | 1.3613
Q9Y6K9 | NF-kappa-B essential modulator | IKBKG | 1:1 | Block_4 | 0.001 | 6.2813 | TRUE | 12.347 | 17.2948 | NA | 0.68079 | 2.0551 | 0.38371 | 2.0645 | 4.0212
P29459_P29460 | Interleukin-12 | IL12A_IL12B | 1:1 | Block_4 | -2.7837 | 5.9191 | FALSE | 9.3354 | 5.4586 | -1.1937 | 0.51366 | 1.8858 | -0.97827 | 1.355 | 3.8145
P29460 | Interleukin-12 subunit beta | IL12B | 1:1 | Block_4 | 0.5 | NA | NA | 14.2601 | 22.4532 | NA | NA | 1.03 | NA | NA | 2.03
P40933 | Interleukin-15 | IL15 | 1:1 | Block_4 | -1.375 | 10.1441 | TRUE | 12.6704 | 12.3015 | -0.60072 | 0.027482 | 0.51854 | -0.25451 | 0.42425 | 1.5639
Q14005 | Pro-interleukin-16 | IL16 | 1:1 | Block_4 | -2.7264 | 8.5273 | TRUE | 12.3669 | 8.2443 | -1.003 | -0.010868 | 0.65567 | -0.22985 | 1.0021 | 4.4465
Q9NRM6 | Interleukin-17 receptor B | IL17RB | 1:1 | Block_4 | -0.54304 | 5.765 | TRUE | 6.2422 | 9.9969 | NA | 0.23991 | 0.7492 | -0.15219 | 0.7295 | 1.738
Q13478 | Interleukin-18 receptor 1 | IL18R1 | 1:1 | Block_4 | -5.9751 | 6.0418 | TRUE | 6.7311 | 2.4994 | -0.50271 | -0.02666 | 0.68064 | -0.12669 | 0.6916 | 1.7487
O95256 | Interleukin-18 receptor accessory protein | IL18RAP | 1:1 | Block_4 | -3.1588 | 6.242 | TRUE | 14.1995 | 8.8728 | -2.6824 | -0.79256 | 1.4141 | -2.9016 | -0.84374 | 2.3723
Q9GZX6 | Interleukin-22 | IL22 | 1:1 | Block_4 | -0.32246 | 8.6792 | TRUE | 11.9982 | 10.7814 | NA | 0.28238 | 1.1364 | -0.05206 | 0.99713 | 3.152
Q5VWK5 | Interleukin-23 receptor | IL23R | 1:1 | Block_4 | -3.6031 | 9.6366 | TRUE | 10.6931 | 2.8538 | -0.43179 | -0.07438 | 0.60903 | -0.6645 | 0.19006 | 1.2161
P31785 | Cytokine receptor common subunit gamma | IL2RG | 1:1 | Block_4 | 0.52058 | 5.9048 | FALSE | NA | NA | NA | NA | 0.52914 | NA | NA | NA
P24001 | Interleukin-32 | IL32 | 1:1 | Block_4 | -2.2242 | 7.1877 | TRUE | 10.2851 | 9.5159 | -0.33342 | 0.74261 | 1.8403 | -0.044788 | 0.72613 | 2.068
Q9NZH8 | Interleukin-36 gamma | IL36G | 1:1 | Block_4 | -0.5 | 7.1566 | TRUE | 20.3734 | 14.9515 | NA | -0.058923 | 1.3306 | NA | 0.13354 | 1.7704
P24394 | Interleukin-4 receptor subunit alpha | IL4R | 1:1 | Block_4 | -0.63995 | 8.2416 | TRUE | 9.8872 | 8.0986 | NA | 0.077783 | 0.56859 | -0.23266 | 0.67661 | 1.8001
Q01344 | Interleukin-5 receptor subunit alpha | IL5RA | 1:1 | Block_4 | -3.9008 | 5.7477 | TRUE | 13.8835 | 10.5126 | -1.586 | -0.41537 | 0.21098 | -1.1016 | 0.053451 | 1.0508
P05231 | Interleukin-6 | IL6 | 1:1 | Block_4 | -1.8843 | 9.0807 | TRUE | 13.9525 | 16.6411 | -1.5443 | -0.78447 | 0.19484 | -1.0034 | 0.42493 | 2.8908
P13232 | Interleukin-7 | IL7 | 1:1 | Block_4 | -1.2289 | 11.3209 | TRUE | 11.3787 | 7.3925 | -0.077572 | 1.0482 | 2.1642 | 0.99378 | 2.4327 | 3.5871
Q16891 | MICOS complex subunit MIC60 | IMMT | 1:1 | Block_4 | 0.001 | 9.3217 | FALSE | 11.9304 | 12.5924 | NA | NA | 0.70635 | 0.040929 | 0.78833 | 2.127
P29218 | Inositol monophosphatase 1 | IMPA1 | 1:1 | Block_4 | -0.38534 | 6.7792 | FALSE | 8.0319 | 10.3371 | -0.075789 | 0.76048 | 2.0963 | 0.38344 | 2.4322 | 3.7227
Q9P2X3 | Protein IMPACT | IMPACT | 1:1 | Block_4 | -0.26219 | 6.4672 | FALSE | 9.1166 | 4.8378 | -0.11837 | 1.3426 | 2.574 | 0.94327 | 2.836 | 4.0219
P49441 | Inositol polyphosphate 1-phosphatase | INPP1 | 1:1 | Block_4 | -0.5 | 5.5786 | FALSE | 22.5298 | 11.1616 | NA | 0.22588 | 0.84389 | 0.40969 | 1.1954 | 2.4931
Q92835 | Phosphatidylinositol 3,4,5-trisphosphate 5-phosphatase 1 | INPP5D | 1:1 | Block_4 | -2.5593 | 8.4379 | TRUE | 13.3854 | 5.4225 | -0.63826 | 0.92816 | 1.9044 | 0.77754 | 2.1133 | 4.0337
O15357 | Phosphatidylinositol 3,4,5-trisphosphate 5-phosphatase 2 | INPPL1 | 1:1 | Block_4 | 0.97384 | 8.5998 | TRUE | NA | NA | NA | NA | 3.4751 | 1.1996 | 3.7149 | 5.4717
P51460 | Insulin-like 3 | INSL3 | 1:1 | Block_4 | -2.1994 | 3.6162 | FALSE | 8.5197 | 5.6223 | NA | NA | 2.0991 | NA | NA | 1.8216
Q14641 | Early placenta insulin-like peptide | INSL4 | 1:1 | Block_4 | -1.5015 | 10.077 | TRUE | 13.3152 | 7.8402 | NA | -1.1826 | 0.93589 | NA | -0.71048 | 0.87214
Q9Y5Q6 | Insulin-like peptide INSL5 | INSL5 | 1:1 | Block_4 | -2.4733 | 8.5047 | TRUE | 12.734 | 9.0887 | -1.8305 | -0.95575 | 1.3256 | -1.0773 | 0.35849 | 2.715
P06213 | Insulin receptor | INSR | 1:1 | Block_4 | 0.25986 | 5.1595 | TRUE | 1.2353 | 5.0286 | NA | NA | 0.36007 | NA | NA | 0.52486
Q8WWN9 | Interactor protein for cytohesin exchange factors 1 | IPCEF1 | 1:1 | Block_4 | -3.3763 | 11.1846 | TRUE | 10.5778 | 10.0562 | -1.2505 | 0.70903 | 3.5753 | 0.76673 | 3.6842 | 5.9539
Q13576 | Ras GTPase-activating-like protein IQGAP2 | IQGAP2 | 1:1 | Block_4 | 0.5 | 9.0685 | TRUE | 10.157 | 2.044 | NA | 1.0386 | 2.49 | 0.91806 | 2.2585 | 4.1008
Q12912 | Inositol 1,4,5-triphosphate receptor associated 2 | IRAG2 | 1:1 | Block_4 | -0.9304 | 9.6372 | FALSE | 12.4686 | 13.9524 | -0.90833 | 0.027937 | 1.9714 | -0.016926 | 1.1729 | 4.8021
Q9NWZ3 | Interleukin-1 receptor-associated kinase 4 | IRAK4 | 1:1 | Block_4 | -0.68117 | 7.629 | FALSE | 13.2374 | 15.1582 | -0.46727 | 1.2601 | 5.1094 | 1.8639 | 5.209 | 6.6275
Q6UXK2 | Immunoglobulin superfamily containing leucine-rich repeat protein 2 | ISLR2 | 1:1 | Block_4 | -1.1683 | 7.7508 | TRUE | 12.2358 | 11.506 | -0.86715 | 0.00061492 | 0.61244 | -0.87314 | -0.19102 | 0.52248
B1AKI9 | Isthmin-1 | ISM1 | 1:1 | Block_4 | -4.1167 | 5.3275 | FALSE | 10.7778 | 7.4351 | 0.20648 | 0.95599 | 1.6558 | 0.38764 | 1.1726 | 1.9852
P53990 | IST1 homolog | IST1 | 1:1 | Block_4 | -0.67872 | 5.8343 | FALSE | 18.4879 | 12.9269 | -0.36555 | 1.7288 | 5.2029 | 2.1583 | 5.1392 | 6.0683
Q9UKX5 | Integrin alpha-11 | ITGA11 | 1:1 | Block_4 | -3.0158 | 4.5774 | FALSE | 3.68 | 3.8485 | -0.63001 | 0.26968 | 1.0854 | -1.0923 | -0.24432 | 0.54512
P23229 | Integrin alpha-6 | ITGA6 | 1:1 | Block_4 | 0.08 | 3.5923 | FALSE | NA | NA | NA | NA | 1.5812 | NA | 0.73975 | 2.3392
P20701 | Integrin alpha-L | ITGAL | 1:1 | Block_4 | -3.8423 | 0.06826 | FALSE | 14.7866 | 9.9167 | -0.51094 | 0.56118 | 1.4905 | -0.52721 | 0.77804 | 2.0479
P06756 | Integrin alpha-V | ITGAV | 1:1 | Block_4 | -2.9658 | 4.0178 | TRUE | 10.1425 | 8.3753 | -0.23954 | 0.19013 | 0.57561 | -0.31279 | 0.15613 | 0.58719
Q9UKP3 | Integrin beta-1-binding protein 2 | ITGB1BP2 | 1:1 | Block_4 | 0.85842 | NA | NA | 9.1909 | 6.2072 | NA | 2.5027 | 6.1346 | 2.3646 | 5.8182 | 8.0969
P26010 | Integrin beta-7 | ITGB7 | 1:1 | Block_4 | -2.7307 | 5.4767 | TRUE | 10.1985 | 8.2779 | -0.5921 | 0.024372 | 0.69962 | -0.71658 | 0.10296 | 1.0912
Q86UX2 | Inter-alpha-trypsin inhibitor heavy chain H5 | ITIH5 | 1:1 | Block_4 | 0.31562 | 8.3331 | FALSE | NA | NA | NA | NA | 0.4696 | NA | NA | 1.0151
O43736 | Integral membrane protein 2A | ITM2A | 1:1 | Block_4 | -1.3814 | 5.2019 | FALSE | 7.7487 | 9.1434 | -0.20014 | 0.48738 | 1.0471 | -0.1174 | 0.56695 | 1.4213
Q14643 | Inositol 1,4,5-trisphosphate receptor type 1 | ITPR1 | 1:1 | Block_4 | 0.5 | 12.3086 | TRUE | 12.2357 | 12.5492 | NA | 1.7663 | 2.9621 | 1.4989 | 3.1918 | 4.2219
Q9H910 | Jupiter microtubule associated homolog 2 | JPT2 | 1:1 | Block_4 | -0.5 | 8.2786 | TRUE | 30.991 | 14.4149 | -0.15383 | 1.3453 | 4.6655 | 1.9239 | 4.7009 | 6.3756
Q96I82 | Kazal-type serine protease inhibitor domain-containing protein 1 | KAZALD1 | 1:1 | Block_4 | -2.2984 | 6.5225 | TRUE | 8.5773 | 3.5609 | -0.38868 | 0.56025 | 1.418 | 0.11051 | 0.90829 | 2.056
Q674X7 | Kazrin | KAZN | 1:1 | Block_4 | 1.2492 | 4.3517 | FALSE | NA | NA | NA | NA | 1.7447 | NA | 1.5366 | 3.5499
Q9NXV2 | BTB/POZ domain-containing protein KCTD5 | KCTD5 | 1:1 | Block_4 | 0.036028 | 10.1415 | FALSE | 11.4487 | NA | NA | NA | 0.60578 | 0.39075 | 1.1381 | 2.2268
P23276 | Kell blood group glycoprotein | KEL | 1:1 | Block_4 | -1.6308 | 6.1749 | FALSE | 6.5568 | 6.4916 | -0.20466 | 0.34825 | 0.7193 | -0.087898 | 0.57499 | 1.3832
P50053 | Ketohexokinase | KHK | 1:1 | Block_4 | -1 | NA | NA | 15.8966 | 11.9432 | 1.1356 | 2.4971 | 4.3035 | 2.2395 | 3.9317 | 5.1567
Q5VV43 | Dyslexia-associated protein KIAA0319 | KIAA0319 | 1:1 | Block_4 | -2 | 8.1605 | TRUE | 10.4303 | 16.0091 | -0.69232 | -0.087839 | 0.50725 | -0.31165 | 0.30785 | 1.2356
Q8IYS2 | Uncharacterized protein KIAA2013 | KIAA2013 | 1:1 | Block_4 | 0.69726 | 3.7394 | FALSE | NA | NA | NA | NA | 1.038 | NA | NA | NA
Q14807 | Kinesin-like protein KIF22 | KIF22 | 1:1 | Block_4 | 0.32241 | 11.7881 | TRUE | 13.2683 | 15.77 | NA | 2.0733 | 5.1766 | 1.3354 | 4.2571 | 6.1926
Q9BVG8 | Kinesin-like protein KIFC3 | KIFC3 | 1:1 | Block_4 | -0.69918 | 11.6271 | TRUE | 6.3279 | 4.3035 | NA | 0.43513 | 2.4847 | 0.57377 | 2.1218 | 5.9504
P43627_P43628 | Killer cell immunoglobulin-like receptor 2DL2/2DL3 | KIR2DL2_KIR2DL3 | 1:1 | Block_4 | -0.27778 | 5.4705 | TRUE | 22.1395 | 12.5067 | NA | 0.76807 | 1.8178 | NA | 0.82223 | 2.2165
Q86Z14 | Beta-klotho | KLB | 1:1 | Block_4 | -2.9775 | 6.4451 | FALSE | 6.7872 | 6.7308 | -1.1467 | 0.033963 | 1.345 | -0.50597 | 0.65574 | 1.7812
Q9UKR3 | Kallikrein-13 | KLK13 | 1:1 | Block_4 | -2 | 7.8873 | TRUE | 17.3305 | 9.3216 | 0.099935 | 1.3147 | 2.0672 | -0.29286 | 0.81437 | 1.979
Q9P0G3 | Kallikrein-14 | KLK14 | 1:1 | Block_4 | -7 | 2.3934 | TRUE | 6.0074 | 1.3938 | -1.4239 | -0.18445 | 0.94118 | -1.3033 | -0.10895 | 0.94498
Q9H2R5 | Kallikrein-15 | KLK15 | 1:1 | Block_4 | -1.7273 | 9.4499 | TRUE | 15.5376 | 6.1158 | -0.81394 | -0.002589 | 0.62216 | -0.30523 | 0.36182 | 1.2735
Q9Y5K2 | Kallikrein-4 | KLK4 | 1:1 | Block_4 | -4.5 | 8.8803 | TRUE | 8.5496 | 6.5123 | -0.23715 | 0.58911 | 1.5662 | 0.20978 | 1.6837 | 3.5305
Q12918 | Killer cell lectin-like receptor subfamily B member 1 | KLRB1 | 1:1 | Block_4 | -5.5373 | 4.7427 | TRUE | 5.7407 | 6.5216 | -0.58191 | 0.14827 | 1.1937 | -0.15166 | 0.92928 | 2.2621
Q13241 | Natural killer cells antigen CD94 | KLRD1 | 1:1 | Block_4 | -3.1005 | 5.7159 | TRUE | 31.1622 | 7.1724 | -0.69475 | 0.35393 | 1.3557 | -0.25676 | 0.72845 | 1.6315
P26718 | NKG2-D type II integral membrane protein | KLRK1 | 1:1 | Block_4 | -2.8005 | 7.0118 | TRUE | 10.5431 | 8.8848 | -0.58578 | 0.23858 | 0.86815 | -0.2794 | 0.56501 | 1.343
P05783 | Keratin, type I cytoskeletal 18 | KRT18 | 1:1 | Block_4 | -4.9533 | 7.5448 | TRUE | 5.8942 | 5.5693 | -1.4162 | -0.9041 | 1.0363 | -0.92678 | 0.49718 | 2.241
P08727 | Keratin, type I cytoskeletal 19 | KRT19 | 1:1 | Block_4 | -1.4444 | 6.5451 | TRUE | 11.7334 | 10.7543 | -0.91157 | -0.33964 | 1.3725 | -0.55585 | 1.011 | 2.5438
P48668 | Keratin, type II cytoskeletal 6C | KRT6C | 1:1 | Block_4 | -4.042 | 6.7432 | TRUE | 10.965 | 9.6816 | NA | NA | -3.0381 | NA | -3.901 | -2.0218
Q16719 | Kynureninase | KYNU | 1:1 | Block_4 | -2 | 4.6914 | TRUE | 12.546 | 15.4471 | 0.080643 | 1.2513 | 1.8285 | 0.92985 | 1.9665 | 2.7617
Q9GZZ8 | Extracellular glycoprotein lacritin | LACRT | 1:1 | Block_4 | -0.35843 | 5.1446 | FALSE | 10.4977 | 3.4713 | -0.357 | 0.41198 | 1.8472 | NA | -0.137 | 0.81751
P18627 | Lymphocyte activation gene 3 protein | LAG3 | 1:1 | Block_4 | -2.5632 | 7.9462 | FALSE | 8.9494 | 5.4969 | -0.42114 | 0.54437 | 0.98114 | -0.11992 | 0.80009 | 1.9428
Q6ISS4 | Leukocyte-associated immunoglobulin-like receptor 2 | LAIR2 | 1:1 | Block_4 | -3.8489 | 7.2118 | TRUE | 8.0419 | 7.9461 | -2.2435 | -0.43232 | 0.63713 | -2.0009 | 0.093139 | 2.3127
P07942 | Laminin subunit beta-1 | LAMB1 | 1:1 | Block_4 | -4.5 | 4.8251 | TRUE | 20.3254 | 8.0049 | -0.48446 | 0.12217 | 0.66818 | -0.16345 | 0.52462 | 1.1889
Q9UQV4 | Lysosome-associated membrane glycoprotein 3 | LAMP3 | 1:1 | Block_4 | -5 | 2.8913 | TRUE | 10.0705 | 8.6852 | -1.2983 | 0.49499 | 1.3602 | -0.6078 | 0.43296 | 1.2968
O43504 | Ragulator complex protein LAMTOR5 | LAMTOR5 | 1:1 | Block_4 | -1.3535 | 6.8316 | TRUE | 11.9761 | 9.5088 | NA | -0.81639 | 0.28577 | -0.76632 | 0.37409 | 2.2793
Q6PKG0 | La-related protein 1 | LARP1 | 1:1 | Block_4 | -1.9652 | 9.4562 | FALSE | 11.0266 | 6.3554 | NA | -1.2702 | 0.55691 | -0.83356 | 0.90863 | 2.6575
O43561 | Linker for activation of T-cells family member 1 | LAT | 1:1 | Block_4 | -0.67063 | 8.8375 | TRUE | 15.0396 | 14.8976 | 0.19377 | 2.6037 | 5.9169 | 2.8899 | 6.0582 | 6.8715
Q9GZY6 | Linker for activation of T-cells family member 2 | LAT2 | 1:1 | Block_4 | -1 | 9.8418 | FALSE | 16.7787 | 24.6738 | -0.31957 | 1.3198 | 5.4147 | 1.3454 | 4.7541 | 6.4656
O95835 | Serine/threonine-protein kinase LATS1 | LATS1 | 1:1 | Block_4 | -0.11387 | 12.1428 | TRUE | 19.0007 | 13.163 | -0.10681 | 1.5436 | 4.3127 | 1.481 | 4.2078 | 5.6899
Q6UX15 | Layilin | LAYN | 1:1 | Block_4 | -1.808 | 7.5857 | TRUE | 11.9449 | 12.2569 | -0.32907 | 0.413 | 0.76959 | 0.25134 | 1.0864 | 2.4499
Q6UWW0 | Lipocalin-15 | LCN15 | 1:1 | Block_4 | -2 | 5.7439 | FALSE | 16.0421 | 8.8741 | -1.3023 | -0.46404 | 0.77988 | -0.71846 | 0.10629 | 1.4388
Q5SW96 | Low density lipoprotein receptor adapter protein 1 | LDLRAP1 | 1:1 | Block_4 | 1 | 5.078 | FALSE | NA | NA | NA | 1.586 | 5.9108 | 1.4791 | 5.5003 | 7.1637
O00292 | Left-right determination factor 2 | LEFTY2 | 1:1 | Block_4 | -4.8959 | 5.4814 | FALSE | 6.9547 | 6.5696 | -0.86097 | -0.079243 | 1.1113 | -0.70057 | 0.35187 | 1.3247
O95202 | Mitochondrial proton/calcium exchanger protein | LETM1 | 1:1 | Block_4 | 0.14512 | 10.8108 | TRUE | 14.4242 | 12.2695 | NA | 0.87591 | 3.2325 | 1.1487 | 2.4197 | 4.0359
P47929 | Galectin-7 | LGALS7 | 1:1 | Block_4 | -2 | 2.232 | FALSE | 10.9165 | 8.9389 | -0.79069 | -0.13865 | 0.92557 | -0.27196 | 0.73673 | 1.9037
P42702 | Leukemia inhibitory factor receptor | LIFR | 1:1 | Block_4 | -1.971 | 5.524 | TRUE | 9.7963 | 7.3392 | -0.24149 | 0.36545 | 0.72441 | -0.10767 | 0.53086 | 1.2274
P07098 | Gastric triacylglycerol lipase | LIPF | 1:1 | Block_4 | -1.3526 | 7.2707 | FALSE | 9.4441 | 4.576 | -0.54825 | 0.27484 | 0.98361 | -0.76466 | 0.27077 | 1.2268
P20700 | Lamin-B1 | LMNB1 | 1:1 | Block_4 | -1.1378 | 8.1699 | FALSE | 21.6106 | 16.6514 | NA | NA | -0.97234 | NA | -0.87517 | 0.91348
Q03252 | Lamin-B2 | LMNB2 | 1:1 | Block_4 | -1.8381 | 5.4143 | FALSE | 6.6166 | 12.378 | -0.66983 | -0.086393 | 0.38383 | 0.12516 | 0.88384 | 2.474
P29536 | Leiomodin-1 | LMOD1 | 1:1 | Block_4 | -2.5447 | 10.4073 | TRUE | 6.172 | 2.8738 | -0.33716 | 0.54145 | 1.1055 | -0.023336 | 1.2752 | 2.5542
P36776 | Lon protease homolog, mitochondrial | LONP1 | 1:1 | Block_4 | 0.97993 | 8.1931 | TRUE | NA | NA | NA | NA | 2.4953 | NA | 1.7557 | 3.343
Q9Y4K0 | Lysyl oxidase homolog 2 | LOXL2 | 1:1 | Block_4 | 0.001 | 6.5571 | TRUE | 15.504 | 6.1677 | 0.027608 | 0.95084 | 1.5407 | NA | 1.3248 | 2.5105
Q7L5N7 | Lysophosphatidylcholine acyltransferase 2 | LPCAT2 | 1:1 | Block_4 | 0.37157 | 4.1454 | FALSE | NA | NA | NA | NA | 0.56908 | NA | NA | 1.0947
P22079 | Lactoperoxidase | LPO | 1:1 | Block_4 | -4.16 | 6.7027 | TRUE | 6.0309 | 8.0644 | -1.8449 | -0.93864 | 0.13552 | -1.8123 | -0.99708 | 0.51849
Q93052 | Lipoma-preferred partner | LPP | 1:1 | Block_4 | 0.80054 | 4.4652 | FALSE | NA | NA | NA | 0.93798 | 2.703 | NA | 2.4092 | 3.749
O75427 | Leucine-rich repeat and calponin homology domain-containing protein 4 | LRCH4 | 1:1 | Block_4 | -1.5473 | 10.0442 | TRUE | 11.5074 | 9.3543 | -0.64067 | 0.42694 | 1.994 | 0.55344 | 2.2445 | 5.9941
Q96JA1 | Leucine-rich repeats and immunoglobulin-like domains protein 1 | LRIG1 | 1:1 | Block_4 | -3.332 | NA | NA | 8.6979 | 5.5851 | -0.82876 | 0.21757 | 0.8021 | -0.13699 | 0.58001 | 1.2969
Q07954 | Prolow-density lipoprotein receptor-related protein 1 | LRP1 | 1:1 | Block_4 | -0.32062 | 4.095 | TRUE | 14.5324 | 12.8176 | -0.119 | 0.2659 | 0.73367 | -0.15124 | 0.3934 | 0.86307
Q8N386 | Leucine-rich repeat-containing protein 25 | LRRC25 | 1:1 | Block_4 | -1.8417 | 5.1994 | TRUE | 7.7995 | 9.7569 | -0.67972 | 0.18516 | 0.65481 | -0.34177 | 0.59768 | 1.4405
A6NM11 | Leucine-rich repeat-containing protein 37A2 | LRRC37A2 | 1:1 | Block_4 | -3.07 | 5.3141 | TRUE | 13.029 | 7.1127 | -0.49327 | 0.14574 | 0.74002 | -0.49555 | 0.40059 | 1.3165
Q8N967 | Leucine-rich repeat and transmembrane domain-containing protein 2 | LRTM2 | 1:1 | Block_4 | -0.44723 | 8.0929 | TRUE | 13.6541 | 8.1778 | -0.42695 | 0.13263 | 0.64992 | -0.39564 | 0.19084 | 0.90854
P33241 | Lymphocyte-specific protein 1 | LSP1 | 1:1 | Block_4 | -1.0483 | 3.7358 | TRUE | 16.7977 | 12.1361 | -0.83707 | -0.081302 | 0.59003 | -0.63999 | 0.39026 | 2.0407
P09960 | Leukotriene A-4 hydrolase | LTA4H | 1:1 | Block_4 | -3.1248 | 2.0256 | FALSE | 8.1286 | 8.5677 | -0.19841 | 0.69058 | 1.1546 | 0.19664 | 0.82249 | 1.5104
Q06643 | Lymphotoxin-beta | LTB | 1:1 | Block_4 | 1.3579 | 8.7483 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86TE4 | Leucine zipper protein 2 | LUZP2 | 1:1 | Block_4 | -2.5817 | 9.1146 | TRUE | 11.3236 | 5.0333 | -0.44718 | 0.16462 | 0.63928 | -0.25105 | 0.42069 | 1.4104
Q9BS40 | Latexin | LXN | 1:1 | Block_4 | -0.72336 | 8.4656 | FALSE | 13.1099 | 10.8509 | 0.44612 | 1.5104 | 3.3244 | 1.7545 | 2.8058 | 4.6527
O60449 | Lymphocyte antigen 75 | LY75 | 1:1 | Block_4 | -2.4356 | 6.4543 | TRUE | 6.2389 | 11.2372 | -0.22318 | 0.37614 | 1.1 | 0.1524 | 0.80577 | 1.5623
Q9HBG7 | T-lymphocyte surface antigen Ly-9 | LY9 | 1:1 | Block_4 | -4.7438 | 4.5579 | TRUE | 6.302 | 5.1211 | -0.40995 | 0.12852 | 0.93296 | -0.1063 | 0.52214 | 1.3789
Q9Y6Y9 | Lymphocyte antigen 96 | LY96 | 1:1 | Block_4 | -1.9383 | 4.6534 | FALSE | 8.0667 | 11.534 | -0.63606 | 0.05588 | 0.39605 | -0.22419 | 0.36182 | 1.1742
P07948 | Tyrosine-protein kinase Lyn | LYN | 1:1 | Block_4 | 0.2007 | NA | NA | 11.1424 | 21.5549 | NA | 1.4939 | 4.51 | 1.1645 | 4.2956 | 5.4317
Q9Y250 | Leucine zipper putative tumor suppressor 1 | LZTS1 | 1:1 | Block_4 | -0.37887 | 11.8619 | TRUE | 8.0641 | 8.0897 | NA | -0.030379 | 0.44123 | -0.13641 | 0.83151 | 2.6405
P20645 | Cation-dependent mannose-6-phosphate receptor | M6PR | 1:1 | Block_4 | 0.56915 | 8.8896 | TRUE | 0.64948 | 3.7693 | NA | NA | 0.6963 | NA | NA | 1.1251
Q9Y6D9 | Mitotic spindle assembly checkpoint protein MAD1 | MAD1L1 | 1:1 | Block_4 | -3.4416 | 6.289 | FALSE | 8.6358 | 7.1771 | -1.493 | -0.52478 | 0.42629 | 0.17487 | 1.0741 | 2.7623
P20916 | Myelin-associated glycoprotein | MAG | 1:1 | Block_4 | 0.5 | 4.1687 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O60476 | Mannosyl-oligosaccharide 1,2-alpha-mannosidase IB | MAN1A2 | 1:1 | Block_4 | -5.26 | 6.1245 | TRUE | 10.6819 | 9.7476 | 0.00091299 | 0.67013 | 1.2113 | 0.23559 | 0.84912 | 1.4766
P11137 | Microtubule-associated protein 2 | MAP2 | 1:1 | Block_4 | 0.57913 | 5.1301 | FALSE | NA | NA | NA | NA | 0.6015 | NA | 0.60281 | 1.743
Q02750 | Dual specificity mitogen-activated protein kinase kinase 1 | MAP2K1 | 1:1 | Block_4 | 1.1446 | 7.9518 | FALSE | NA | NA | NA | NA | 2.5636 | 1.4557 | 3.3725 | 5.3794
P52564 | Dual specificity mitogen-activated protein kinase kinase 6 | MAP2K6 | 1:1 | Block_4 | 0.39 | 9.1278 | FALSE | 1.8979 | 3.6854 | NA | 1.2354 | 3.9233 | 1.0194 | 3.6954 | 5.3316
Q9Y4K4 | Mitogen-activated protein kinase kinase kinase kinase 5 | MAP4K5 | 1:1 | Block_4 | 0.032249 | 11.5101 | TRUE | 26.0861 | 9.2506 | NA | 1.6724 | 5.4987 | 2.029 | 5.0996 | 7.3267
P45984 | Mitogen-activated protein kinase 9 | MAPK9 | 1:1 | Block_4 | -1.1627 | 6.7422 | FALSE | 9.692 | 8.9612 | -0.47907 | 0.23655 | 1.0949 | 0.62046 | 1.4289 | 2.6313
P48740 | Mannan-binding lectin serine protease 1 | MASP1 | 1:1 | Block_4 | -3.1831 | 1.1474 | FALSE | 5.917 | 1.5734 | -1.4987 | -1.0899 | -0.32075 | -1.3495 | -0.78933 | 0.12021
O15232 | Matrilin-3 | MATN3 | 1:1 | Block_4 | -6.61 | 6.1764 | TRUE | 6.0845 | 2.5957 | 0.20004 | 0.69888 | 1.257 | 0.18242 | 0.95006 | 2.2581
Q96PE7 | Methylmalonyl-CoA epimerase, mitochondrial | MCEE | 1:1 | Block_4 | 0.4 | 10.209 | FALSE | 6.2723 | 11.0551 | NA | 0.82972 | 3.6331 | 0.69614 | 2.0061 | 3.8001
Q8IX19 | Mast cell-expressed membrane protein 1 | MCEMP1 | 1:1 | Block_4 | 1.7616 | 3.6057 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.8728
O15068 | Guanine nucleotide exchange factor DBS | MCF2L | 1:1 | Block_4 | -2.0008 | 11.6577 | TRUE | 10.6746 | 5.2013 | -0.36027 | 0.78596 | 3.4214 | 1.2934 | 2.6554 | 5.1301
Q8NFP4 | MAM domain-containing glycosylphosphatidylinositol anchor protein 1 | MDGA1 | 1:1 | Block_4 | -3.065 | 6.1797 | FALSE | 10.3465 | 11.5329 | -0.7822 | 0.6002 | 1.4387 | -0.16505 | 0.81717 | 2.0571
Q9BV79 | Enoyl-[acyl-carrier-protein] reductase, mitochondrial | MECR | 1:1 | Block_4 | -1.33 | 11.6606 | FALSE | 22.413 | 21.141 | -0.48313 | 2.0043 | 5.0503 | 2.1183 | 4.9794 | 6.7966
Q16819 | Meprin A subunit alpha | MEP1A | 1:1 | Block_4 | -1.3646 | 6.3204 | FALSE | 12.3093 | 8.7303 | 0.041329 | 0.90965 | 2.7273 | -0.39222 | 1.3468 | 2.8899
Q6UB28 | Methionine aminopeptidase 1D, mitochondrial | METAP1D | 1:1 | Block_4 | 0.50466 | 9.5047 | TRUE | NA | NA | NA | 1.1777 | 3.43 | 1.2171 | 3.3749 | 4.7852
P50579 | Methionine aminopeptidase 2 | METAP2 | 1:1 | Block_4 | -0.45623 | 6.1155 | FALSE | 21.7993 | 13.5488 | NA | 1.5612 | 3.1468 | 1.3626 | 3.4748 | 4.5789
P16455 | Methylated-DNA--protein-cysteine methyltransferase | MGMT | 1:1 | Block_4 | -1.68 | 6.7571 | TRUE | 11.5171 | 11.1861 | -1.4665 | 1.1077 | 4.6953 | 1.4382 | 4.8221 | 6.1881
Q29980_Q29983 | MHC class I polypeptide-related sequence A_MHC class I polypeptide-related sequence B | MICA_MICB | 1:1 | Block_4 | -3.8718 | 3.7041 | TRUE | 4.0532 | 8.4542 | -3.0798 | -0.39599 | 0.62796 | -1.7452 | 0.48156 | 1.3414
A9UHW6 | MIF4G domain-containing protein | MIF4GD | 1:1 | Block_4 | -0.70426 | 9.8848 | FALSE | 16.7346 | 11.2485 | NA | 2.1698 | 4.3416 | 1.6768 | 4.3501 | 5.897
Q8N5J2 | Ubiquitin carboxyl-terminal hydrolase MINDY-1 | MINDY1 | 1:1 | Block_4 | 0.66 | 7.297 | TRUE | 12.4548 | 17.4232 | NA | NA | 2.7071 | NA | 2.2333 | 4.9212
Q8N4C8 | Misshapen-like kinase 1 | MINK1 | 1:1 | Block_4 | 0.55 | 10.2306 | TRUE | 15.3085 | 11.1564 | 0.58178 | 2.4497 | 6.643 | 2.8598 | 6.6951 | 8.0679
Q8WV92 | MIT domain-containing protein 1 | MITD1 | 1:1 | Block_4 | 0.044178 | 9.0664 | TRUE | 10.0433 | 5.4577 | NA | 1.1713 | 4.6039 | 1.1689 | 4.3805 | 5.896
Q9BUB5 | MAP kinase-interacting serine/threonine-protein kinase 1 | MKNK1 | 1:1 | Block_4 | -0.51238 | 8.9166 | FALSE | 6.246 | 4.1717 | -0.20207 | 1.768 | 3.9645 | 0.57153 | 3.8093 | 5.3436
P12872 | Promotilin | MLN | 1:1 | Block_4 | -2.2003 | 4.987 | TRUE | 11.5269 | 10.9875 | -0.094732 | 1.5689 | 2.4934 | 0.46567 | 2.3688 | 4.0853
P08473 | Neprilysin | MME | 1:1 | Block_4 | -4 | 2.9814 | TRUE | 11.2738 | 7.12 | -1.8322 | -0.87483 | 0.047657 | -1.3275 | -0.31961 | 1.228
P39900 | Macrophage metalloelastase | MMP12 | 1:1 | Block_4 | -6.81 | 6.852 | TRUE | 6.1557 | 4.7694 | -0.71449 | 0.25582 | 1.166 | -0.21699 | 1.0195 | 2.4957
Q16653 | Myelin-oligodendrocyte glycoprotein | MOG | 1:1 | Block_4 | -1.4793 | 4.4415 | TRUE | 15.5128 | 5.6186 | -0.047333 | 0.3448 | 0.72933 | 0.28266 | 0.92394 | 1.8703
Q14149 | MORC family CW-type zinc finger protein 3 | MORC3 | 1:1 | Block_4 | -1.6222 | 5.5415 | FALSE | 12.7737 | 8.9251 | NA | -1.007 | -0.059985 | -0.92312 | 0.19225 | 2.3163
Q99549 | M-phase phosphoprotein 8 | MPHOSPH8 | 1:1 | Block_4 | -0.0023056 | 10.8522 | TRUE | NA | 6.6106 | NA | 0.47185 | 2.4902 | 1.2309 | 2.9145 | 4.6738
P34949 | Mannose-6-phosphate isomerase | MPI | 1:1 | Block_4 | 0.56303 | 6.8533 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5T2T1 | MAGUK p55 subfamily member 7 | MPP7 | 1:1 | Block_4 | -2.32 | 10.2476 | FALSE | 13.5052 | 15.202 | -0.24265 | 1.3493 | 3.5908 | 1.5312 | 3.5225 | 5.0027
Q6WCQ1 | Myosin phosphatase Rho-interacting protein | MPRIP | 1:1 | Block_4 | 0.3377 | 8.5779 | TRUE | 1.5676 | 13.3771 | NA | 0.38505 | 1.0906 | NA | 0.893 | 1.8086
Q96A35 | 39S ribosomal protein L24, mitochondrial | MRPL24 | 1:1 | Block_4 | 0.54925 | 8.6081 | FALSE | 0.33863 | 1.6121 | NA | NA | 1.4109 | NA | 0.80428 | 1.6564
Q13084 | 39S ribosomal protein L28, mitochondrial | MRPL28 | 1:1 | Block_4 | 0.55433 | 6.1862 | FALSE | NA | NA | NA | NA | 0.59661 | NA | NA | 0.62274
Q14197 | Peptidyl-tRNA hydrolase ICT1, mitochondrial | MRPL58 | 1:1 | Block_4 | 1.1215 | 8.6062 | FALSE | NA | NA | NA | NA | 3.8854 | NA | 2.2698 | 4.4397
P11836 | B-lymphocyte antigen CD20 | MS4A1 | 1:1 | Block_4 | -2.5494 | 6.3291 | TRUE | 10.7659 | 7.6318 | -0.48986 | 1.4261 | 4.1066 | -1.3372 | 0.76092 | 2.7344
Q13421 | Mesothelin | MSLN | 1:1 | Block_4 | -3.3328 | 4.9692 | FALSE | 11.25 | 7.491 | -0.75796 | 0.42843 | 1.8543 | -0.95181 | 0.16344 | 2.0629
O14793 | Growth/differentiation factor 8 | MSTN | 1:1 | Block_4 | -2.1284 | 3.6677 | TRUE | 20.7979 | 9.9501 | -0.58816 | -0.09656 | 1.0129 | -1.2213 | -0.083438 | 0.81912
Q86UE4 | Protein LYRIC | MTDH | 1:1 | Block_4 | 0.8409 | 7.901 | TRUE | NA | NA | NA | NA | 2.7327 | NA | 2.5333 | 4.1628
P13995 | Bifunctional methylenetetrahydrofolate dehydrogenase/cyclohydrolase, mitochondrial | MTHFD2 | 1:1 | Block_4 | 0.71482 | 8.2972 | TRUE | NA | NA | NA | NA | 2.3309 | NA | 1.8468 | 3.1355
Q9H2K0 | Translation initiation factor IF-3, mitochondrial | MTIF3 | 1:1 | Block_4 | 1.8576 | 8.4405 | FALSE | NA | NA | NA | NA | 5.7536 | NA | 4.6553 | 6.5505
P58546 | Myotrophin | MTPN | 1:1 | Block_4 | -2.5114 | 6.3099 | TRUE | NA | NA | NA | NA | -2.1222 | NA | -2.2347 | -0.71041
O43312 | Protein MTSS 1 | MTSS1 | 1:1 | Block_4 | 0.97752 | 11.1249 | FALSE | NA | NA | NA | 1.3563 | 4.0743 | NA | 2.9082 | 5.2303
Q765P7 | Protein MTSS 2 | MTSS2 | 1:1 | Block_4 | 0.70959 | 7.7223 | TRUE | 2.8211 | 2.4336 | NA | NA | 3.0078 | NA | 2.7388 | 4.9413
Q8WXI7 | Mucin-16 | MUC16 | 1:1 | Block_4 | -3.56 | 4.113 | FALSE | 8.5771 | 19.7365 | -1.5358 | 0.14537 | 1.39 | -0.73317 | 0.39688 | 3.0629
Q14324 | Myosin-binding protein C, fast-type | MYBPC2 | 1:1 | Block_4 | -1.312 | 8.8728 | TRUE | 8.5067 | 7.1786 | -1.032 | -0.27118 | 1.1366 | -0.8896 | 0.33371 | 3.362
O75592 | E3 ubiquitin-protein ligase MYCBP2 | MYCBP2 | 1:1 | Block_4 | 0.24745 | 10.8207 | TRUE | 9.9696 | 4.2751 | NA | 0.29635 | 1.2843 | NA | 0.99858 | 2.8529
P12829 | Myosin light chain 4 | MYL4 | 1:1 | Block_4 | -0.99286 | 7.5046 | TRUE | NA | NA | NA | NA | -0.40499 | NA | -0.73004 | 0.93335
P41227 | N-alpha-acetyltransferase 10 | NAA10 | 1:1 | Block_4 | 1.6418 | 7.5894 | FALSE | NA | NA | NA | NA | 2.9175 | NA | 3.3563 | 5.2086
Q02083 | N-acylethanolamine-hydrolyzing acid amidase | NAAA | 1:1 | Block_4 | -0.66509 | 3.9998 | FALSE | 18.1113 | 11.7848 | -0.62577 | 0.10561 | 0.83452 | -0.3104 | 0.17268 | 0.94474
Q96RE7 | Nucleus accumbens-associated protein 1 | NACC1 | 1:1 | Block_4 | -0.94685 | 10.5239 | TRUE | 8.7606 | 9.7587 | -0.54298 | 0.2366 | 0.81451 | 0.23798 | 1.424 | 2.7524
P17050 | Alpha-N-acetylgalactosaminidase | NAGA | 1:1 | Block_4 | 0.29504 | 5.0569 | FALSE | NA | 7.21 | NA | 0.44634 | 1.1149 | NA | 0.886 | 1.8664
Q9UJ70 | N-acetyl-D-glucosamine kinase | NAGK | 1:1 | Block_4 | -0.59179 | 5.7059 | TRUE | 11.0048 | 7.3244 | -0.18149 | 1.1752 | 2.7866 | 1.0382 | 2.4927 | 3.7309
Q9NR45 | Sialic acid synthase | NANS | 1:1 | Block_4 | -0.60146 | 5.859 | FALSE | 9.8723 | 7.2113 | 0.53985 | 1.4887 | 3.1101 | 1.9353 | 3.6115 | 4.6728
Q99733 | Nucleosome assembly protein 1-like 4 | NAP1L4 | 1:1 | Block_4 | -1.2472 | 5.6206 | FALSE | 7.5183 | 3.4163 | -0.42607 | 2.0938 | 3.324 | 1.1644 | 3.1342 | 3.8477
Q6XQN6 | Nicotinate phosphoribosyltransferase | NAPRT | 1:1 | Block_4 | 0.01 | 5.8661 | FALSE | 26.6426 | 24.8773 | 0.19403 | 1.5741 | 2.6126 | 1.2279 | 2.6083 | 3.8437
O43776 | Asparagine--tRNA ligase, cytoplasmic | NARS1 | 1:1 | Block_4 | 1.5929 | 8.2886 | FALSE | NA | NA | NA | NA | 2.4433 | NA | 2.6005 | 3.4263
P19878 | Neutrophil cytosol factor 2 | NCF2 | 1:1 | Block_4 | -4.6687 | 6.4034 | TRUE | 9.3966 | 3.6041 | -1.9219 | -1.1711 | 0.82216 | -1.6964 | -0.44277 | 5.1785
O76036 | Natural cytotoxicity triggering receptor 1 | NCR1 | 1:1 | Block_4 | -1.4372 | 7.7623 | TRUE | 10.2538 | 4.6291 | -0.27561 | 0.30934 | 1.0579 | -0.035189 | 0.98925 | 2.118
Q68D85 | Natural cytotoxicity triggering receptor 3 ligand 1 | NCR3LG1 | 1:1 | Block_4 | -3.34 | 4.9633 | TRUE | 15.8791 | 14.4268 | -0.5057 | 0.20345 | 0.7739 | 0.035261 | 0.98043 | 2.3477
P62166 | Neuronal calcium sensor 1 | NCS1 | 1:1 | Block_4 | -5.2996 | 4.7391 | TRUE | 6.398 | 3.1822 | -0.56895 | 0.28573 | 0.67129 | 0.14188 | 0.90649 | 1.7964
P52848 | Bifunctional heparan sulfate N-deacetylase/N-sulfotransferase 1 | NDST1 | 1:1 | Block_4 | 1.6315 | 11.0833 | FALSE | NA | NA | NA | NA | 1.654 | NA | NA | 1.7106
Q16718 | NADH dehydrogenase [ubiquinone] 1 alpha subcomplex subunit 5 | NDUFA5 | 1:1 | Block_4 | 0.60132 | 4.8302 | FALSE | NA | NA | NA | NA | 0.72139 | NA | 0.69414 | 1.736
P17568 | NADH dehydrogenase [ubiquinone] 1 beta subcomplex subunit 7 | NDUFB7 | 1:1 | Block_4 | 0.45566 | 9.2974 | FALSE | NA | 22.2418 | NA | 0.93104 | 3.7624 | 0.89949 | 2.912 | 5.061
Q15223 | Nectin-1 | NECTIN1 | 1:1 | Block_4 | 0.80144 | 7.5896 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.3767
Q96PU5 | E3 ubiquitin-protein ligase NEDD4-like | NEDD4L | 1:1 | Block_4 | -0.54492 | 4.398 | TRUE | 6.1094 | 7.3434 | 0.18581 | 1.318 | 2.7608 | 0.98382 | 2.461 | 3.5283
Q14511 | Enhancer of filamentation 1 | NEDD9 | 1:1 | Block_4 | 0.36486 | NA | NA | 11.1545 | 8.9715 | NA | NA | 0.40025 | NA | NA | 5.3002
Q92859 | Neogenin | NEO1 | 1:1 | Block_4 | -0.46922 | 6.4385 | TRUE | 11.0917 | 10.8603 | -0.20317 | 0.17104 | 0.45793 | 0.16427 | 0.52594 | 0.91783
O94856 | Neurofascin | NFASC | 1:1 | Block_4 | -0.49443 | 4.462 | TRUE | 11.5468 | 4.2393 | NA | 0.26513 | 1.0105 | -0.077459 | 0.51775 | 1.3254
O94916 | Nuclear factor of activated T-cells 5 | NFAT5 | 1:1 | Block_4 | -0.0081963 | 7.5023 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.2907
O95644 | Nuclear factor of activated T-cells, cytoplasmic 1 | NFATC1 | 1:1 | Block_4 | -0.64023 | 9.4183 | FALSE | 14.2507 | 12.577 | -0.59487 | 0.53618 | 3.0036 | 0.1389 | 2.1915 | 4.1672
Q16621 | Transcription factor NF-E2 45 kDa subunit | NFE2 | 1:1 | Block_4 | 1.0005 | 6.672 | FALSE | NA | NA | NA | NA | 2.5761 | NA | 2.5989 | 4.3551
Q9UMS0 | NFU1 iron-sulfur cluster scaffold homolog, mitochondrial | NFU1 | 1:1 | Block_4 | -1 | 11.6738 | TRUE | 32.7503 | 19.5066 | -0.20804 | 2.3193 | 7.0495 | 3.4696 | 6.5232 | 8.1472
Q9NPE2 | Neugrin | NGRN | 1:1 | Block_4 | 1.7047 | 2.078 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H9Q4 | Non-homologous end-joining factor 1 | NHEJ1 | 1:1 | Block_4 | -1.0104 | 6.9112 | FALSE | 8.0253 | 9.9699 | -0.27901 | 0.73371 | 2.7742 | 1.1846 | 2.6273 | 3.9622
Q86X76 | Deaminated glutathione amidase | NIT1 | 1:1 | Block_4 | -3.77 | 8.2769 | FALSE | 13.2314 | 6.6897 | 0.60904 | 2.626 | 4.2626 | 1.6208 | 3.9847 | 4.7789
Q13287 | N-myc-interactor | NMI | 1:1 | Block_4 | -0.045889 | 9.8517 | FALSE | NA | 2.3396 | NA | 0.23475 | 1.2736 | 0.46164 | 1.5146 | 2.809
Q9HAN9 | Nicotinamide/nicotinic acid mononucleotide adenylyltransferase 1 | NMNAT1 | 1:1 | Block_4 | -1.819 | 9.0077 | TRUE | 8.0623 | 7.0768 | -1.4733 | -0.88644 | 0.64174 | -1.2509 | -0.013331 | 1.8893
P30419 | Glycylpeptide N-tetradecanoyltransferase 1 | NMT1 | 1:1 | Block_4 | 1.09 | 10.9339 | FALSE | NA | NA | NA | 1.1559 | 4.0226 | NA | 4.0082 | 5.5091
O60500 | Nephrin | NPHS1 | 1:1 | Block_4 | -0.017784 | 5.7448 | FALSE | 9.1512 | 16.3338 | NA | 0.10805 | 0.64301 | NA | 0.044459 | 0.7219
P06748 | Nucleophosmin | NPM1 | 1:1 | Block_4 | -1.57 | 3.4949 | FALSE | 5.9227 | 10.4995 | NA | -1.332 | -0.39126 | -1.3455 | -0.38359 | 2.268
P23582 | C-type natriuretic peptide | NPPC | 1:1 | Block_4 | -1.2422 | 7.9677 | TRUE | 75.4926 | 17.458 | -1.0801 | 0.87031 | 1.5051 | -1.049 | 1.2898 | 2.8056
P01303 | Pro-neuropeptide Y | NPY | 1:1 | Block_4 | -1.9252 | 4.8022 | FALSE | 16.0796 | 12.2699 | 0.26333 | 2.0633 | 3.2285 | 0.66432 | 2.2846 | 3.3007
Q9BXI3 | Cytosolic 5'-nucleotidase 1A | NT5C1A | 1:1 | Block_4 | -1.2271 | 8.4105 | TRUE | 13.4456 | 11.8414 | -0.43905 | 0.41085 | 0.96056 | 0.18228 | 1.197 | 2.0971
Q9H0P0 | Cytosolic 5'-nucleotidase 3A | NT5C3A | 1:1 | Block_4 | -0.10017 | 7.3998 | TRUE | 2.7795 | 2.2162 | NA | 1.4376 | 5.2384 | 1.9306 | 5.3316 | 6.5555
P20783 | Neurotrophin-3 | NTF3 | 1:1 | Block_4 | -0.23321 | 9.5609 | TRUE | 18.0673 | 5.9749 | NA | 0.34025 | 0.91525 | NA | 0.35094 | 1.1302
Q9Y5A7 | NEDD8 ultimate buster 1 | NUB1 | 1:1 | Block_4 | 0.21011 | 9.698 | TRUE | 6.1479 | 5.4362 | NA | 0.98932 | 2.4657 | 0.75004 | 2.4749 | 4.2648
P53384 | Cytosolic Fe-S cluster assembly factor NUBP1 | NUBP1 | 1:1 | Block_4 | 1.6555 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
P80303 | Nucleobindin-2 | NUCB2 | 1:1 | Block_4 | -3.4984 | 4.809 | TRUE | 10.2313 | 5.9651 | -0.70849 | 0.13577 | 1.215 | 0.85059 | 1.473 | 2.496
Q9Y266 | Nuclear migration protein nudC | NUDC | 1:1 | Block_4 | -2.5629 | 5.4799 | TRUE | 18.248 | 10.5817 | -1.4435 | -0.35632 | 1.5271 | 0.40515 | 1.6539 | 2.5443
Q8IVD9 | NudC domain-containing protein 3 | NUDCD3 | 1:1 | Block_4 | -1.5603 | 5.8845 | FALSE | 11.4966 | 7.506 | -0.19435 | 1.0556 | 2.7645 | 1.9058 | 3.1228 | 4.2654
Q8NFP7 | Diphosphoinositol polyphosphate phosphohydrolase 3-alpha | NUDT10 | 1:1 | Block_4 | 2.0588 | 6.1469 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96DE0 | U8 snoRNA-decapping enzyme | NUDT16 | 1:1 | Block_4 | -0.98599 | 7.9334 | TRUE | 12.4232 | 7.3848 | -0.55378 | 1.0956 | 2.7315 | 1.3262 | 3.7773 | 5.7843
P50583 | Bis(5'-nucleosyl)-tetraphosphatase [asymmetrical] | NUDT2 | 1:1 | Block_4 | 1.3234 | 8.7152 | TRUE | NA | NA | NA | NA | 3.1667 | NA | 2.3302 | 3.4654
P49757 | Protein numb homolog | NUMB | 1:1 | Block_4 | -2.36 | 6.6301 | FALSE | 13.7227 | 18.7407 | -0.4611 | 1.009 | 4.362 | 1.038 | 4.3095 | 5.5604
O95157 | Neurexophilin-3 | NXPH3 | 1:1 | Block_4 | -2.2993 | 6.8399 | TRUE | 7.2482 | 6.8733 | -0.17329 | 0.49629 | 0.89222 | -0.12134 | 0.53899 | 1.3735
Q9NPH6 | Odorant-binding protein 2b | OBP2B | 1:1 | Block_4 | -3.4083 | 7.0709 | TRUE | 8.805 | 4.8419 | -1.3041 | -0.33681 | 0.94106 | -0.97515 | 0.015828 | 1.077
Q16625 | Occludin | OCLN | 1:1 | Block_4 | -2.6187 | 8.6999 | TRUE | 10.2689 | 4.9909 | -0.17323 | 0.45882 | 1.2632 | 0.40911 | 1.2744 | 2.6356
O60502 | Protein O-GlcNAcase | OGA | 1:1 | Block_4 | -1.6684 | 6.2604 | FALSE | 7.8453 | 13.2169 | -0.97057 | 0.28506 | 2.4241 | 0.79997 | 2.4577 | 4.0468
Q9NZT2 | Opioid growth factor receptor | OGFR | 1:1 | Block_4 | -1.5 | 5.3082 | FALSE | 16.7368 | 11.1455 | -0.11522 | 0.98133 | 1.7288 | 0.61938 | 2.0285 | 3.0037
P23515 | Oligodendrocyte-myelin glycoprotein | OMG | 1:1 | Block_4 | -2.4118 | 7.9399 | TRUE | 9.7917 | 6.6856 | -1.1261 | 0.057002 | 0.70803 | -1.3852 | -0.17416 | 1.0188
O60890 | Oligophrenin-1 | OPHN1 | 1:1 | Block_4 | -4.3981 | 8.7598 | TRUE | 16.9218 | 16.7306 | -4.3299 | -2.4251 | 0.76982 | -2.5547 | 1.1528 | 2.8851
O14841 | 5-oxoprolinase | OPLAH | 1:1 | Block_4 | -0.48298 | 5.5476 | FALSE | 8.9717 | 8.934 | NA | 1.4337 | 3.7899 | 1.0062 | 3.532 | 5.22
Q9UBM4 | Opticin | OPTC | 1:1 | Block_4 | -0.68233 | 8.0298 | TRUE | 19.9437 | 9.9735 | -0.6339 | 0.028875 | 0.54759 | -0.44463 | 0.21204 | 0.92383
Q9H1P3 | Oxysterol-binding protein-related protein 2 | OSBPL2 | 1:1 | Block_4 | 0.72817 | 8.3721 | FALSE | NA | NA | NA | NA | 1.451 | NA | 0.84278 | 2.2259
P13725 | Oncostatin-M | OSM | 1:1 | Block_4 | -1.6816 | 11.3094 | TRUE | 10.5724 | 9.7223 | -0.12462 | 1.5864 | 2.4977 | 0.43558 | 2.2783 | 4.8586
Q8N6M0 | Deubiquitinase OTUD6B | OTUD6B | 1:1 | Block_4 | 2.3681 | 6.3382 | FALSE | NA | NA | NA | NA | 3.1224 | NA | 3.2024 | 4.2805
Q6GQQ9 | OTU domain-containing protein 7B | OTUD7B | 1:1 | Block_4 | -0.50276 | 9.4962 | FALSE | 14.9991 | 15.774 | -0.34174 | 0.675 | 3.1681 | 1.596 | 2.9549 | 4.2118
P55809 | Succinyl-CoA:3-ketoacid coenzyme A transferase 1, mitochondrial | OXCT1 | 1:1 | Block_4 | -0.39184 | 7.186 | FALSE | 7.9439 | 12.1726 | 0.38309 | 1.5826 | 2.7581 | 0.93456 | 2.2661 | 4.0254
Q86VP3 | Phosphofurin acidic cluster sorting protein 2 | PACS2 | 1:1 | Block_4 | 0.49907 | 14.0645 | TRUE | 6.4049 | 11.1278 | 0.50697 | 2.0074 | 3.787 | 2.2192 | 3.9703 | 5.8633
P09466 | Glycodelin | PAEP | 1:1 | Block_4 | -5.662 | 3.3353 | TRUE | 8.7625 | 6.6857 | NA | NA | -2.4951 | NA | -5.5223 | -3.2816
Q15102 | Platelet-activating factor acetylhydrolase IB subunit alpha1 | PAFAH1B3 | 1:1 | Block_4 | -3.5981 | NA | NA | NA | NA | NA | NA | NA | NA | NA | -3.5406
Q9NWQ8 | Phosphoprotein associated with glycosphingolipid-enriched microdomains 1 | PAG1 | 1:1 | Block_4 | -0.066612 | 5.9914 | TRUE | 9.3197 | 5.8903 | NA | 0.27818 | 0.89444 | 0.24637 | 1.0658 | 3.1754
Q9BTK6 | PAXIP1-associated glutamate-rich protein 1 | PAGR1 | 1:1 | Block_4 | 0.66394 | 9.2267 | TRUE | NA | NA | NA | NA | 1.156 | NA | 1.3803 | 2.6771
O75781 | Paralemmin-1 | PALM | 1:1 | Block_4 | -3.6879 | 5.1756 | TRUE | 11.8924 | 3.515 | -0.66766 | 0.28252 | 0.63899 | 0.13919 | 0.90326 | 2.0411
Q8IXS6 | Paralemmin-2 | PALM2 | 1:1 | Block_4 | -3.0175 | 9.0504 | TRUE | 7.4304 | 3.2792 | -0.73843 | -0.17899 | 0.4142 | -0.012315 | 0.96294 | 2.3306
A6NDB9 | Paralemmin-3 | PALM3 | 1:1 | Block_4 | 0.34568 | 10.1666 | TRUE | 9.101 | 4.5878 | NA | NA | 0.74845 | NA | 0.48246 | 1.8131
Q96AQ6 | Pre-B-cell leukemia transcription factor-interacting protein 1 | PBXIP1 | 1:1 | Block_4 | -1.2213 | 6.4027 | FALSE | 9.5464 | 6.3485 | 0.24182 | 0.67873 | 0.97057 | 0.3111 | 0.83201 | 1.4182
Q15366 | Poly(rC)-binding protein 2 | PCBP2 | 1:1 | Block_4 | 0.46795 | 4.5311 | FALSE | 2.4562 | 11.1931 | NA | 1.4889 | 3.5133 | 0.72406 | 3.9738 | 5.6243
Q9NPG4 | Protocadherin-12 | PCDH12 | 1:1 | Block_4 | -2.9653 | 2.2462 | FALSE | 10.8566 | 11.4743 | -0.45164 | 0.36764 | 0.57667 | 0.0082972 | 0.51154 | 0.98739
O60245 | Protocadherin-7 | PCDH7 | 1:1 | Block_4 | -2.1948 | 4.8712 | FALSE | 9.0087 | 9.4915 | -1.1517 | -0.37035 | 0.32877 | -1.1502 | -0.15371 | 0.61738
Q9HC56 | Protocadherin-9 | PCDH9 | 1:1 | Block_4 | -2.24 | 4.5762 | FALSE | 20.3185 | 12.4651 | -0.40904 | 0.45658 | 0.97484 | 0.12228 | 0.71763 | 1.291
Q9Y5E8 | Protocadherin beta-15 | PCDHB15 | 1:1 | Block_4 | -1.5876 | 6.7211 | FALSE | 12.0602 | 10.986 | -0.18696 | 0.40417 | 1.3454 | -0.11302 | 0.69257 | 1.5875
Q16549 | Proprotein convertase subtilisin/kexin type 7 | PCSK7 | 1:1 | Block_4 | -1.5926 | 8.1932 | TRUE | 7.4726 | 6.523 | -0.34885 | 0.41605 | 1.3979 | -0.077193 | 0.7386 | 1.3527
Q99447 | Ethanolamine-phosphate cytidylyltransferase | PCYT2 | 1:1 | Block_4 | -1.5 | 6.2252 | TRUE | 15.2383 | 8.5659 | -1.1912 | 0.84309 | 2.8579 | 0.58711 | 3.387 | 4.406
Q15116 | Programmed cell death protein 1 | PDCD1 | 1:1 | Block_4 | -3.56 | 6.5633 | TRUE | 7.9928 | 5.9546 | -0.59493 | 0.16747 | 1.3529 | -0.048201 | 0.86526 | 2.0737
Q9BQ51 | Programmed cell death 1 ligand 2 | PDCD1LG2 | 1:1 | Block_4 | -2.3571 | 5.5579 | TRUE | 6.7485 | 7.0684 | -0.65185 | -0.02627 | 0.28693 | -0.58377 | 0.13072 | 0.93296
Q96HC4 | PDZ and LIM domain protein 5 | PDLIM5 | 1:1 | Block_4 | 0.27102 | 10.1196 | TRUE | 12.2253 | 8.4891 | NA | 1.148 | 5.2274 | 1.7926 | 4.5713 | 6.345
Q5T2W1 | Na(+)/H(+) exchange regulatory cofactor NHE-RF3 | PDZK1 | 1:1 | Block_4 | -0.53297 | 4.0145 | FALSE | 12.3212 | 10.5995 | NA | -0.38658 | 0.27652 | NA | 0.054164 | 1.6509
Q9BY49 | Peroxisomal trans-2-enoyl-CoA reductase | PECR | 1:1 | Block_4 | 2.0994 | 5.2824 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.4481
P56645 | Period circadian protein homolog 3 | PER3 | 1:1 | Block_4 | 1.2609 | 10.2888 | TRUE | NA | NA | NA | NA | 4.1989 | NA | 3.1952 | 6.074
Q9UHV9 | Prefoldin subunit 2 | PFDN2 | 1:1 | Block_4 | 0.19519 | 7.3484 | FALSE | 0.51165 | 6.2735 | NA | NA | 0.85747 | 0.39064 | 1.0744 | 1.881
O60825 | 6-phosphofructo-2-kinase/fructose-2,6-bisphosphatase 2 | PFKFB2 | 1:1 | Block_4 | -2.283 | 8.2255 | FALSE | 7.5252 | 6.0295 | -1.3778 | 0.41665 | 2.0121 | 0.41723 | 3.0588 | 5.2462
P49763 | Placenta growth factor | PGF | 1:1 | Block_4 | -3.3482 | 8.0797 | TRUE | 9.1048 | 4.3559 | -0.40208 | 0.20977 | 0.78506 | 0.43718 | 1.0047 | 2.1975
Q96LB9 | Peptidoglycan recognition protein 3 | PGLYRP3 | 1:1 | Block_4 | -3.6521 | 8.1507 | TRUE | 6.1533 | 4.2882 | -0.99685 | -0.17079 | 0.9448 | -0.79051 | 0.097598 | 1.1744
O75167 | Phosphatase and actin regulator 2 | PHACTR2 | 1:1 | Block_4 | 0.21135 | 11.2481 | TRUE | NA | NA | NA | 1.669 | 5.0337 | 0.95431 | 4.6519 | 6.8137
Q86UU1 | Pleckstrin homology-like domain family B member 1 | PHLDB1 | 1:1 | Block_4 | -2.9289 | 9.6078 | TRUE | 9.1752 | 4.9583 | -0.33491 | 0.38666 | 1.6126 | 0.87367 | 2.0193 | 2.9406
Q8TCT1 | Phosphoethanolamine/phosphocholine phosphatase | PHOSPHO1 | 1:1 | Block_4 | 0.094722 | 7.2701 | TRUE | 11.1548 | 13.3683 | NA | 0.21653 | 0.63045 | 0.18569 | 0.78475 | 1.4612
Q8WXW3 | Progesterone-induced-blocking factor 1 | PIBF1 | 1:1 | Block_4 | 1.1235 | 11.6118 | TRUE | NA | NA | NA | 1.3114 | 3.4426 | 1.195 | 3.5913 | 5.2083
Q6ZUJ8 | Phosphoinositide 3-kinase adapter protein 1 | PIK3AP1 | 1:1 | Block_4 | -1.673 | 8.5299 | FALSE | 20.9113 | 20.0589 | -0.33898 | 1.307 | 2.8429 | 0.89305 | 2.4649 | 4.4025
Q13563 | Polycystin-2 | PKD2 | 1:1 | Block_4 | -1.5986 | 12.0473 | TRUE | 8.2588 | 5.6174 | -0.55368 | 0.25299 | 2.1271 | 0.69486 | 2.1684 | 4.2945
O15496 | Group 10 secretory phospholipase A2 | PLA2G10 | 1:1 | Block_4 | -4.4314 | 8.1158 | TRUE | 8.7804 | 8.3877 | -1.3303 | -0.093625 | 0.63591 | -0.76805 | 0.30731 | 1.6848
Q6P1J6 | Phospholipase B1, membrane-associated | PLB1 | 1:1 | Block_4 | -1.4226 | 7.0815 | TRUE | 7.1288 | 4.2935 | -0.99536 | -0.020962 | 0.74617 | -1.1828 | -0.33078 | 0.60876
Q00722 | 1-phosphatidylinositol 4,5-bisphosphate phosphodiesterase beta-2 | PLCB2 | 1:1 | Block_4 | 1.0183 | 9.6594 | TRUE | NA | NA | NA | 1.1187 | 5.5508 | 1.4458 | 6.1764 | 7.9035
Q53GL0 | Pleckstrin homology domain-containing family O member 1 | PLEKHO1 | 1:1 | Block_4 | 1.3288 | 7.6271 | TRUE | NA | NA | NA | NA | 3.7595 | NA | 3.1852 | 5.5286
Q9NRY6 | Phospholipid scramblase 3 | PLSCR3 | 1:1 | Block_4 | 0.85764 | 8.397 | TRUE | NA | NA | NA | 0.91044 | 2.1871 | 1.1686 | 2.1395 | 3.4104
Q6UX71 | Plexin domain-containing protein 2 | PLXDC2 | 1:1 | Block_4 | -1.0407 | 6.2853 | TRUE | 10.8735 | 20.3041 | -0.12986 | 0.62955 | 1.2948 | 0.095505 | 0.96584 | 1.9151
Q9HCM2 | Plexin-A4 | PLXNA4 | 1:1 | Block_4 | -1.3976 | 7.9522 | TRUE | 14.6739 | 11.6434 | -0.23211 | 0.5787 | 2.1244 | 0.73972 | 2.0323 | 3.6109
P20382 | Pro-MCH | PMCH | 1:1 | Block_4 | -1.4987 | 7.7824 | TRUE | 5.8335 | 7.275 | NA | NA | 0.11674 | NA | -1.4686 | 0.2276
P54315 | Inactive pancreatic lipase-related protein 1 | PNLIPRP1 | 1:1 | Block_4 | -5.3249 | 5.423 | FALSE | 6.6623 | 5.5442 | -1.5648 | -0.0080736 | 1.3255 | -1.0805 | 0.34478 | 1.9933
O00592 | Podocalyxin | PODXL | 1:1 | Block_4 | -4.6636 | 4.046 | TRUE | 7.251 | 3.9004 | -0.22059 | 0.17457 | 0.34163 | -0.16183 | 0.090042 | 0.34474
Q8WVV4 | Protein POF1B | POF1B | 1:1 | Block_4 | 0.49052 | 11.4387 | TRUE | 7.8804 | 2.8519 | NA | NA | 0.96454 | NA | 0.5189 | 1.9079
P61218 | DNA-directed RNA polymerases I, II, and III subunit RPABC2 | POLR2F | 1:1 | Block_4 | -0.03278 | 6.7451 | TRUE | 10.5859 | 7.2768 | NA | NA | 0.23887 | NA | 0.55543 | 2.6029
P01189 | Pro-opiomelanocortin | POMC | 1:1 | Block_4 | -0.22488 | 10.9635 | TRUE | 14.8245 | 4.5881 | 0.44613 | 1.7422 | 2.8389 | 0.21924 | 1.6747 | 3.0857
Q15165 | Serum paraoxonase/arylesterase 2 | PON2 | 1:1 | Block_4 | -3.4186 | 4.1932 | FALSE | 14.6667 | 5.0535 | -0.146 | 0.29547 | 0.69847 | -0.12972 | 0.41536 | 1.0937
O60437 | Periplakin | PPL | 1:1 | Block_4 | -0.52221 | 9.5094 | TRUE | 19.0057 | 9.9941 | -0.38267 | 0.11719 | 1.1096 | 0.18372 | 0.88513 | 2.0378
O75688 | Protein phosphatase 1B | PPM1B | 1:1 | Block_4 | 2.7045 | 11.2558 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y570 | Protein phosphatase methylesterase 1 | PPME1 | 1:1 | Block_4 | 0.38891 | 5.3379 | FALSE | NA | NA | NA | 1.0216 | 2.7879 | 1.2794 | 2.5746 | 3.6505
P36873 | Serine/threonine-protein phosphatase PP1-gamma catalytic subunit | PPP1CC | 1:1 | Block_4 | -0.051552 | 6.0741 | FALSE | 6.5634 | 10.7893 | NA | 0.69619 | 2.6046 | 0.57231 | 2.5601 | 4.5655
Q96A00 | Protein phosphatase 1 regulatory subunit 14A | PPP1R14A | 1:1 | Block_4 | 1.4628 | NA | NA | NA | NA | NA | NA | 2.4436 | NA | 2.1157 | 3.6328
Q9NXH3 | Protein phosphatase 1 regulatory subunit 14D | PPP1R14D | 1:1 | Block_4 | -0.11878 | 8.7727 | TRUE | 8.3258 | 7.1726 | NA | 0.078257 | 1.1134 | NA | 0.68473 | 2.2983
Q6NYC8 | Phostensin | PPP1R18 | 1:1 | Block_4 | -2.3182 | 7.9266 | TRUE | 5.9171 | 6.9367 | 0.078713 | 1.4167 | 2.2294 | 0.44163 | 2.2084 | 3.3872
Q96SB3 | Neurabin-2 | PPP1R9B | 1:1 | Block_4 | 0.84977 | 7.3341 | TRUE | NA | NA | NA | 0.9806 | 4.2168 | 1.3127 | 3.6369 | 5.166
Q15172 | Serine/threonine-protein phosphatase 2A 56 kDa regulatory subunit alpha isoform | PPP2R5A | 1:1 | Block_4 | 2.0215 | 8.6576 | FALSE | 30.7279 | 14.214 | NA | NA | 5.0746 | NA | 5.678 | 7.5229
P63098 | Calcineurin subunit B type 1 | PPP3R1 | 1:1 | Block_4 | -0.75139 | 2.1058 | TRUE | 14.3725 | 11.0526 | -0.43106 | 0.53867 | 1.462 | 0.66473 | 1.3905 | 2.1424
Q06830 | Peroxiredoxin-1 | PRDX1 | 1:1 | Block_4 | 0.02498 | 6.6463 | TRUE | 15.6306 | 2.94 | NA | 0.73643 | 1.5593 | 0.51211 | 2.0789 | 3.2409
P30044 | Peroxiredoxin-5, mitochondrial | PRDX5 | 1:1 | Block_4 | -1.3687 | NA | NA | 13.1443 | 15.406 | -0.8177 | 0.98562 | 3.314 | 0.647 | 3.5642 | 4.684
P30041 | Peroxiredoxin-6 | PRDX6 | 1:1 | Block_4 | 1.6937 | 3.4958 | FALSE | NA | NA | NA | 2.0681 | 3.7212 | 2.0487 | 3.635 | 5.1615
P13861 | cAMP-dependent protein kinase type II-alpha regulatory subunit | PRKAR2A | 1:1 | Block_4 | -1.4181 | 7.3817 | FALSE | 11.0207 | 16.687 | -1.2667 | 0.89264 | 2.4114 | 1.1323 | 2.5336 | 3.9273
Q9UKY0 | Prion-like protein doppel | PRND | 1:1 | Block_4 | -1.3487 | 6.2661 | TRUE | 21.5479 | 7.5904 | -0.78454 | -0.094297 | 0.98566 | -0.82347 | 0.64186 | 1.9183
P58294 | Prokineticin-1 | PROK1 | 1:1 | Block_4 | -2.5214 | 7.2904 | TRUE | 10.4692 | 3.7001 | -1.747 | -0.38885 | 1.4865 | -1.4226 | -0.41734 | 1.1567
Q5FWE3 | Proline-rich transmembrane protein 3 | PRRT3 | 1:1 | Block_4 | -4.5729 | 5.0719 | TRUE | 11.357 | 9.7664 | -0.45361 | 0.35353 | 1.1778 | 0.067478 | 1.0446 | 2.4539
Q2L4Q9 | Serine protease 53 | PRSS53 | 1:1 | Block_4 | -0.3664 | 8.1763 | TRUE | 8.5201 | 12.0848 | NA | 0.023323 | 0.941 | NA | 0.10104 | 0.71298
Q2VWP7 | Protogenin | PRTG | 1:1 | Block_4 | -3.3208 | 7.3511 | TRUE | 6.0363 | 6.7289 | -0.011237 | 0.43767 | 1.1319 | 0.18187 | 0.74844 | 1.2699
Q6NUJ1 | Proactivator polypeptide-like 1 | PSAPL1 | 1:1 | Block_4 | -4.02 | 6.7999 | TRUE | 6.9224 | 8.9661 | -0.73657 | -0.060721 | 1.0457 | -0.30767 | 0.70485 | 1.8441
O43653 | Prostate stem cell antigen | PSCA | 1:1 | Block_4 | -4.8 | 4.9941 | TRUE | 10.5422 | 11.4399 | NA | -0.18327 | 1.9221 | NA | 0.16975 | 2.6051
P11464 | Pregnancy-specific beta-1-glycoprotein 1 | PSG1 | 1:1 | Block_4 | -10 | 0.41868 | TRUE | 5.0115 | 4.8475 | -6.9391 | -5.9685 | -4.0087 | -7.851 | -5.3937 | -2.895
O75475 | PC4 and SFRS1-interacting protein | PSIP1 | 1:1 | Block_4 | -1.2116 | 5.9488 | FALSE | 12.7314 | 10.6209 | NA | -0.86264 | 0.12155 | -0.71613 | 0.50364 | 3.0682
Q99460 | 26S proteasome non-ATPase regulatory subunit 1 | PSMD1 | 1:1 | Block_4 | 0.44078 | 11.1106 | TRUE | NA | NA | NA | 1.4002 | 3.3675 | 1.939 | 3.3881 | 5.0422
O00233 | 26S proteasome non-ATPase regulatory subunit 9 | PSMD9 | 1:1 | Block_4 | -0.093473 | 10.0857 | FALSE | 7.9879 | 11.0241 | 0.49022 | 1.73 | 2.562 | 1.4634 | 2.2549 | 3.4438
Q06323 | Proteasome activator complex subunit 1 | PSME1 | 1:1 | Block_4 | -2.6203 | 7.032 | FALSE | 7.3372 | 6.1084 | -0.086514 | 1.2978 | 1.8594 | 1.0019 | 2.0146 | 2.6428
Q9UL46 | Proteasome activator complex subunit 2 | PSME2 | 1:1 | Block_4 | -3.1458 | 6.6516 | TRUE | 7.1468 | 8.2915 | 0.058883 | 1.2537 | 1.9096 | 0.90328 | 1.9003 | 2.8115
Q9BT73 | Proteasome assembly chaperone 3 | PSMG3 | 1:1 | Block_4 | 0.72633 | 8.2895 | TRUE | 4.0902 | 9.3693 | 1.7752 | 3.0834 | 3.9093 | 2.8697 | 4.0533 | 4.9143
Q5JS54 | Proteasome assembly chaperone 4 | PSMG4 | 1:1 | Block_4 | 0.32242 | 10.2115 | TRUE | 5.8318 | 5.5954 | 0.34811 | 1.6167 | 2.6402 | 1.3783 | 2.3146 | 3.4918
O60542 | Persephin | PSPN | 1:1 | Block_4 | -0.92281 | 8.6421 | FALSE | 11.1388 | 8.3548 | NA | -0.51009 | 1.1916 | NA | 0.31265 | 1.9715
O43586 | Proline-serine-threonine phosphatase-interacting protein 1 | PSTPIP1 | 1:1 | Block_4 | -3.0719 | 8.5617 | FALSE | 8.5715 | 5.0137 | -0.89868 | -0.047295 | 0.52644 | -0.52423 | 0.46959 | 4.2697
Q9H7Z7 | Prostaglandin E synthase 2 | PTGES2 | 1:1 | Block_4 | 1.85 | 8.1851 | FALSE | NA | NA | NA | NA | 3.1537 | NA | 3.1311 | 5.1732
P01270 | Parathyroid hormone | PTH | 1:1 | Block_4 | -3.1708 | 8.9351 | FALSE | 13.446 | 11.3241 | -1.0792 | 0.33156 | 1.3709 | -0.96902 | 0.80989 | 2.4784
Q03431 | Parathyroid hormone/parathyroid hormone-related peptide receptor | PTH1R | 1:1 | Block_4 | -1.7958 | 8.829 | TRUE | 9.0894 | 8.8287 | -0.19044 | 0.50373 | 1.1978 | 0.15198 | 0.78864 | 1.6426
Q13308 | Inactive tyrosine-protein kinase 7 | PTK7 | 1:1 | Block_4 | -0.52931 | 8.0824 | TRUE | 9.8544 | 5.7259 | -0.25993 | 0.22384 | 0.74384 | 0.13319 | 0.68325 | 1.7672
P21246 | Pleiotrophin | PTN | 1:1 | Block_4 | 0.61979 | 6.7395 | TRUE | 12.9824 | 10.8421 | NA | 1.0344 | 2.0072 | 0.70852 | 2.1019 | 3.7759
P18031 | Tyrosine-protein phosphatase non-receptor type 1 | PTPN1 | 1:1 | Block_4 | -1.5275 | 4.2132 | FALSE | 11.6315 | 12.3713 | -1.1456 | 0.8796 | 3.9362 | 1.2766 | 3.8512 | 5.5621
Q9HD43 | Receptor-type tyrosine-protein phosphatase H | PTPRH | 1:1 | Block_4 | -1.0287 | 7.9928 | TRUE | 13.9818 | 11.0484 | -0.86603 | -0.059088 | 0.62448 | -0.67177 | 0.46025 | 1.7046
Q15262 | Receptor-type tyrosine-protein phosphatase kappa | PTPRK | 1:1 | Block_4 | -0.021995 | 4.8678 | FALSE | 7.1408 | 6.8609 | NA | 0.12243 | 0.72889 | NA | 0.36754 | 1.1687
Q16849 | Receptor-type tyrosine-protein phosphatase-like N | PTPRN | 1:1 | Block_4 | -5.3989 | 6.0649 | TRUE | 5.3492 | 4.3482 | -0.43184 | 0.16709 | 0.69798 | -0.20635 | 0.6443 | 1.4711
Q92932 | Receptor-type tyrosine-protein phosphatase N2 | PTPRN2 | 1:1 | Block_4 | -2.2371 | 7.961 | TRUE | 7.7936 | 5.1304 | -0.43785 | 0.15307 | 0.78975 | -0.070428 | 0.74273 | 1.9591
Q15256 | Receptor-type tyrosine-protein phosphatase R | PTPRR | 1:1 | Block_4 | -2.4828 | 8.4152 | TRUE | 13.4481 | 3.6771 | -0.66474 | 0.077632 | 0.59311 | -0.18097 | 0.48298 | 1.4614
Q9NZH5 | Securin-2 | PTTG2 | 1:1 | Block_4 | 0.37266 | 10.1905 | TRUE | 1.1604 | NA | NA | NA | 0.65994 | NA | NA | 1.0802
P20472 | Parvalbumin alpha | PVALB | 1:1 | Block_4 | -1.2398 | 4.9151 | FALSE | 8.325 | 3.0874 | -0.28112 | 1.005 | 3.838 | 1.3179 | 3.5663 | 5.1806
A1KZ92 | Probable oxidoreductase PXDNL | PXDNL | 1:1 | Block_4 | -0.48937 | 9.8109 | TRUE | 20.7558 | 7.1915 | NA | -0.026455 | 1.2073 | NA | 0.17537 | 1.1391
Q8WXC3 | Pyrin domain-containing protein 1 | PYDC1 | 1:1 | Block_4 | -0.31215 | 12.1741 | TRUE | 18.6147 | 7.9297 | NA | -0.036881 | 0.74285 | -0.18437 | 0.8317 | 2.1354
Q6K0P9 | Pyrin and HIN domain-containing protein 1 | PYHIN1 | 1:1 | Block_4 | -4.9381 | 8.8273 | TRUE | 32.1781 | 9.8686 | -3.2343 | -1.8376 | -0.65441 | -2.3959 | -1.3868 | 2.511
Q8WUD1 | Ras-related protein Rab-2B | RAB2B | 1:1 | Block_4 | -2.9653 | 6.3499 | FALSE | NA | NA | NA | -2.6018 | -0.94434 | -2.5476 | -0.73637 | 0.69658
Q8TBN0 | Guanine nucleotide exchange factor for Rab-3A | RAB3IL1 | 1:1 | Block_4 | -1.7468 | 9.7665 | TRUE | 11.6511 | 7.608 | -1.1429 | 0.3553 | 1.6135 | -0.14056 | 1.1752 | 2.7145
Q96QF0 | Rab-3A-interacting protein | RAB3IP | 1:1 | Block_4 | -0.088534 | 8.1648 | TRUE | NA | 7.9547 | NA | 0.71512 | 2.7193 | 0.27001 | 2.5331 | 5.4998
Q9H5N1 | Rab GTPase-binding effector protein 2 | RABEP2 | 1:1 | Block_4 | -0.2 | 8.5853 | TRUE | 23.6635 | 6.5333 | NA | 0.84181 | 3.4635 | 1.1695 | 3.0551 | 5.5624
Q5VY80_Q9BZM5 | UL16-binding protein 2/6 | RAET1L_ULBP2 | 1:1 | Block_4 | -1.7392 | 6.2759 | TRUE | 10.2368 | 4.1755 | -0.30775 | 0.23969 | 0.76644 | 0.18964 | 1.093 | 2.3433
P11234 | Ras-related protein Ral-B | RALB | 1:1 | Block_4 | 2.3021 | 6.5967 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.7279
Q9UKM9 | RNA-binding protein Raly | RALY | 1:1 | Block_4 | -0.18428 | 7.6852 | TRUE | NA | 12.2101 | NA | NA | -0.020729 | NA | 0.074634 | 2.0524
P43487 | Ran-specific GTPase-activating protein | RANBP1 | 1:1 | Block_4 | 0.23623 | 7.077 | TRUE | NA | 7.0561 | 0.2977 | 3.1064 | 4.6435 | 2.1668 | 4.4888 | 5.3012
Q9Y4G8 | Rap guanine nucleotide exchange factor 2 | RAPGEF2 | 1:1 | Block_4 | 0.59686 | 11.2357 | FALSE | NA | NA | NA | 0.97529 | 3.263 | NA | 3.134 | 4.7529
Q7LDG7 | RAS guanyl-releasing protein 2 | RASGRP2 | 1:1 | Block_4 | -0.75877 | 10.6572 | FALSE | 11.5859 | 19.0077 | -0.70792 | 1.49 | 5.8677 | 1.3538 | 5.5636 | 8.0029
Q9H477 | Ribokinase | RBKS | 1:1 | Block_4 | -4.1722 | 5.3609 | TRUE | 18.5403 | 8.55 | -0.68992 | 0.66312 | 2.0399 | 0.55183 | 1.4126 | 3.0811
Q96I25 | Splicing factor 45 | RBM17 | 1:1 | Block_4 | 1.3425 | 9.4839 | FALSE | NA | NA | NA | NA | 2.6927 | NA | 2.2051 | 3.5204
P50120 | Retinol-binding protein 2 | RBP2 | 1:1 | Block_4 | -2.9631 | 3.9862 | TRUE | 7.1624 | 6.6512 | -1.2002 | -0.37538 | 1.3111 | -0.69459 | 0.64307 | 2.4023
P82980 | Retinol-binding protein 5 | RBP5 | 1:1 | Block_4 | -1.88 | 3.3537 | TRUE | 25.5739 | 14.3626 | -1.249 | -0.11737 | 0.71588 | -0.41937 | 0.75291 | 1.8282
Q96R05 | Retinoid-binding protein 7 | RBP7 | 1:1 | Block_4 | -1.2668 | 4.2579 | TRUE | 11.9068 | 10.8823 | -1.2516 | -0.39373 | 0.39467 | -0.77066 | 0.65387 | 3.8418
Q93062 | RNA-binding protein with multiple splicing | RBPMS | 1:1 | Block_4 | 0.69494 | 10.7603 | TRUE | NA | NA | NA | NA | 1.3253 | NA | NA | 2.43
Q6ZRY4 | RNA-binding protein with multiple splicing 2 | RBPMS2 | 1:1 | Block_4 | -0.31422 | 8.7159 | TRUE | NA | 6.919 | NA | 1.6276 | 4.4766 | 1.608 | 4.1096 | 6.3997
O95980 | Reversion-inducing cysteine-rich protein with Kazal motifs | RECK | 1:1 | Block_4 | -3.7955 | 2.8604 | TRUE | 5.6036 | 3.7296 | -0.3112 | 0.14319 | 0.45382 | -0.3266 | 0.18294 | 0.63945
Q9H6H4 | Receptor expression-enhancing protein 4 | REEP4 | 1:1 | Block_4 | 0.52832 | 6.6199 | FALSE | 7.4663 | 3.8919 | NA | NA | 1.487 | NA | 1.5297 | 3.2301
Q6UW15 | Regenerating islet-derived protein 3-gamma | REG3G | 1:1 | Block_4 | -0.87987 | 7.89 | TRUE | 9.8321 | 9.2288 | -0.76373 | -0.13436 | 0.69907 | -0.62616 | 0.1566 | 1.1196
P07949 | Proto-oncogene tyrosine-protein kinase receptor Ret | RET | 1:1 | Block_4 | -4.4091 | 6.0562 | TRUE | 8.0626 | 2.8623 | -1.0562 | -0.34886 | 0.65281 | -1.0409 | -0.043393 | 1.07
P35249 | Replication factor C subunit 4 | RFC4 | 1:1 | Block_4 | -1.6151 | 9.4834 | FALSE | 12.5179 | 8.9504 | -1.3316 | 0.21334 | 1.8325 | 0.73139 | 2.3073 | 3.9642
Q9H4X1 | Regulator of cell cycle RGCC | RGCC | 1:1 | Block_4 | -0.076807 | 7.0088 | FALSE | 9.2755 | 5.6079 | NA | 0.55977 | 3.6195 | 1.3447 | 3.4704 | 4.9348
O43665 | Regulator of G-protein signaling 10 | RGS10 | 1:1 | Block_4 | 1.42 | 7.987 | TRUE | NA | NA | NA | NA | 3.5438 | NA | 3.4394 | 5.1371
P08134 | Rho-related GTP-binding protein RhoC | RHOC | 1:1 | Block_4 | -1.07 | 6.7901 | TRUE | 15.1488 | 15.0469 | 0.024254 | 1.432 | 4.8048 | 1.9542 | 4.7052 | 5.9571
Q96NA2 | Rab-interacting lysosomal protein | RILP | 1:1 | Block_4 | -1.3245 | 7.7096 | TRUE | 19.4141 | 10.284 | 0.14526 | 2.4033 | 4.3919 | 2.3571 | 4.0351 | 5.533
P04808 | Prorelaxin H1 | RLN1 | 1:1 | Block_4 | -4.483 | 5.2977 | FALSE | 16.7725 | 10.8393 | NA | -3.6388 | 1.6083 | -4.463 | -2.706 | -0.26424
Q5GAN6 | Inactive ribonuclease-like protein 10 | RNASE10 | 1:1 | Block_4 | -1.8192 | 9.4442 | TRUE | 10.7746 | 9.5211 | -0.46695 | 0.21374 | 1.2529 | 0.022338 | 1.1375 | 2.8691
O75792 | Ribonuclease H2 subunit A | RNASEH2A | 1:1 | Block_4 | 1 | 3.0337 | FALSE | NA | NA | NA | NA | NA | NA | 1.0906 | 2.3951
Q8NC42 | E3 ubiquitin-protein ligase RNF149 | RNF149 | 1:1 | Block_4 | -1.6507 | 5.7696 | TRUE | 9.0741 | 8.5061 | -0.26501 | 0.36837 | 1.0267 | 0.2975 | 1.1456 | 2.2976
Q8IYW5 | E3 ubiquitin-protein ligase RNF168 | RNF168 | 1:1 | Block_4 | 0.74077 | 8.3575 | FALSE | NA | NA | NA | NA | NA | NA | NA | 0.81268
Q99942 | E3 ubiquitin-protein ligase RNF5 | RNF5 | 1:1 | Block_4 | 1.3291 | 11.0021 | FALSE | NA | NA | NA | NA | 2.6594 | NA | 2.9509 | 4.4501
Q8WZ75 | Roundabout homolog 4 | ROBO4 | 1:1 | Block_4 | -2.0181 | 2.4889 | TRUE | 8.6415 | 11.4129 | -0.18421 | 0.23311 | 0.49374 | 0.1904 | 0.50412 | 0.84
O75695 | Protein XRP2 | RP2 | 1:1 | Block_4 | 0.45362 | 5.2629 | FALSE | 21.843 | 2.4816 | NA | 0.8697 | 2.284 | 0.85541 | 2.1804 | 3.4551
P50914 | 60S ribosomal protein L14 | RPL14 | 1:1 | Block_4 | -2 | 5.6876 | FALSE | 18.5107 | 12.3619 | -1.5159 | 0.0037193 | 0.48687 | -0.64073 | 0.84348 | 3.1167
Q2MKA7 | R-spondin-1 | RSPO1 | 1:1 | Block_4 | -1.9658 | 10.2876 | TRUE | 7.2048 | 4.2008 | 0.39249 | 1.3749 | 2.2641 | 0.99003 | 2.2742 | 3.3649
Q9BXY4 | R-spondin-3 | RSPO3 | 1:1 | Block_4 | -3 | 9.2672 | TRUE | 10.7824 | 10.2413 | 0.19546 | 1.2483 | 2.008 | 0.69226 | 2.0678 | 3.0038
Q9BSG5 | Retbindin | RTBDN | 1:1 | Block_4 | -2.1812 | 9.2908 | TRUE | 8.5787 | 6.8605 | -0.46122 | 0.60018 | 1.2076 | 0.11254 | 0.98551 | 2.0033
Q8WWV3 | Reticulon-4-interacting protein 1, mitochondrial | RTN4IP1 | 1:1 | Block_4 | -1 | 12.4942 | FALSE | 27.0778 | 10.2954 | -0.57521 | 2.7991 | 5.7082 | 2.0484 | 5.0868 | 6.931
Q9BZR6 | Reticulon-4 receptor | RTN4R | 1:1 | Block_4 | -5 | 5.5691 | FALSE | 14.9814 | 12.6412 | -0.51091 | 0.17658 | 0.81453 | -0.089373 | 0.78664 | 1.5295
P80511 | Protein S100-A12 | S100A12 | 1:1 | Block_4 | -2.2667 | 2.5893 | FALSE | 9.4304 | 8.0989 | -0.54659 | 0.6567 | 1.8651 | -0.45054 | 1.2488 | 3.056
Q99584 | Protein S100-A13 | S100A13 | 1:1 | Block_4 | -4.1158 | 8.8848 | TRUE | 6.6573 | 2.3108 | -1.4463 | -0.080082 | 0.98188 | -1.3585 | 0.54811 | 1.5609
Q96FQ6 | Protein S100-A16 | S100A16 | 1:1 | Block_4 | -2.2267 | 5.7639 | FALSE | 10.5481 | 7.6215 | -0.53272 | 0.22004 | 1.8152 | -1.0079 | -0.08555 | 1.14
P29377 | Protein S100-G | S100G | 1:1 | Block_4 | 0.30479 | 9.2829 | TRUE | NA | NA | NA | NA | 1.0096 | NA | NA | 1.2941
P25815 | Protein S100-P | S100P | 1:1 | Block_4 | -0.77471 | 4.0573 | TRUE | 10.6283 | 7.0946 | -0.17734 | 0.39276 | 3.454 | -0.17543 | 1.4302 | 3.8516
Q14151 | Scaffold attachment factor B2 | SAFB2 | 1:1 | Block_4 | 0.0012144 | 7.9416 | TRUE | NA | NA | NA | NA | 0.047512 | NA | 0.021047 | 1.4078
Q8IVG5 | Sterile alpha motif domain-containing protein 9-like | SAMD9L | 1:1 | Block_4 | -1.5149 | 9.1323 | FALSE | 7.9759 | 5.0734 | -1.1258 | 0.20376 | 2.3768 | 0.99872 | 3.0753 | 4.5992
O75995 | SAM and SH3 domain-containing protein 3 | SASH3 | 1:1 | Block_4 | -2.1768 | 7.9735 | TRUE | 9.3234 | 13.3406 | -0.87901 | 0.46841 | 1.2802 | 0.064059 | 1.5904 | 5.1579
Q01826 | DNA-binding protein SATB1 | SATB1 | 1:1 | Block_4 | -0.28417 | 9.7534 | FALSE | 6.1618 | 1.3113 | NA | -0.085846 | 0.42507 | 0.02882 | 0.4997 | 1.5987
O15127 | Secretory carrier-associated membrane protein 2 | SCAMP2 | 1:1 | Block_4 | 0.028263 | 8.3566 | TRUE | 16.0049 | 11.7662 | NA | 1.6459 | 4.9105 | 1.7644 | 4.3147 | 5.9526
O14828 | Secretory carrier-associated membrane protein 3 | SCAMP3 | 1:1 | Block_4 | -2.391 | 8.4774 | TRUE | 8.9935 | 3.0918 | -0.37488 | 1.0281 | 3.5719 | 1.9574 | 3.7294 | 5.4505
Q14108 | Lysosome membrane protein 2 | SCARB2 | 1:1 | Block_4 | -3.0879 | 6.2405 | TRUE | 11.1369 | 5.4161 | -0.32229 | 0.19862 | 0.60965 | 0.10345 | 0.96033 | 2.5305
P13521 | Secretogranin-2 | SCG2 | 1:1 | Block_4 | -3.0386 | 3.5475 | FALSE | 8.4097 | 4.7818 | -0.14708 | 0.61703 | 1.1923 | 0.16084 | 0.89642 | 1.9476
Q96NL6 | Sodium channel and clathrin linker 1 | SCLT1 | 1:1 | Block_4 | 0.10639 | 12.753 | TRUE | NA | NA | NA | NA | 0.99682 | NA | 0.6315 | 2.3649
Q96I15 | Selenocysteine lyase | SCLY | 1:1 | Block_4 | -1.1459 | 5.5859 | TRUE | 10.5855 | 14.9629 | -1.0903 | -0.077249 | 1.4391 | -0.28667 | 0.6502 | 1.5595
Q99250 | Sodium channel protein type 2 subunit alpha | SCN2A | 1:1 | Block_4 | 0.56146 | 8.6548 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IWT1 | Sodium channel subunit beta-4 | SCN4B | 1:1 | Block_4 | -3.2456 | 6.2262 | TRUE | 9.3948 | 10.0799 | -0.4889 | 0.1634 | 0.70062 | -0.39726 | 0.64994 | 1.8357
Q9HB40 | Retinoid-inducible serine carboxypeptidase | SCPEP1 | 1:1 | Block_4 | -0.85745 | 6.6929 | TRUE | 19.4088 | 7.702 | -0.082945 | 1.0606 | 1.7122 | 0.73397 | 1.6578 | 2.6706
Q14160 | Protein scribble homolog | SCRIB | 1:1 | Block_4 | -1.2594 | 7.8231 | TRUE | 7.9194 | 9.5703 | -0.78855 | -0.041338 | 3.2657 | 0.39355 | 1.3208 | 3.0135
Q12765 | Secernin-1 | SCRN1 | 1:1 | Block_4 | -1.5047 | 6.9789 | FALSE | 17.0705 | 12.6514 | -0.35168 | 1.1236 | 3.6102 | 2.2381 | 3.972 | 5.4282
Q86SQ7 | Serologically defined colon cancer antigen 8 | SDCCAG8 | 1:1 | Block_4 | -1.5041 | 13.0173 | TRUE | 19.1512 | 9.7686 | NA | -0.21627 | 2.9225 | 0.20267 | 2.6178 | 5.3542
P21912 | Succinate dehydrogenase [ubiquinone] iron-sulfur subunit, mitochondrial | SDHB | 1:1 | Block_4 | -0.88223 | 6.7737 | FALSE | 15.1103 | 9.6204 | -0.86098 | 0.87635 | 2.5128 | 0.15313 | 1.8854 | 3.5496
O94979 | Protein transport protein Sec31A | SEC31A | 1:1 | Block_4 | -2 | NA | NA | 30.1757 | 15.0503 | 0.27094 | 1.808 | 4.1614 | 1.7661 | 3.9854 | 4.7285
Q9UBV2 | Protein sel-1 homolog 1 | SEL1L | 1:1 | Block_4 | -2.1367 | 8.3976 | FALSE | 10.8913 | 3.6082 | -0.32066 | 0.27556 | 0.91486 | 0.47657 | 1.1459 | 2.2846
Q9C0C4 | Semaphorin-4C | SEMA4C | 1:1 | Block_4 | -0.25732 | 6.7632 | FALSE | 8.8821 | 14.6746 | 0.725 | 1.9253 | 2.4281 | 0.60751 | 1.7748 | 2.9691
Q9H3T2 | Semaphorin-6C | SEMA6C | 1:1 | Block_4 | -0.63133 | 5.7619 | FALSE | NA | NA | NA | NA | NA | NA | NA | -0.60772
Q9UH03 | Neuronal-specific septin-3 | SEPTIN3 | 1:1 | Block_4 | -0.98419 | 7.2421 | FALSE | 6.739 | 3.6131 | -0.80067 | 0.19814 | 1.0295 | -0.80535 | 0.28427 | 1.4049
Q92599 | Septin-8 | SEPTIN8 | 1:1 | Block_4 | -1.4145 | 8.1328 | FALSE | 16.1051 | 12.8733 | 0.36238 | 1.0296 | 1.848 | 1.3894 | 2.1138 | 3.2307
Q9UHD8 | Septin-9 | SEPTIN9 | 1:1 | Block_4 | 0.46662 | 7.0473 | FALSE | NA | NA | NA | 0.66473 | 2.1708 | 0.46999 | 2.2049 | 3.6636
Q86WD7 | Serpin A9 | SERPINA9 | 1:1 | Block_4 | -2.7378 | 8.6417 | TRUE | 7.4916 | 4.5151 | -1.7126 | -0.47221 | 0.54521 | -1.4924 | -0.27182 | 1.2655
P48594 | Serpin B4 | SERPINB4 | 1:1 | Block_4 | -2.5 | 4.0055 | TRUE | 12.3658 | 11.7442 | -0.69534 | -0.13572 | 0.71909 | -0.38421 | 0.50556 | 1.7549
P50452 | Serpin B8 | SERPINB8 | 1:1 | Block_4 | -2 | 6.5602 | TRUE | 13.4163 | 15.1332 | -1.5793 | -0.81078 | 0.07161 | -0.90003 | 0.092143 | 2.7663
P50453 | Serpin B9 | SERPINB9 | 1:1 | Block_4 | -1.697 | 4.8974 | FALSE | 10.6547 | 9.8018 | -0.70551 | 0.80682 | 1.812 | 0.51951 | 1.8417 | 2.4757
O75830 | Serpin I2 | SERPINI2 | 1:1 | Block_4 | -2.3691 | 6.5007 | TRUE | 8.5415 | 4.6589 | -1.7494 | -0.46277 | 1.3017 | -1.5014 | -0.24424 | 1.0188
Q53H47 | Histone-lysine N-methyltransferase SETMAR | SETMAR | 1:1 | Block_4 | -3.1085 | 2.877 | FALSE | 7.2661 | 3.3124 | -0.74854 | 0.14537 | 0.59104 | 0.19889 | 0.81627 | 1.7816
Q53EL9 | Seizure protein 6 homolog | SEZ6 | 1:1 | Block_4 | -1.4888 | 7.2852 | TRUE | 15.0443 | 9.6504 | -0.3224 | 0.15526 | 0.76395 | -0.14955 | 0.47135 | 1.3666
Q6UXD5 | Seizure 6-like protein 2 | SEZ6L2 | 1:1 | Block_4 | -2.9423 | 7.2652 | TRUE | 9.4403 | 12.1748 | -0.53992 | 0.15858 | 0.56004 | -0.061393 | 0.55719 | 1.2867
Q15427 | Splicing factor 3B subunit 4 | SF3B4 | 1:1 | Block_4 | 0.12279 | 6.5771 | TRUE | NA | 1.2284 | NA | 0.60312 | 2.4551 | 0.56025 | 2.7025 | 4.0876
Q8N474 | Secreted frizzled-related protein 1 | SFRP1 | 1:1 | Block_4 | -4.1696 | 7.643 | TRUE | 11.2518 | 13.765 | 0.40655 | 1.1927 | 3.3588 | 0.82655 | 2.0008 | 3.4075
O60880 | SH2 domain-containing protein 1A | SH2D1A | 1:1 | Block_4 | -0.15694 | 9.3417 | TRUE | 8.8425 | 13.2621 | NA | 0.28173 | 1.2757 | 0.045376 | 1.5061 | 3.4201
Q9Y3L3 | SH3 domain-binding protein 1 | SH3BP1 | 1:1 | Block_4 | 1.3754 | 6.274 | FALSE | NA | NA | NA | NA | 1.841 | NA | 1.631 | 3.5606
Q9NR46 | Endophilin-B2 | SH3GLB2 | 1:1 | Block_4 | -0.39665 | 5.374 | FALSE | NA | 0.53123 | 0.02159 | 1.9543 | 3.2145 | 1.2465 | 3.0439 | 4.0243
A1X283 | SH3 and PX domain-containing protein 2B | SH3PXD2B | 1:1 | Block_4 | -2.4609 | 9.3103 | TRUE | 12.4022 | 3.8474 | -0.6112 | 0.13926 | 2.0038 | 0.64345 | 2.3802 | 4.6571
Q15465 | Sonic hedgehog protein | SHH | 1:1 | Block_4 | 2.5981 | 5.3639 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96LC7 | Sialic acid-binding Ig-like lectin 10 | SIGLEC10 | 1:1 | Block_4 | -2.6336 | 6.636 | FALSE | 7.6295 | 9.067 | -0.58611 | 0.13784 | 0.79635 | 0.054224 | 0.76853 | 1.7306
O43699 | Sialic acid-binding Ig-like lectin 6 | SIGLEC6 | 1:1 | Block_4 | -4.5501 | 7.4032 | TRUE | 25.6249 | 14.9646 | -1.1799 | 0.3446 | 0.74392 | -0.27445 | 0.66843 | 1.5841
Q9NYZ4 | Sialic acid-binding Ig-like lectin 8 | SIGLEC8 | 1:1 | Block_4 | -5 | 5.5709 | TRUE | 15.3329 | 8.4759 | -0.56797 | 0.061072 | 0.88896 | 0.067606 | 1.1424 | 2.6032
Q9Y336 | Sialic acid-binding Ig-like lectin 9 | SIGLEC9 | 1:1 | Block_4 | -3.005 | 3.8743 | TRUE | 11.0156 | 6.9127 | -0.89432 | -0.0015526 | 0.5241 | -0.39043 | 0.21876 | 0.72965
Q9H173 | Nucleotide exchange factor SIL1 | SIL1 | 1:1 | Block_4 | -0.73112 | 7.7836 | FALSE | 8.9909 | 7.0528 | -0.5814 | -0.1781 | 0.40209 | -0.37046 | 0.14636 | 0.81011
Q8IXJ6 | NAD-dependent protein deacetylase sirtuin-2 | SIRT2 | 1:1 | Block_4 | -0.4744 | 7.9735 | TRUE | 21.6163 | 10.5624 | -0.011206 | 1.9132 | 4.9374 | 2.5014 | 5.1285 | 6.5724
Q86WV1 | Src kinase-associated phosphoprotein 1 | SKAP1 | 1:1 | Block_4 | -0.58585 | 7.4919 | TRUE | 9.6792 | 6.7699 | NA | 0.32072 | 1.2775 | NA | 0.36789 | 2.3512
Q9H6Q3 | Src-like-adapter 2 | SLA2 | 1:1 | Block_4 | 0.69392 | 7.8973 | FALSE | NA | NA | NA | 0.77616 | 3.6638 | NA | 3.0857 | 5.3257
Q9NQ25 | SLAM family member 7 | SLAMF7 | 1:1 | Block_4 | -3.7441 | 6.9617 | FALSE | 5.1507 | 3.8892 | -1.7337 | -0.31831 | 0.21991 | -1.2089 | -0.19421 | 1.2786
Q9BZW2 | Solute carrier family 13 member 1 | SLC13A1 | 1:1 | Block_4 | -0.19143 | 10.8707 | FALSE | NA | NA | NA | NA | 0.15632 | NA | -0.048404 | 0.78015
Q6ZMH5 | Zinc transporter ZIP5 | SLC39A5 | 1:1 | Block_4 | -2.6336 | 4.8579 | TRUE | 8.4982 | 4.3885 | -0.6071 | -0.10497 | 1.1079 | -0.25751 | 0.62497 | 2.0335
Q96PX8 | SLIT and NTRK-like protein 1 | SLITRK1 | 1:1 | Block_4 | -2.2538 | 7.3131 | TRUE | 10.9686 | 7.5151 | -0.28243 | 0.27961 | 0.74083 | -0.27092 | 0.52404 | 1.342
Q9H156 | SLIT and NTRK-like protein 2 | SLITRK2 | 1:1 | Block_4 | -2.4185 | 6.9972 | FALSE | 7.1216 | 3.568 | -0.53182 | 0.53965 | 1.1005 | 0.3149 | 1.0821 | 1.9587
Q9H5Y7 | SLIT and NTRK-like protein 6 | SLITRK6 | 1:1 | Block_4 | -1.9326 | 9.1495 | TRUE | 11.8826 | 2.6341 | -0.36256 | 0.45523 | 1.1769 | 0.0089381 | 0.68141 | 1.4522
Q14BN4 | Sarcolemmal membrane-associated protein | SLMAP | 1:1 | Block_4 | -0.77319 | 8.7009 | TRUE | 11.7476 | 11.1698 | -0.1851 | 1.2143 | 2.8767 | 1.4452 | 2.7308 | 5.1446
P84022 | Mothers against decapentaplegic homolog 3 | SMAD3 | 1:1 | Block_4 | 1 | 8.3546 | FALSE | NA | 23.9138 | NA | NA | 3.4467 | NA | 3.3106 | 5.2302
Q99717 | Mothers against decapentaplegic homolog 5 | SMAD5 | 1:1 | Block_4 | -2.1613 | 5.8084 | TRUE | 12.1904 | 10.2828 | 0.033982 | 0.55864 | 0.99386 | 0.055564 | 0.76073 | 1.1029
O75940 | Survival of motor neuron-related-splicing factor 30 | SMNDC1 | 1:1 | Block_4 | -1.1719 | 4.9804 | TRUE | 10.1283 | 12.0739 | NA | -1.1347 | -0.22352 | NA | -0.14012 | 4.0156
Q9H4F8 | SPARC-related modular calcium-binding protein 1 | SMOC1 | 1:1 | Block_4 | -3.0572 | 3.3702 | TRUE | 7.5651 | 4.606 | -0.33609 | 0.50215 | 1.2475 | 0.47572 | 1.2799 | 2.1391
Q92485 | Acid sphingomyelinase-like phosphodiesterase 3b | SMPDL3B | 1:1 | Block_4 | 0.49637 | 5.6959 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P53814 | Smoothelin | SMTN | 1:1 | Block_4 | -0.10984 | 5.9096 | FALSE | 10.9518 | 11.0649 | NA | 0.65791 | 2.6967 | 0.97983 | 2.5678 | 5.7193
O95295 | SNARE-associated protein Snapin | SNAPIN | 1:1 | Block_4 | 0.19644 | 5.609 | TRUE | 2.4705 | 3.5527 | NA | NA | 0.39265 | NA | 0.3232 | 0.85118
O76070 | Gamma-synuclein | SNCG | 1:1 | Block_4 | -3.4162 | 6.6861 | TRUE | 10.9906 | 8.2321 | -1.1547 | 0.18709 | 0.73964 | -0.73872 | 1.0637 | 2.7814
Q8TER0 | Sushi, nidogen and EGF-like domain-containing protein 1 | SNED1 | 1:1 | Block_4 | -3 | 5.0313 | FALSE | 11.5283 | 14.5506 | 0.34476 | 0.93887 | 1.5097 | 0.48697 | 1.0558 | 1.707
P08579 | U2 small nuclear ribonucleoprotein B'' | SNRPB2 | 1:1 | Block_4 | -0.92603 | 7.3666 | FALSE | 8.164 | 10.6758 | NA | NA | 0.17963 | -0.57477 | 0.62335 | 3.6409
Q96RF0 | Sorting nexin-18 | SNX18 | 1:1 | Block_4 | 0.5 | 5.1923 | FALSE | 14.6741 | 19.8201 | NA | 1.035 | 3.1951 | NA | 2.8051 | 4.3606
O60749 | Sorting nexin-2 | SNX2 | 1:1 | Block_4 | 0.34688 | 7.2955 | TRUE | NA | NA | NA | 1.629 | 4.1828 | 0.99007 | 3.8681 | 5.3443
Q8TEQ0 | Sorting nexin-29 | SNX29 | 1:1 | Block_4 | -1.8644 | 9.96 | TRUE | 13.2941 | 14.9059 | -1.3811 | 0.61813 | 5.0298 | 0.69946 | 3.6611 | 6.9356
Q8WV41 | Sorting nexin-33 | SNX33 | 1:1 | Block_4 | -3 | 7.658 | TRUE | 15.1864 | 17.6815 | -0.44692 | 0.13845 | 0.66432 | 0.31243 | 0.9509 | 1.9812
Q9Y5X3 | Sorting nexin-5 | SNX5 | 1:1 | Block_4 | -7 | 2.4952 | FALSE | 19.3634 | 16.4784 | -0.192 | 0.39801 | 1.2543 | 0.1446 | 1.5714 | 2.5719
Q00796 | Sorbitol dehydrogenase | SORD | 1:1 | Block_4 | -4.7271 | NA | NA | 8.4145 | 7.1865 | -2.222 | -1.0869 | 0.022826 | -1.8178 | 0.74732 | 2.4591
Q53LP3 | Ankyrin repeat domain-containing protein SOWAHC | SOWAHC | 1:1 | Block_4 | -1.5089 | 8.6098 | TRUE | 10.7558 | 5.2012 | -0.57967 | 0.045873 | 0.71059 | 0.032438 | 1.1935 | 2.499
O60575 | Serine protease inhibitor Kazal-type 4 | SPINK4 | 1:1 | Block_4 | -2.0298 | 3.7711 | TRUE | 11.0838 | 7.1588 | -1.2958 | -0.20441 | 0.46584 | -0.81353 | 0.54955 | 2.0647
Q6UWN8 | Serine protease inhibitor Kazal-type 6 | SPINK6 | 1:1 | Block_4 | -4.7733 | 4.766 | TRUE | 5.9716 | 3.7775 | -0.97736 | 0.041683 | 0.68036 | -0.46141 | 0.59468 | 1.6851
P49223 | Kunitz-type protease inhibitor 3 | SPINT3 | 1:1 | Block_4 | -3.5 | 6.5047 | TRUE | 7.4775 | 5.3758 | NA | -1.9031 | 1.0562 | NA | -0.97487 | 1.3417
Q08629 | Testican-1 | SPOCK1 | 1:1 | Block_4 | -3.5156 | 8.1886 | FALSE | 9.5233 | 6.0717 | -0.38947 | 0.36469 | 0.97387 | 0.0054813 | 0.65736 | 1.3288
Q9UBC9 | Small proline-rich protein 3 | SPRR3 | 1:1 | Block_4 | -3 | 6.1564 | TRUE | 9.3051 | 9.0524 | -0.56317 | 0.31203 | 1.2579 | -0.72364 | 0.41664 | 1.8369
O15269 | Serine palmitoyltransferase 1 | SPTLC1 | 1:1 | Block_4 | -0.012066 | 8.3394 | FALSE | 11.8141 | 11.7673 | NA | 0.48978 | 1.5142 | 0.33143 | 1.0805 | 2.0595
P37108 | Signal recognition particle 14 kDa protein | SRP14 | 1:1 | Block_4 | -0.27674 | 6.0311 | TRUE | NA | NA | NA | NA | 0.88681 | -0.055145 | 0.91063 | 3.0448
P78362 | SRSF protein kinase 2 | SRPK2 | 1:1 | Block_4 | 0.69911 | 8.1405 | TRUE | NA | NA | NA | NA | 1.077 | NA | 1.4151 | 3.3704
O43805 | Microtubule nucleation factor SSNA1 | SSNA1 | 1:1 | Block_4 | 0.73468 | 9.3001 | FALSE | NA | NA | NA | NA | 2.1757 | 0.93915 | 2.7752 | 4.3153
Q11201 | CMP-N-acetylneuraminate-beta-galactosamide-alpha-2,3-sialyltransferase 1 | ST3GAL1 | 1:1 | Block_4 | -2.0811 | 3.8871 | FALSE | 11.055 | 10.9315 | 0.0055599 | 0.45296 | 1.0626 | -0.054806 | 0.60262 | 1.3134
Q92783 | Signal transducing adapter molecule 1 | STAM | 1:1 | Block_4 | 1.1036 | 5.3908 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.2467
O95630 | STAM-binding protein | STAMBP | 1:1 | Block_4 | -3.53 | 6.3067 | TRUE | 13.7003 | 12.5509 | 0.15885 | 1.7868 | 4.0444 | 1.7149 | 4.0697 | 5.4751
P52630 | Signal transducer and activator of transcription 2 | STAT2 | 1:1 | Block_4 | -0.83372 | 6.3958 | FALSE | 15.128 | 13.8632 | -0.77343 | 0.19182 | 2.4178 | 0.94285 | 2.5498 | 4.089
P51692 | Signal transducer and activator of transcription 5B | STAT5B | 1:1 | Block_4 | -1.5724 | 9.8736 | TRUE | 25.2025 | 9.6268 | -0.60914 | 1.6059 | 6.8948 | 2.1963 | 6.9053 | 8.9615
Q15831 | Serine/threonine-protein kinase STK11 | STK11 | 1:1 | Block_4 | 0.001 | 10.1555 | FALSE | 15.2616 | 6.6073 | 0.014734 | 1.3363 | 3.5742 | 1.7092 | 3.6124 | 5.135
Q13277 | Syntaxin-3 | STX3 | 1:1 | Block_4 | 1.2619 | 10.0184 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.4417
Q12846 | Syntaxin-4 | STX4 | 1:1 | Block_4 | -1.6917 | 7.4467 | FALSE | 7.287 | 5.3808 | -0.35444 | 0.70385 | 2.4478 | 1.3039 | 2.2672 | 4.5252
O15400 | Syntaxin-7 | STX7 | 1:1 | Block_4 | -1.5 | 6.5997 | TRUE | 11.3563 | 7.7386 | 0.25289 | 1.383 | 2.1215 | 0.99781 | 2.1134 | 3.625
Q9UNK0 | Syntaxin-8 | STX8 | 1:1 | Block_4 | -0.75193 | 9.536 | TRUE | 10.6525 | 14.0134 | -0.1852 | 0.84883 | 2.2824 | 1.3289 | 2.76 | 4.491
P61764 | Syntaxin-binding protein 1 | STXBP1 | 1:1 | Block_4 | 1.7221 | 7.1216 | FALSE | NA | NA | NA | 1.7803 | 4.0734 | NA | 3.5711 | 4.857
Q8IWZ8 | SURP and G-patch domain-containing protein 1 | SUGP1 | 1:1 | Block_4 | -0.11184 | 8.2843 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.5499
Q06520 | Sulfotransferase 2A1 | SULT2A1 | 1:1 | Block_4 | -2.9758 | 3.5144 | FALSE | 4.8661 | 9.1993 | -2.7455 | -2.387 | -0.38649 | -2.7036 | -1.6017 | 0.8368
P51687 | Sulfite oxidase, mitochondrial | SUOX | 1:1 | Block_4 | -0.67743 | 6.3214 | TRUE | 12.2822 | 10.3194 | NA | -0.37935 | 0.58834 | -0.013113 | 0.89663 | 2.5687
Q5VX71 | Sushi domain-containing protein 4 | SUSD4 | 1:1 | Block_4 | -0.6272 | 10.7224 | TRUE | 10.0347 | 4.1237 | -0.62472 | -0.0092594 | 0.62954 | -0.25424 | 0.33294 | 1.2889
Q7L0J3 | Synaptic vesicle glycoprotein 2A | SV2A | 1:1 | Block_4 | 1.1944 | 13.764 | TRUE | 12.2513 | 19.3568 | NA | NA | 3.4041 | NA | 2.0613 | 4.195
Q9UH65 | Switch-associated protein 70 | SWAP70 | 1:1 | Block_4 | -0.5 | 9.8404 | TRUE | 21.7855 | 7.1671 | NA | 0.31464 | 1.163 | 0.62647 | 1.3365 | 3.2264
P21579 | Synaptotagmin-1 | SYT1 | 1:1 | Block_4 | -0.26056 | 8.0314 | TRUE | 8.4566 | 6.623 | NA | 0.095549 | 0.54262 | -0.095145 | 0.57041 | 1.5191
Q9Y6A5 | Transforming acidic coiled-coil-containing protein 3 | TACC3 | 1:1 | Block_4 | 1 | 8.2606 | TRUE | NA | NA | NA | 1.1196 | 4.8192 | 1.5075 | 5.1588 | 6.935
P09758 | Tumor-associated calcium signal transducer 2 | TACSTD2 | 1:1 | Block_4 | -1.584 | 6.6016 | TRUE | 9.1619 | 4.2805 | -0.46618 | 0.21272 | 0.72977 | 0.022029 | 0.5836 | 1.3783
O75528 | Transcriptional adapter 3 | TADA3 | 1:1 | Block_4 | 0.22429 | 7.7942 | FALSE | NA | 0.58838 | NA | 0.7311 | 3.6961 | 0.93828 | 3.5011 | 5.0249
Q7Z5A7 | Chemokine-like protein TAFA-5 | TAFA5 | 1:1 | Block_4 | -1.5906 | 7.1968 | TRUE | 8.6765 | 6.4014 | -0.28398 | 0.68056 | 1.3946 | 0.44421 | 1.3529 | 2.7748
Q86VP1 | Tax1-binding protein 1 | TAX1BP1 | 1:1 | Block_4 | -1.4313 | 11.32 | TRUE | 9.5732 | 6.6791 | -0.84681 | 0.33369 | 1.6548 | 0.25935 | 1.298 | 3.6582
Q9NUY8 | TBC1 domain family member 23 | TBC1D23 | 1:1 | Block_4 | 0.55883 | 6.1001 | TRUE | NA | NA | NA | 0.59324 | 2.4907 | NA | 1.8717 | 3.9097
Q99426 | Tubulin-folding cofactor B | TBCB | 1:1 | Block_4 | 1.0446 | 9.9115 | TRUE | NA | NA | NA | 2.1415 | 6.0737 | 2.643 | 6.5392 | 7.6876
Q15814 | Tubulin-specific chaperone C | TBCC | 1:1 | Block_4 | -1.2762 | 6.9019 | FALSE | 10.9555 | 12.9117 | NA | -0.080819 | 2.6359 | 0.16623 | 3.0705 | 4.5307
O14776 | Transcription elongation regulator 1 | TCERG1 | 1:1 | Block_4 | -3.5709 | 8.1736 | TRUE | 7.4704 | 5.6622 | -2.7237 | -1.8096 | -0.95853 | -1.8762 | -0.79683 | 1.6223
P56279 | T-cell leukemia/lymphoma protein 1A | TCL1A | 1:1 | Block_4 | -2.2489 | 6.2608 | TRUE | 5.2482 | 4.4221 | -1.3611 | -0.27017 | 1.3623 | -1.8698 | 0.61772 | 2.8385
Q9NUW8 | Tyrosyl-DNA phosphodiesterase 1 | TDP1 | 1:1 | Block_4 | 0.42724 | 7.93 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.0965
Q9H7E2 | Tudor domain-containing protein 3 | TDRD3 | 1:1 | Block_4 | 0.86065 | 9.1908 | TRUE | NA | NA | NA | NA | 1.3541 | NA | 1.2746 | 3.0495
Q9BY14 | Testis-expressed protein 101 | TEX101 | 1:1 | Block_4 | 0.036456 | 11.1797 | TRUE | 8.6662 | 7.882 | NA | NA | 1.2523 | NA | 0.10845 | 2.2351
P01266 | Thyroglobulin | TG | 1:1 | Block_4 | -0.52213 | 8.7981 | TRUE | 6.3608 | 9.3541 | NA | -0.047201 | 1.8015 | NA | 0.43398 | 2.0907
P01135 | Protransforming growth factor alpha | TGFA | 1:1 | Block_4 | -2.0445 | 8.9564 | TRUE | 8.793 | 9.6 | -0.16582 | 0.21852 | 0.87053 | 0.1495 | 0.94653 | 3.0639
P01137 | Transforming growth factor beta-1 proprotein | TGFB1 | 1:1 | Block_4 | -1.8482 | 5.9453 | TRUE | 12.8789 | 6.3217 | -0.36707 | 0.65279 | 1.6073 | 1.0093 | 1.6262 | 2.4833
P61812 | Transforming growth factor beta-2 proprotein | TGFB2 | 1:1 | Block_4 | 0.46853 | 8.6454 | FALSE | NA | NA | NA | NA | 0.55597 | NA | NA | 1.233
P36897 | TGF-beta receptor type-1 | TGFBR1 | 1:1 | Block_4 | -1 | 8.3443 | TRUE | 18.4528 | 14.3998 | -0.97207 | -0.30051 | 0.2237 | -0.78408 | 0.093043 | 1.5443
P40225 | Thrombopoietin | THPO | 1:1 | Block_4 | -2.459 | 4.7171 | FALSE | 7.2302 | 5.2662 | -0.047026 | 0.61655 | 1.2871 | 0.50623 | 1.3279 | 2.1197
Q9BU02 | Thiamine-triphosphatase | THTPA | 1:1 | Block_4 | 0.10463 | 9.3979 | TRUE | 11.7468 | 7.5012 | 0.32624 | 2.478 | 4.0915 | 1.9849 | 3.8621 | 5.577
Q9NQ88 | Fructose-2,6-bisphosphatase TIGAR | TIGAR | 1:1 | Block_4 | -1.7246 | 4.8798 | FALSE | 17.0973 | 22.4163 | -1.4207 | -0.67273 | -0.17346 | -0.71796 | 0.36048 | 2.9058
O60220 | Mitochondrial import inner membrane translocase subunit Tim8 A | TIMM8A | 1:1 | Block_4 | 1.2025 | 7.7804 | TRUE | NA | NA | NA | NA | 1.7407 | NA | 1.6884 | 2.8366
Q15399 | Toll-like receptor 1 | TLR1 | 1:1 | Block_4 | -0.29421 | 10.2854 | TRUE | 14.9351 | 9.1169 | NA | 0.12902 | 0.67811 | NA | 0.29658 | 1.153
O15455 | Toll-like receptor 3 | TLR3 | 1:1 | Block_4 | -6 | 6.971 | TRUE | 5.1837 | 6.3398 | -1.0272 | 0.13831 | 0.9108 | -1.2324 | 0.43142 | 1.217
Q6PL24 | Protein TMED8 | TMED8 | 1:1 | Block_4 | 0.40383 | 10.1267 | TRUE | NA | NA | NA | 1.0226 | 4.963 | 1.1322 | 4.9539 | 6.6667
Q96A25 | Transmembrane protein 106A | TMEM106A | 1:1 | Block_4 | -0.35198 | 8.8636 | TRUE | 6.7551 | 11.3515 | -0.16661 | 0.97216 | 2.8591 | 0.76234 | 3.031 | 4.5596
Q24JP5 | Transmembrane protein 132A | TMEM132A | 1:1 | Block_4 | -0.74071 | NA | NA | 8.4454 | 7.441 | -0.55276 | 0.15372 | 0.78579 | NA | 0.28526 | 1.4233
Q86YD3 | Transmembrane protein 25 | TMEM25 | 1:1 | Block_4 | -3 | 6.76 | TRUE | 17.3619 | 12.0923 | -0.1881 | 0.32751 | 0.95907 | -0.4326 | 0.35747 | 1.3777
P82094 | TATA element modulatory factor | TMF1 | 1:1 | Block_4 | -1.657 | 10.2348 | TRUE | 8.3215 | 8.8385 | -1.3839 | 0.35189 | 2.2441 | -0.28599 | 1.7332 | 4.8498
O60235 | Transmembrane protease serine 11D | TMPRSS11D | 1:1 | Block_4 | -2.5 | 8.3348 | TRUE | 6.2597 | 3.6253 | -0.59597 | -0.1285 | 0.55127 | -0.4735 | 0.22984 | 1.5511
P98073 | Enteropeptidase | TMPRSS15 | 1:1 | Block_4 | -0.83065 | 6.7743 | TRUE | 12.7923 | 9.1147 | -0.46489 | 0.67945 | 3.0498 | -0.64061 | 0.93146 | 2.6638
Q9H3S3 | Transmembrane protease serine 5 | TMPRSS5 | 1:1 | Block_4 | -4.1932 | 7.5092 | TRUE | 10.3394 | 7.3294 | -0.72846 | -0.096535 | 0.60745 | -0.30282 | 0.33458 | 1.041
Q03169 | Tumor necrosis factor alpha-induced protein 2 | TNFAIP2 | 1:1 | Block_4 | -0.62944 | 6.6738 | TRUE | 17.9867 | 7.892 | -0.4316 | 0.20903 | 0.72775 | 0.2401 | 1.0287 | 2.3902
P98066 | Tumor necrosis factor-inducible gene 6 protein | TNFAIP6 | 1:1 | Block_4 | -3.444 | 1.3955 | FALSE | 9.3572 | 11.3017 | -0.80756 | 0.1108 | 0.67647 | -0.26789 | 0.81349 | 2.0826
O00220 | Tumor necrosis factor receptor superfamily member 10A | TNFRSF10A | 1:1 | Block_4 | -0.77053 | 8.1311 | TRUE | 8.2756 | 6.751 | -0.64221 | -0.086551 | 0.45659 | -0.24698 | 0.50009 | 1.5213
O14763 | Tumor necrosis factor receptor superfamily member 10B | TNFRSF10B | 1:1 | Block_4 | -4.6889 | 8.4989 | TRUE | 5.6488 | 3.3892 | -0.39219 | 0.17679 | 0.47218 | -0.050042 | 1.1075 | 2.3565
Q9Y6Q6 | Tumor necrosis factor receptor superfamily member 11A | TNFRSF11A | 1:1 | Block_4 | -3.3483 | 5.4905 | TRUE | 8.319 | 11.9733 | -0.77937 | 0.091336 | 0.63832 | 0.25213 | 1.2319 | 2.7846
Q9NS68 | Tumor necrosis factor receptor superfamily member 19 | TNFRSF19 | 1:1 | Block_4 | -3.5969 | 4.5642 | TRUE | 8.0864 | 10.2613 | -0.48973 | 0.11639 | 1.0136 | 0.27365 | 1.037 | 2.3929
P43489 | Tumor necrosis factor receptor superfamily member 4 | TNFRSF4 | 1:1 | Block_4 | -5.2468 | 6.8283 | TRUE | 7.3394 | 5.0701 | -0.54365 | 0.15068 | 0.7384 | 0.1497 | 1.0747 | 2.5192
O95407 | Tumor necrosis factor receptor superfamily member 6B | TNFRSF6B | 1:1 | Block_4 | -3.6009 | 6.7107 | TRUE | 8.2704 | 6.4202 | -0.9122 | -0.090201 | 0.52342 | -0.19949 | 0.75721 | 2.5083
Q07011 | Tumor necrosis factor receptor superfamily member 9 | TNFRSF9 | 1:1 | Block_4 | -3.5 | 7.4551 | TRUE | 15.2689 | 13.0217 | -0.7273 | 0.097935 | 0.90068 | -0.015399 | 1.0066 | 2.4614
P50591 | Tumor necrosis factor ligand superfamily member 10 | TNFSF10 | 1:1 | Block_4 | -6.1517 | 4.9754 | TRUE | 3.6126 | 4.0459 | -0.23846 | 0.444 | 0.99691 | -0.12422 | 0.54244 | 1.0755
O14788 | Tumor necrosis factor ligand superfamily member 11 | TNFSF11 | 1:1 | Block_4 | -2.0128 | 9.4513 | TRUE | 10.8126 | 7.3599 | -0.56637 | 0.44087 | 1.2781 | -0.76046 | 0.51892 | 2.4559
O43508 | Tumor necrosis factor ligand superfamily member 12 | TNFSF12 | 1:1 | Block_4 | -4.4593 | 7.2118 | TRUE | 8.3188 | 7.1809 | -0.08515 | 0.80695 | 1.3728 | 0.44813 | 1.2527 | 1.9962
O43557 | Tumor necrosis factor ligand superfamily member 14 | TNFSF14 | 1:1 | Block_4 | -3.7921 | 6.8325 | TRUE | 7.5216 | 3.4829 | -0.15179 | 0.78285 | 1.856 | 1.0244 | 2.0319 | 4.2156
Q15025 | TNFAIP3-interacting protein 1 | TNIP1 | 1:1 | Block_4 | -0.83999 | 8.531 | FALSE | NA | 12.841 | NA | -0.0246 | 2.1457 | -0.25514 | 1.9538 | 4.5737
Q92752 | Tenascin-R | TNR | 1:1 | Block_4 | -2.5 | 5.461 | TRUE | 11.2104 | 6.2819 | -0.10133 | 0.42304 | 1.1136 | 0.042754 | 0.72072 | 1.4187
Q6ZVM7 | TOM1-like protein 2 | TOM1L2 | 1:1 | Block_4 | 2.222 | NA | NA | NA | NA | NA | NA | 5.0312 | NA | 4.6511 | 6.0684
P11387 | DNA topoisomerase 1 | TOP1 | 1:1 | Block_4 | -1.7492 | 0.3719 | FALSE | 11.8926 | 10.0863 | NA | -0.44508 | 0.72297 | NA | -0.90912 | 0.80713
Q02880 | DNA topoisomerase 2-beta | TOP2B | 1:1 | Block_4 | 0.22884 | 8.2501 | TRUE | NA | NA | NA | NA | 1.0279 | 0.25919 | 1.5903 | 3.6556
Q5JTV8 | Torsin-1A-interacting protein 1 | TOR1AIP1 | 1:1 | Block_4 | -1.3633 | 6.3173 | FALSE | 9.6813 | 5.5101 | -0.21501 | 1.5436 | 3.0626 | 0.69097 | 3.2186 | 6.6177
Q96A56 | Tumor protein p53-inducible nuclear protein 1 | TP53INP1 | 1:1 | Block_4 | 0.71645 | 7.0086 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q15661 | Tryptase alpha/beta-1 | TPSAB1 | 1:1 | Block_4 | -2.5 | 3.1812 | FALSE | 9.9445 | 5.3256 | -0.63093 | -0.064434 | 0.60891 | -0.56068 | 0.18468 | 1.2144
Q12933 | TNF receptor-associated factor 2 | TRAF2 | 1:1 | Block_4 | -0.34715 | 1.2424 | FALSE | 10.9401 | 7.301 | NA | 0.34792 | 2.3371 | 0.63327 | 1.6434 | 3.4763
O43715 | TP53-regulated inhibitor of apoptosis 1 | TRIAP1 | 1:1 | Block_4 | -0.54078 | 6.4219 | TRUE | 8.6074 | 5.3711 | -0.38876 | 0.2714 | 1.4766 | 1.0262 | 2.0182 | 3.2131
P19474 | E3 ubiquitin-protein ligase TRIM21 | TRIM21 | 1:1 | Block_4 | -0.79193 | 8.112 | FALSE | 13.8219 | 6.4536 | 0.62327 | 2.08 | 3.304 | 2.0229 | 3.5018 | 4.875
Q14258 | E3 ubiquitin/ISG15 ligase TRIM25 | TRIM25 | 1:1 | Block_4 | -1.7096 | 9.6225 | TRUE | 17.0585 | 9.8318 | -1.4185 | 1.2322 | 3.3342 | 0.41771 | 3.4091 | 5.694
Q9C035 | Tripartite motif-containing protein 5 | TRIM5 | 1:1 | Block_4 | 1.6758 | 7.7248 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.4198
Q8NG06 | E3 ubiquitin-protein ligase TRIM58 | TRIM58 | 1:1 | Block_4 | -0.32606 | 7.074 | TRUE | 11.9306 | 8.3334 | NA | 1.6775 | 4.1001 | 1.391 | 4.1128 | 6.0528
Q15643 | Thyroid receptor-interacting protein 11 | TRIP11 | 1:1 | Block_4 | -1.4329 | 10.556 | TRUE | 14.0699 | 15.7846 | NA | 0.69157 | 3.3899 | 0.47763 | 3.0093 | 6.091
Q15714 | TSC22 domain family protein 1 | TSC22D1 | 1:1 | Block_4 | 0.72199 | 8.8603 | TRUE | NA | NA | NA | NA | 1.9009 | 0.73865 | 1.937 | 3.2635
O60635 | Tetraspanin-1 | TSPAN1 | 1:1 | Block_4 | 0.27273 | 4.4371 | TRUE | 4.1921 | 3.3861 | NA | NA | 0.58571 | NA | NA | 1.3053
P19075 | Tetraspanin-8 | TSPAN8 | 1:1 | Block_4 | -3.1965 | 7.7082 | TRUE | 12.401 | 7.5572 | -2.6076 | -0.38626 | 0.9273 | -1.3598 | 0.29008 | 2.1376
Q9H0U9 | Testis-specific Y-encoded-like protein 1 | TSPYL1 | 1:1 | Block_4 | 0.32419 | 8.3067 | FALSE | 8.092 | 8.219 | NA | 0.37312 | 1.3462 | 0.53627 | 1.4626 | 2.6179
P40222 | Alpha-taxilin | TXLNA | 1:1 | Block_4 | -0.70494 | 6.4175 | TRUE | 16.6195 | 11.4008 | -0.6912 | 0.2771 | 2.3273 | 1.0947 | 2.7506 | 4.3383
O14530 | Thioredoxin domain-containing protein 9 | TXNDC9 | 1:1 | Block_4 | 1.5014 | 8.1645 | TRUE | NA | NA | NA | NA | 2.1947 | NA | 2.3161 | 4.5464
P17643 | 5,6-dihydroxyindole-2-carboxylic acid oxidase | TYRP1 | 1:1 | Block_4 | 0.14872 | 3.6777 | FALSE | 5.0055 | 6.7828 | NA | 0.26539 | 0.88697 | NA | 0.45877 | 1.5569
Q9BSL1 | Ubiquitin-associated domain-containing protein 1 | UBAC1 | 1:1 | Block_4 | -2.7512 | 4.6455 | TRUE | 7.3878 | 7.8283 | 0.73296 | 1.5687 | 2.2185 | 1.5215 | 2.0902 | 2.725
P63146 | Ubiquitin-conjugating enzyme E2 B | UBE2B | 1:1 | Block_4 | 1.1744 | 7.65 | TRUE | NA | NA | NA | 1.222 | 2.5989 | NA | 2.3817 | 3.887
P11441 | Ubiquitin-like protein 4A | UBL4A | 1:1 | Block_4 | -1.7814 | 10.5169 | TRUE | 10.4516 | 12.4241 | 0.16529 | 1.2108 | 2.4731 | 1.5803 | 2.5731 | 3.6605
Q14CS0 | UBX domain-containing protein 2B | UBXN2B | 1:1 | Block_4 | -2.1109 | NA | NA | 16.321 | 12.8229 | -0.52089 | 0.61072 | 2.4668 | 0.83674 | 2.4372 | 4.49
Q92890 | Ubiquitin recognition factor in ER-associated degradation protein 1 | UFD1 | 1:1 | Block_4 | 1.6487 | 12.2151 | TRUE | NA | NA | NA | NA | 6.5168 | 1.8427 | 6.0749 | 8.3636
Q96PU4 | E3 ubiquitin-protein ligase UHRF2 | UHRF2 | 1:1 | Block_4 | 0.8773 | 11.9475 | TRUE | NA | NA | NA | NA | 1.0258 | NA | NA | 1.2677
P13051 | Uracil-DNA glycosylase | UNG | 1:1 | Block_4 | 1.874 | 7.562 | FALSE | NA | NA | NA | NA | 2.6005 | NA | 1.9477 | 3.2928
O75631 | Uroplakin-3a | UPK3A | 1:1 | Block_4 | 0.32472 | 6.8056 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P10746 | Uroporphyrinogen-III synthase | UROS | 1:1 | Block_4 | -0.14671 | 10.0034 | TRUE | 10.6005 | 8.1507 | NA | 0.61492 | 1.8273 | 0.28837 | 1.1294 | 2.4457
O60763 | General vesicular transport factor p115 | USO1 | 1:1 | Block_4 | -0.83346 | 4.0591 | FALSE | 18.3498 | 19.0291 | NA | 1.2781 | 3.4721 | 1.5405 | 4.0884 | 5.0487
Q9UHP3 | Ubiquitin carboxyl-terminal hydrolase 25 | USP25 | 1:1 | Block_4 | -0.21239 | 9.562 | TRUE | 7.2935 | 10.5337 | NA | 0.90355 | 2.7902 | 1.6277 | 2.9385 | 4.5241
Q96RU2 | Ubiquitin carboxyl-terminal hydrolase 28 | USP28 | 1:1 | Block_4 | 0.29776 | 9.9178 | FALSE | NA | 4.208 | NA | NA | 0.40232 | NA | 1.0477 | 2.5747
P40818 | Ubiquitin carboxyl-terminal hydrolase 8 | USP8 | 1:1 | Block_4 | -1.8114 | 6.7471 | FALSE | 12.082 | 7.0152 | -0.067082 | 1.8386 | 4.1813 | 1.1793 | 4.4837 | 6.1542
Q9BV40 | Vesicle-associated membrane protein 8 | VAMP8 | 1:1 | Block_4 | -2 | 9.2642 | FALSE | 21.2601 | 23.5767 | -0.48884 | 1.1103 | 3.8647 | 1.2418 | 4.2135 | 5.5617
Q7L8A9 | Tubulinyl-Tyr carboxypeptidase 1 | VASH1 | 1:1 | Block_4 | 1.0249 | 7.1558 | TRUE | NA | NA | NA | NA | 2.6156 | NA | 2.2384 | 4.4825
Q99536 | Synaptic vesicle membrane protein VAT-1 homolog | VAT1 | 1:1 | Block_4 | -1.2353 | 1.1792 | FALSE | 23.0222 | 12.0171 | NA | -0.14488 | 0.20631 | -0.30394 | 0.2693 | 1.0335
Q9UKW4 | Guanine nucleotide exchange factor VAV3 | VAV3 | 1:1 | Block_4 | 1.0641 | 9.5572 | TRUE | NA | NA | NA | 1.4277 | 4.0006 | 1.1086 | 4.364 | 6.5419
P15692 | Vascular endothelial growth factor A | VEGFA | 1:1 | Block_4 | -4.4849 | 6.5154 | TRUE | 6.4532 | 5.3198 | -0.31832 | 0.43992 | 0.91584 | 0.58802 | 1.4249 | 2.7178
P49765 | Vascular endothelial growth factor B | VEGFB | 1:1 | Block_4 | -2.5 | 4.5078 | TRUE | 30.8068 | 19.1444 | -0.54857 | 0.30541 | 0.87421 | -0.14601 | 0.65814 | 1.4087
O43915 | Vascular endothelial growth factor D | VEGFD | 1:1 | Block_4 | -6.1683 | 5.3633 | TRUE | 5.4466 | 6.4729 | -0.77004 | -0.13488 | 0.38956 | -0.44171 | 0.30217 | 1.1245
O15240 | Neurosecretory protein VGF | VGF | 1:1 | Block_4 | -1.4124 | 5.9173 | FALSE | 11.7914 | 11.9282 | 0.12078 | 0.81771 | 1.4021 | 0.030574 | 1.0715 | 2.3152
Q7Z5L0 | Vitelline membrane outer layer protein 1 homolog | VMO1 | 1:1 | Block_4 | -0.39347 | 4.0412 | TRUE | 17.8936 | 8.9803 | NA | -0.13401 | 0.76944 | NA | 0.60119 | 1.8475
O95498 | Pantetheine hydrolase VNN2 | VNN2 | 1:1 | Block_4 | -2.8162 | 2.5444 | FALSE | 10.464 | 4.0111 | -0.22509 | 0.51059 | 1.0646 | 0.05801 | 0.99582 | 1.7614
Q8NEZ2 | Vacuolar protein sorting-associated protein 37A | VPS37A | 1:1 | Block_4 | 0.41656 | 8.351 | TRUE | 5.2138 | 16.4853 | NA | 0.86437 | 3.553 | 1.32 | 3.5896 | 4.9784
O75351 | Vacuolar protein sorting-associated protein 4B | VPS4B | 1:1 | Block_4 | 1.0786 | 9.6263 | FALSE | NA | NA | NA | 1.2895 | 3.1829 | NA | 3.6366 | 5.0942
Q5VIR6 | Vacuolar protein sorting-associated protein 53 homolog | VPS53 | 1:1 | Block_4 | 0.5 | 8.3081 | TRUE | NA | NA | NA | NA | 1.7373 | NA | 2.224 | 3.4953
Q8N0Z9 | V-set and immunoglobulin domain-containing protein 10 | VSIG10 | 1:1 | Block_4 | -0.86195 | 8.4331 | TRUE | 16.4572 | 16.6788 | NA | -0.17609 | 0.67426 | -0.60419 | -0.03224 | 0.58409
Q86VR7 | V-set and immunoglobulin domain-containing protein 10-like | VSIG10L | 1:1 | Block_4 | -3 | 8.2659 | FALSE | 11.1375 | 8.2807 | -0.45778 | 0.064105 | 1.0679 | -0.19895 | 0.309 | 1.195
Q96IQ7 | V-set and immunoglobulin domain-containing protein 2 | VSIG2 | 1:1 | Block_4 | -1.2973 | 3.815 | TRUE | 12.698 | 7.8728 | -1.2668 | -0.050572 | 0.69015 | -0.51166 | 0.70443 | 2.6557
P62760 | Visinin-like protein 1 | VSNL1 | 1:1 | Block_4 | -2.8105 | 7.7521 | TRUE | 6.3218 | 7.0701 | -0.61798 | -0.070566 | 0.61063 | -0.32961 | 0.16677 | 1.1083
Q6PCB0 | von Willebrand factor A domain-containing protein 1 | VWA1 | 1:1 | Block_4 | -2.7601 | 4.8289 | TRUE | 13.4511 | 14.5315 | -0.61275 | 0.016956 | 1.0384 | -0.52159 | 0.32158 | 1.4115
Q2TAL6 | Brorin | VWC2 | 1:1 | Block_4 | -3.4166 | 5.2913 | TRUE | 10.557 | 5.6158 | -0.38588 | 0.45657 | 1.1613 | 0.009223 | 0.81266 | 1.9645
B2RUY7 | von Willebrand factor C domain-containing protein 2-like | VWC2L | 1:1 | Block_4 | -2.2261 | 13.1705 | TRUE | 14.0645 | 3.8573 | -0.063404 | 0.23461 | 0.72946 | -0.070937 | 0.56488 | 2.3561
Q92558 | Actin-binding protein WASF1 | WASF1 | 1:1 | Block_4 | 0.83483 | 7.7366 | FALSE | NA | NA | NA | NA | 2.8429 | NA | 2.7623 | 4.7466
Q9UPY6 | Actin-binding protein WASF3 | WASF3 | 1:1 | Block_4 | 0.67007 | 5.4814 | FALSE | 8.9351 | 5.6735 | NA | 0.776 | 4.107 | NA | 3.2401 | 5.3712
Q9Y3C0 | WASH complex subunit 3 | WASHC3 | 1:1 | Block_4 | 1.776 | 9.7153 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.8948
O15213 | WD repeat-containing protein 46 | WDR46 | 1:1 | Block_4 | 0.37979 | NA | NA | 5.9954 | 13.455 | NA | NA | 0.59362 | NA | NA | 2.1777
A4D1P6 | WD repeat-containing protein 91 | WDR91 | 1:1 | Block_4 | -2.9779 | 8.0527 | TRUE | 9.5648 | 7.6302 | -0.2965 | 1.123 | 2.3328 | 0.3097 | 1.3524 | 3.0019
Q9HC57 | WAP four-disulfide core domain protein 1 | WFDC1 | 1:1 | Block_4 | -0.59772 | NA | NA | 11.2142 | 10.4873 | -0.38217 | 0.26042 | 0.7876 | -0.17567 | 0.59427 | 1.1453
Q8WWY7 | WAP four-disulfide core domain protein 12 | WFDC12 | 1:1 | Block_4 | -3.5 | 3.0483 | TRUE | 7.7604 | 11.3574 | -1.1203 | -0.39555 | 1.1434 | -0.48441 | 0.62827 | 2.0271
Q96NZ8 | WAP, Kazal, immunoglobulin, Kunitz and NTR domain-containing protein 1 | WFIKKN1 | 1:1 | Block_4 | -1.6379 | 7.0209 | FALSE | 15.2019 | 6.1518 | -0.74357 | -0.011403 | 0.6362 | -0.57249 | 0.21282 | 1.0656
Q9Y5W5 | Wnt inhibitory factor 1 | WIF1 | 1:1 | Block_4 | -6.3692 | 5.1551 | TRUE | 6.8654 | 3.2379 | -0.72458 | 0.065361 | 0.67903 | -0.21103 | 0.47769 | 1.1106
Q9NZC7 | WW domain-containing oxidoreductase | WWOX | 1:1 | Block_4 | 0.38375 | 10.8101 | TRUE | NA | NA | NA | NA | 0.56718 | NA | NA | 0.68158
O00308 | NEDD4-like E3 ubiquitin-protein ligase WWP2 | WWP2 | 1:1 | Block_4 | -1.6682 | 7.5918 | TRUE | 13.1781 | 5.865 | -0.56407 | 0.56061 | 1.9637 | 0.27602 | 1.6476 | 3.0311
P47992 | Lymphotactin | XCL1 | 1:1 | Block_4 | -2.1963 | 8.5623 | TRUE | 11.6005 | 16.6961 | -2.0101 | -0.80476 | 0.36412 | -1.6258 | -0.46017 | 0.67173
P54577 | Tyrosine--tRNA ligase, cytoplasmic | YARS1 | 1:1 | Block_4 | 0.23039 | 8.4534 | FALSE | 7.4592 | 6.7818 | NA | 2.0951 | 7.1026 | 2.7705 | 7.6511 | 9.048
P07947 | Tyrosine-protein kinase Yes | YES1 | 1:1 | Block_4 | -1.5 | 9.6438 | TRUE | 16.7209 | 21.1006 | -0.4264 | 1.2873 | 4.3454 | 1.3294 | 3.9436 | 5.6201
Q9BW85 | Splicing factor YJU2 | YJU2 | 1:1 | Block_4 | 0.066649 | 7.925 | TRUE | NA | NA | NA | NA | 0.83176 | NA | 1.0178 | 2.3643
Q5VVQ6 | Ubiquitin thioesterase OTU1 | YOD1 | 1:1 | Block_4 | -1.4138 | 7.4858 | TRUE | 12.7318 | 22.1023 | -0.91199 | 1.0421 | 2.1899 | 0.016316 | 2.3305 | 3.8098
Q9H171 | Z-DNA-binding protein 1 | ZBP1 | 1:1 | Block_4 | 0.15784 | 7.8933 | FALSE | 9.4627 | 13.5037 | NA | 0.3818 | 0.94325 | NA | 0.90104 | 5.0629
Q13105 | Zinc finger and BTB domain-containing protein 17 | ZBTB17 | 1:1 | Block_4 | -1.3809 | 10.0321 | TRUE | 6.4056 | 5.2574 | -0.31548 | 0.36933 | 1.0899 | 0.35273 | 1.1122 | 2.7962
Q7Z7L7 | Protein zer-1 homolog | ZER1 | 1:1 | Block_4 | -0.64412 | 5.0795 | FALSE | 10.9012 | 5.1776 | -0.33691 | 0.17382 | 0.68249 | 0.31791 | 0.8908 | 1.6625
Q8WV99 | AN1-type zinc finger protein 2B | ZFAND2B | 1:1 | Block_4 | -2.4095 | 6.7221 | TRUE | 12.5651 | 6.3738 | -0.11949 | 0.54866 | 2.2327 | 0.72262 | 2.4415 | 3.8818
Q96K21 | Abscission/NoCut checkpoint regulator | ZFYVE19 | 1:1 | Block_4 | 0.37036 | 5.9924 | FALSE | 10.7684 | 7.7144 | NA | 2.0809 | 4.8838 | 2.1389 | 4.7956 | 6.8382
Q9Y6X8 | Zinc fingers and homeoboxes protein 2 | ZHX2 | 1:1 | Block_4 | 0.72922 | 9.0443 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7296
O60232 | Protein ZNRD2 | ZNRD2 | 1:1 | Block_4 | 0.69511 | 6.0949 | FALSE | 9.6384 | 8.2751 | NA | NA | 1.4177 | NA | 1.2305 | 3.2708
P21754 | Zona pellucida sperm-binding protein 3 | ZP3 | 1:1 | Block_4 | -1.5412 | 5.5357 | TRUE | 25.6983 | 19.4245 | NA | 0.16874 | 1.7786 | NA | 0.35112 | 2.0759
Q9BZC7 | ATP-binding cassette sub-family A member 2 | ABCA2 | 1:1 | Block_3 | 1.865 | 10.0443 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.2155
Q15018 | BRISC complex subunit Abraxas 2 | ABRAXAS2 | 1:1 | Block_3 | 4.0956 | 11.0332 | TRUE | NA | NA | NA | NA | 4.1815 | NA | NA | 5.8168
P09110 | 3-ketoacyl-CoA thiolase, peroxisomal | ACAA1 | 1:1 | Block_3 | -0.88263 | 7.4243 | FALSE | 11.5887 | 10.6365 | -0.60903 | 0.54379 | 1.666 | 0.49964 | 2.0549 | 3.7709
P11310 | Medium-chain specific acyl-CoA dehydrogenase, mitochondrial | ACADM | 1:1 | Block_3 | 0.75981 | 2.9946 | FALSE | 8.9933 | 4.3217 | NA | NA | 0.99405 | NA | NA | 1.4887
P15309 | Prostatic acid phosphatase | ACP3 | 1:1 | Block_3 | -2.07 | 6.712 | TRUE | 15.68 | 13.7365 | -0.89346 | 0.10572 | 0.66064 | -0.74464 | 0.11806 | 1.2035
P33121 | Long-chain-fatty-acid--CoA ligase 1 | ACSL1 | 1:1 | Block_3 | 1 | 9.5107 | FALSE | 15.9571 | 19.3795 | NA | NA | 1.3435 | NA | NA | 1.9917
P63267 | Actin, gamma-enteric smooth muscle | ACTG2 | 1:1 | Block_3 | 0.69703 | 9.5997 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O43707 | Alpha-actinin-4 | ACTN4 | 1:1 | Block_3 | 1 | 5.3125 | FALSE | NA | NA | NA | NA | 1.1909 | NA | NA | 2.2825
Q96HD9 | N-acyl-aromatic-L-amino acid amidohydrolase | ACY3 | 1:1 | Block_3 | -5.3189 | 4.6345 | TRUE | 8.2922 | 13.5071 | -0.54366 | -0.043508 | 0.97416 | -0.55429 | 0.56777 | 1.4143
O43184 | Disintegrin and metalloproteinase domain-containing protein 12 | ADAM12 | 1:1 | Block_3 | 1.3 | 8.534 | FALSE | 10.671 | 0.18327 | NA | NA | NA | NA | NA | 1.9832
O75173 | A disintegrin and metalloproteinase with thrombospondin motifs 4 | ADAMTS4 | 1:1 | Block_3 | 2 | 12.6174 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P41586 | Pituitary adenylate cyclase-activating polypeptide type I receptor | ADCYAP1R1 | 1:1 | Block_3 | 0.57064 | 7.4615 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86Y34 | Adhesion G protein-coupled receptor G3 | ADGRG3 | 1:1 | Block_3 | -1.1136 | 7.7985 | FALSE | 14.0511 | 6.6472 | NA | -0.69936 | 0.69043 | NA | -0.35022 | 0.59628
Q8WXG9 | Adhesion G-protein coupled receptor V1 | ADGRV1 | 1:1 | Block_3 | 1.1 | 11.1857 | TRUE | NA | NA | NA | NA | 1.4092 | NA | NA | 1.7929
Q9BV57 | Acireductone dioxygenase | ADI1 | 1:1 | Block_3 | 2.0391 | 8.3313 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P08913 | Alpha-2A adrenergic receptor | ADRA2A | 1:1 | Block_3 | 0.577 | 7.6565 | FALSE | 2.066 | 5.7543 | NA | NA | NA | NA | NA | NA
Q8N4X5 | Actin filament-associated protein 1-like 2 | AFAP1L2 | 1:1 | Block_3 | -0.52946 | 11.1394 | TRUE | NA | NA | NA | NA | -0.22981 | NA | 0.27526 | 2.1654
Q96P47 | Arf-GAP with GTPase, ANK repeat and PH domain-containing protein 3 | AGAP3 | 1:1 | Block_3 | 1.009 | 10.1196 | TRUE | 9.0118 | 8.9264 | NA | NA | 1.1953 | NA | NA | 1.8965
Q8TD06 | Anterior gradient protein 3 | AGR3 | 1:1 | Block_3 | -0.38 | 4.3573 | TRUE | 25.2693 | 20.1419 | NA | 0.078249 | 1.5944 | NA | 0.10471 | 1.5642
Q96BJ3 | Axin interactor, dorsalization-associated protein | AIDA | 1:1 | Block_3 | 1.76 | 6.4654 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.8956
P55008 | Allograft inflammatory factor 1 | AIF1 | 1:1 | Block_3 | 0.52152 | 8.5521 | FALSE | NA | NA | NA | NA | 0.85593 | NA | NA | 1.7213
Q9UKA4 | A-kinase anchor protein 11 | AKAP11 | 1:1 | Block_3 | -2.3014 | 8.2825 | TRUE | 17.4855 | 3.8708 | NA | -2.1751 | -0.67423 | NA | -1.8588 | -0.30893
Q5T1N1 | Protein AKNAD1 | AKNAD1 | 1:1 | Block_3 | 0.71119 | 12.2824 | FALSE | NA | NA | NA | NA | 0.90026 | NA | NA | NA
P15121 | Aldo-keto reductase family 1 member B1 | AKR1B1 | 1:1 | Block_3 | 0.82655 | 4.2743 | FALSE | 6.339 | 9.5059 | NA | 1.3618 | 3.0118 | 1.4501 | 2.9812 | 4.2551
O60218 | Aldo-keto reductase family 1 member B10 | AKR1B10 | 1:1 | Block_3 | 0.001 | 9.2828 | FALSE | NA | 14.0571 | NA | NA | 2.7786 | NA | NA | 1.5326
P17516 | Aldo-keto reductase family 1 member C4 | AKR1C4 | 1:1 | Block_3 | 0.78862 | 7.2176 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NHP1 | Aflatoxin B1 aldehyde reductase member 4 | AKR7L | 1:1 | Block_3 | 1.9406 | 8.7898 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.5588
Q9Y243 | RAC-gamma serine/threonine-protein kinase | AKT3 | 1:1 | Block_3 | 2.7219 | 12.8355 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P49189 | 4-trimethylaminobutyraldehyde dehydrogenase | ALDH9A1 | 1:1 | Block_3 | 2.443 | 8.4926 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.2132
Q8TCU4 | Centrosome-associated protein ALMS1 | ALMS1 | 1:1 | Block_3 | 1.07 | 11.0037 | TRUE | NA | 2.7037 | NA | NA | 1.2104 | NA | NA | NA
P09923 | Intestinal-type alkaline phosphatase | ALPI | 1:1 | Block_3 | -2.948 | 4.1105 | TRUE | 15.2248 | 15.8644 | NA | -0.65271 | 1.0658 | NA | -0.76552 | 1.6416
Q9NP70 | Ameloblastin | AMBN | 1:1 | Block_3 | 0.77 | 10.355 | FALSE | NA | 5.9417 | NA | NA | 1.0127 | NA | NA | 1.0752
Q9UKV5 | E3 ubiquitin-protein ligase AMFR | AMFR | 1:1 | Block_3 | 0.64304 | 8.7833 | FALSE | NA | 4.9112 | NA | NA | NA | NA | NA | 1.0999
Q9BXJ7 | Protein amnionless | AMN | 1:1 | Block_3 | -0.84 | 10.5301 | FALSE | 20.2508 | 27.5326 | -0.32047 | 0.65586 | 1.3194 | -0.70639 | 0.63827 | 2.0213
Q9Y2J4 | Angiomotin-like protein 2 | AMOTL2 | 1:1 | Block_3 | 1.2074 | 7.4798 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.5534
Q01484 | Ankyrin-2 | ANK2 | 1:1 | Block_3 | -0.57952 | 10.9534 | TRUE | 10.4493 | 20.2883 | NA | NA | 1.5422 | NA | NA | 2.0955
Q9H9E1 | Ankyrin repeat family A protein 2 | ANKRA2 | 1:1 | Block_3 | 1.7811 | 8.5863 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6UB99 | Ankyrin repeat domain-containing protein 11 | ANKRD11 | 1:1 | Block_3 | -1.7756 | 10.9746 | TRUE | 14.5069 | 11.5797 | NA | -0.33129 | 2.4263 | -1.5911 | -0.37605 | 2.5961
Q8N283 | Ankyrin repeat domain-containing protein 35 | ANKRD35 | 1:1 | Block_3 | -0.46671 | 12.4363 | TRUE | 20.184 | 19.3576 | NA | -0.039006 | 0.81341 | -0.20427 | 0.88776 | 1.8714
Q8N9V6 | Ankyrin repeat domain-containing protein 53 | ANKRD53 | 1:1 | Block_3 | -1.051 | 9.8871 | TRUE | 9.2254 | 19.0382 | -1.05 | -0.60749 | 0.44247 | -1.0407 | -0.53296 | 0.05633
Q6NXT1 | Ankyrin repeat domain-containing protein 54 | ANKRD54 | 1:1 | Block_3 | 0.23878 | 11.5435 | FALSE | 19.7708 | 17.3182 | NA | 0.29525 | 1.4702 | NA | 0.92355 | 1.8908
Q3KP44 | Ankyrin repeat domain-containing protein 55 | ANKRD55 | 1:1 | Block_3 | -2.0827 | 10.3429 | FALSE | 16.9191 | 6.2091 | NA | NA | -0.8022 | NA | NA | -0.85382
Q9Y2G4 | Ankyrin repeat domain-containing protein 6 | ANKRD6 | 1:1 | Block_3 | 0.37499 | 10.3358 | TRUE | 0.60415 | 4.6568 | NA | NA | 0.72826 | NA | NA | 0.7995
Q9BTT0 | Acidic leucine-rich nuclear phosphoprotein 32 family member E | ANP32E | 1:1 | Block_3 | -3.9401 | 7.8692 | TRUE | 19.0882 | 8.9175 | -2.9069 | -1.8576 | -0.37215 | -1.8344 | -0.64849 | 3.4499
P08758 | Annexin A5 | ANXA5 | 1:1 | Block_3 | 3.7375 | 13.9938 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P19801 | Amiloride-sensitive amine oxidase [copper-containing] | AOC1 | 1:1 | Block_3 | 0.73 | 7.7257 | FALSE | NA | NA | NA | NA | 0.99801 | NA | NA | 1.0373
Q63HQ0 | AP-1 complex-associated regulatory protein | AP1AR | 1:1 | Block_3 | -0.0016917 | 12.3688 | TRUE | 13.4057 | 16.2432 | NA | 0.57376 | 1.8973 | 0.86397 | 2.055 | 3.6745
O43747 | AP-1 complex subunit gamma-1 | AP1G1 | 1:1 | Block_3 | 0.52539 | 10.5428 | FALSE | NA | NA | NA | NA | NA | NA | NA | 0.81291
O75843 | AP-1 complex subunit gamma-like 2 | AP1G2 | 1:1 | Block_3 | 1.9129 | 12.6723 | TRUE | NA | NA | NA | NA | 2.4236 | NA | 2.184 | 4.2736
P63010 | AP-2 complex subunit beta | AP2B1 | 1:1 | Block_3 | 1.4404 | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1.8939
P10398 | Serine/threonine-protein kinase A-Raf | ARAF | 1:1 | Block_3 | -0.67232 | 8.4056 | TRUE | NA | 2.0243 | NA | NA | -0.43953 | NA | NA | 0.16795
P78540 | Arginase-2, mitochondrial | ARG2 | 1:1 | Block_3 | 3.0495 | 8.4347 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q07960 | Rho GTPase-activating protein 1 | ARHGAP1 | 1:1 | Block_3 | 0.001 | 6.8928 | TRUE | 29.9309 | 22.9493 | 0.002942 | 0.94481 | 2.5575 | 0.62282 | 1.8696 | 2.8434
Q8IWW6 | Rho GTPase-activating protein 12 | ARHGAP12 | 1:1 | Block_3 | 0.001 | 11.367 | FALSE | 8.4518 | 10.1622 | NA | 0.79165 | 2.3561 | 1.0311 | 2.721 | 4.2274
P42331 | Rho GTPase-activating protein 25 | ARHGAP25 | 1:1 | Block_3 | 0.23842 | 10.438 | FALSE | NA | NA | NA | 0.45865 | 2.0862 | 0.81976 | 2.2653 | 3.878
Q13017 | Rho GTPase-activating protein 5 | ARHGAP5 | 1:1 | Block_3 | 1.4069 | 10.7192 | TRUE | 8.6847 | 8.459 | NA | NA | 1.9876 | NA | 1.5642 | 3.6849
Q12774 | Rho guanine nucleotide exchange factor 5 | ARHGEF5 | 1:1 | Block_3 | -1.431 | 10.1921 | TRUE | 23.8403 | 14.4201 | NA | -0.7681 | 0.79712 | -1.1366 | 0.043192 | 2.751
O14497 | AT-rich interactive domain-containing protein 1A | ARID1A | 1:1 | Block_3 | 1.0092 | 4.2867 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.1465
Q99856 | AT-rich interactive domain-containing protein 3A | ARID3A | 1:1 | Block_3 | 1.0324 | 3.8899 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q4LE39 | AT-rich interactive domain-containing protein 4B | ARID4B | 1:1 | Block_3 | 0.31336 | 12.1223 | TRUE | NA | NA | NA | NA | 0.34102 | NA | NA | 1.2048
Q9Y2Y0 | ADP-ribosylation factor-like protein 2-binding protein | ARL2BP | 1:1 | Block_3 | 0.64346 | 8.6919 | TRUE | 10.3627 | 2.1651 | NA | 1.0179 | 3.0232 | 1.1052 | 3.6015 | 5.1233
Q5W041 | Armadillo repeat-containing protein 3 | ARMC3 | 1:1 | Block_3 | 0.80056 | 10.6472 | TRUE | NA | NA | NA | NA | 1.2313 | NA | NA | NA
Q7L311 | Armadillo repeat-containing X-linked protein 2 | ARMCX2 | 1:1 | Block_3 | 1.1928 | 8.397 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P27540 | Aryl hydrocarbon receptor nuclear translocator | ARNT | 1:1 | Block_3 | 1.9421 | 7.2448 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96L15 | Ecto-ADP-ribosyltransferase 5 | ART5 | 1:1 | Block_3 | 2.2283 | 11.4 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5T4W7 | Artemin | ARTN | 1:1 | Block_3 | 1.4983 | 14.2085 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O43150 | Arf-GAP with SH3 domain, ANK repeat and PH domain-containing protein 2 | ASAP2 | 1:1 | Block_3 | 1.3744 | 8.4276 | FALSE | NA | NA | NA | NA | 4.8795 | NA | 4.1619 | 6.0772
Q86U10 | 60 kDa lysophospholipase | ASPG | 1:1 | Block_3 | 0.64951 | 8.6378 | TRUE | NA | 1.1346 | NA | NA | NA | NA | NA | NA
Q5T9A4 | ATPase family AAA domain-containing protein 3B | ATAD3B | 1:1 | Block_3 | -0.64502 | 12.388 | TRUE | 12.7011 | 32.4385 | -0.54809 | 0.15894 | 1.4699 | 0.45583 | 1.3136 | 2.6112
O95260 | Arginyl-tRNA--protein transferase 1 | ATE1 | 1:1 | Block_3 | 0.4242 | 9.615 | FALSE | 12.9649 | 8.9186 | NA | 1.8872 | 5.1817 | 1.4807 | 4.6896 | 6.4748
P15336 | Cyclic AMP-dependent transcription factor ATF-2 | ATF2 | 1:1 | Block_3 | 0.4667 | 7.3321 | FALSE | 5.7506 | 7.5399 | NA | NA | 1.911 | NA | NA | 1.2528
Q8WYN0 | Cysteine protease ATG4A | ATG4A | 1:1 | Block_3 | 0.001 | 9.1862 | FALSE | 33.9509 | 19.9071 | 0.22915 | 1.8026 | 3.2927 | 0.92484 | 2.6653 | 4.2212
P05026 | Sodium/potassium-transporting ATPase subunit beta-1 | ATP1B1 | 1:1 | Block_3 | 1.2064 | 8.9227 | TRUE | 13.1449 | 8.6333 | NA | NA | NA | NA | NA | NA
P14415 | Sodium/potassium-transporting ATPase subunit beta-2 | ATP1B2 | 1:1 | Block_3 | 2.8294 | 10.5301 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P54709 | Sodium/potassium-transporting ATPase subunit beta-3 | ATP1B3 | 1:1 | Block_3 | 1.8941 | 12.1248 | FALSE | NA | NA | NA | NA | 2.5715 | NA | NA | NA
Q9UN42 | Protein ATP1B4 | ATP1B4 | 1:1 | Block_3 | 0.56586 | 9.7783 | TRUE | NA | NA | NA | NA | 0.73381 | NA | NA | 1.1366
P23634 | Plasma membrane calcium-transporting ATPase 4 | ATP2B4 | 1:1 | Block_3 | 1.5869 | 9.0446 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.0493
P30049 | ATP synthase subunit delta, mitochondrial | ATP5F1D | 1:1 | Block_3 | 3.3516 | 12.2383 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O75947 | ATP synthase subunit d, mitochondrial | ATP5PD | 1:1 | Block_3 | 3.6796 | 14.0416 | FALSE | NA | NA | NA | NA | 4.7126 | NA | NA | 5.3465
P48047 | ATP synthase subunit O, mitochondrial | ATP5PO | 1:1 | Block_3 | 1.8686 | 10.0083 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O75787 | Renin receptor | ATP6AP2 | 1:1 | Block_3 | 0.001 | 3.0294 | TRUE | NA | NA | NA | NA | 0.02876 | NA | NA | 0.59061
Q16864 | V-type proton ATPase subunit F | ATP6V1F | 1:1 | Block_3 | 1.4077 | 6.0833 | FALSE | NA | NA | NA | 2.2863 | 4.7525 | 1.9122 | 4.5481 | 6.2204
Q9UBB4 | Ataxin-10 | ATXN10 | 1:1 | Block_3 | 1.47 | 8.2876 | TRUE | NA | NA | NA | NA | 2.3416 | NA | 2.3065 | 3.6869
Q99700 | Ataxin-2 | ATXN2 | 1:1 | Block_3 | -0.47 | 11.0807 | TRUE | NA | NA | NA | -0.20066 | 1.0405 | NA | 0.34618 | 2.3486
Q9NWV8 | BRISC and BRCA1-A complex member 1 | BABAM1 | 1:1 | Block_3 | 1.8564 | 8.3821 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O95429 | BAG family molecular chaperone regulator 4 | BAG4 | 1:1 | Block_3 | 1.09 | 9.8644 | TRUE | 9.5585 | NA | NA | NA | 1.7222 | NA | 2.305 | 4.7077
Q9UQB8 | Brain-specific angiogenesis inhibitor 1-associated protein 2 | BAIAP2 | 1:1 | Block_3 | -0.55532 | 8.73 | FALSE | 9.7167 | 3.4521 | NA | 0.27504 | 1.8264 | 0.12933 | 1.1645 | 2.8062
Q6UXY1 | Brain-specific angiogenesis inhibitor 1-associated protein 2-like protein 2 | BAIAP2L2 | 1:1 | Block_3 | 0.97235 | 12.5326 | TRUE | 19.0111 | 14.8795 | NA | NA | 1.4105 | NA | 1.0171 | 2.2496
Q13145 | BMP and activin membrane-bound inhibitor homolog | BAMBI | 1:1 | Block_3 | 1.4639 | 5.8957 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7332
Q6ZT62 | Bargin | BARGIN | 1:1 | Block_3 | 0.32304 | 7.5175 | TRUE | NA | 8.0743 | NA | 0.70717 | 2.9219 | 0.33302 | 2.7364 | 4.2784
Q16520 | Basic leucine zipper transcriptional factor ATF-like | BATF | 1:1 | Block_3 | -1.7656 | 13.9258 | TRUE | NA | NA | NA | NA | NA | NA | NA | -1.1053
Q07812 | Apoptosis regulator BAX | BAX | 1:1 | Block_3 | 0.39439 | 9.5436 | TRUE | NA | NA | NA | 1.068 | 2.135 | 1.3119 | 3.3554 | 5.3275
O15382 | Branched-chain-amino-acid aminotransferase, mitochondrial | BCAT2 | 1:1 | Block_3 | 0.68838 | 10.5269 | FALSE | NA | NA | NA | NA | 1.4616 | NA | NA | 0.92783
Q5TBC7 | Bcl-2-like protein 15 | BCL2L15 | 1:1 | Block_3 | 0.69 | 5.4822 | TRUE | NA | NA | NA | NA | 1.2396 | NA | NA | 4.2096
Q4VC05 | B-cell CLL/lymphoma 7 protein family member A | BCL7A | 1:1 | Block_3 | 1.1436 | 8.6262 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BUT1 | Dehydrogenase/reductase SDR family member 6 | BDH2 | 1:1 | Block_3 | -0.015919 | 10.4703 | TRUE | 7.1891 | 10.6794 | NA | NA | 0.64087 | NA | 0.22361 | 1.3594
Q00994 | Protein BEX3 | BEX3 | 1:1 | Block_3 | -2.4552 | 7.3324 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P21810 | Biglycan | BGN | 1:1 | Block_3 | 0.001 | #VALUE! | FALSE | 12.677 | 12.1331 | NA | 0.51497 | 1.6776 | NA | 0.73546 | 1.7949
O14503 | Class E basic helix-loop-helix protein 40 | BHLHE40 | 1:1 | Block_3 | 1.018 | 11.3956 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.1548
Q9H2M3 | S-methylmethionine--homocysteine S-methyltransferase BHMT2 | BHMT2 | 1:1 | Block_3 | 1.3497 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
A1A5D9 | BICD family-like cargo adapter 2 | BICDL2 | 1:1 | Block_3 | 0.70257 | 12.086 | TRUE | 3.8362 | NA | NA | NA | 0.75451 | NA | NA | NA
Q9NQY0 | Bridging integrator 3 | BIN3 | 1:1 | Block_3 | 1.8543 | 8.4239 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.1629
Q13490 | Baculoviral IAP repeat-containing protein 2 | BIRC2 | 1:1 | Block_3 | 2.3648 | 10.743 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6QNY1 | Biogenesis of lysosome-related organelles complex 1 subunit 2 | BLOC1S2 | 1:1 | Block_3 | 1.6187 | 11.4287 | TRUE | 10.7092 | 35.7025 | NA | NA | NA | NA | NA | NA
A0JNW5 | UHRF1-binding protein 1-like | BLTP3B | 1:1 | Block_3 | 0.1264 | 10.3565 | TRUE | NA | 14.5906 | NA | NA | 1.2682 | NA | 0.90528 | 5.2595
P53004 | Biliverdin reductase A | BLVRA | 1:1 | Block_3 | 0.64145 | 7.1978 | FALSE | NA | NA | NA | NA | 1.517 | NA | 1.1316 | 2.3519
Q96MC5 | bMERB domain-containing protein 1 | BMERB1 | 1:1 | Block_3 | 0.046476 | 11.1076 | TRUE | 23.7472 | 23.2289 | NA | 0.47676 | 1.1362 | 0.1117 | 0.83632 | 1.6847
Q9Y3E2 | BolA-like protein 1 | BOLA1 | 1:1 | Block_3 | 1.2575 | 11.2261 | FALSE | NA | NA | NA | NA | 1.9326 | NA | 2.1317 | 3.3539
O95696 | Bromodomain-containing protein 1 | BRD1 | 1:1 | Block_3 | 1.455 | 10.981 | TRUE | NA | NA | NA | NA | 1.8771 | NA | NA | NA
P25440 | Bromodomain-containing protein 2 | BRD2 | 1:1 | Block_3 | 0.43704 | 8.3436 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.1325
Q15059 | Bromodomain-containing protein 3 | BRD3 | 1:1 | Block_3 | 4.2782 | 15.0753 | TRUE | NA | NA | NA | NA | 5.1475 | NA | NA | 7.4103
Q0VDD7 | Break repair meiotic recombinase recruitment factor 1 | BRME1 | 1:1 | Block_3 | 0.90977 | 7.3665 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WZ55 | Barttin | BSND | 1:1 | Block_3 | -0.58723 | 10.5558 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5W0U4 | B box and SPRY domain-containing protein | BSPRY | 1:1 | Block_3 | 0.337 | 9.3331 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P35070 | Probetacellulin | BTC | 1:1 | Block_3 | 0.66537 | 13.4335 | TRUE | NA | NA | NA | NA | 1.5985 | NA | 0.68169 | 1.9497
Q7Z6A9 | B- and T-lymphocyte attenuator | BTLA | 1:1 | Block_3 | 0.49439 | 8.9233 | TRUE | NA | NA | NA | NA | 0.57406 | NA | NA | 1.668
A8MVZ5 | Putative butyrophilin-like protein 10 pseudogene | BTNL10P | 1:1 | Block_3 | 0.39 | 10.0786 | FALSE | 16.7206 | 14.8804 | NA | 0.39318 | 1.6283 | NA | NA | 1.1965
Q9H972 | Uncharacterized protein C14orf93 | C14orf93 | 1:1 | Block_3 | 1.4616 | 10.317 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NSK7 | Protein C19orf12 | C19orf12 | 1:1 | Block_3 | -0.17219 | 9.4826 | TRUE | 8.7715 | 7.943 | NA | -0.15927 | 1.3883 | NA | 0.47034 | 1.9925
O14523 | Phospholipid transfer protein C2CD2L | C2CD2L | 1:1 | Block_3 | 1.023 | 7.9199 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.737
Q8N8R5 | Mitochondrial protein C2orf69 | C2orf69 | 1:1 | Block_3 | 0.7404 | 9.2378 | FALSE | NA | NA | NA | NA | 1.151 | NA | NA | 1.9604
Q8N865 | Uncharacterized protein C7orf31 | C7orf31 | 1:1 | Block_3 | -0.23137 | 10.1012 | FALSE | 1.4026 | 3.3667 | NA | NA | NA | NA | NA | 0.27118
O75493 | Carbonic anhydrase-related protein 11 | CA11 | 1:1 | Block_3 | -0.16031 | 5.6484 | FALSE | 7.6953 | 8.9913 | NA | -0.10688 | 0.48934 | NA | NA | 0.30942
P35218 | Carbonic anhydrase 5A, mitochondrial | CA5A | 1:1 | Block_3 | -0.85959 | 9.8829 | TRUE | 17.8392 | 14.5641 | NA | -0.65302 | -0.15207 | NA | -0.27955 | 1.997
P43166 | Carbonic anhydrase 7 | CA7 | 1:1 | Block_3 | 1.6395 | 10.8025 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P35219 | Carbonic anhydrase-related protein | CA8 | 1:1 | Block_3 | 2.4982 | 7.5073 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NPB3 | Calcium-binding protein 2 | CABP2 | 1:1 | Block_3 | -0.30135 | 8.5972 | FALSE | NA | NA | NA | NA | -0.22673 | NA | NA | 0.33002
Q13936 | Voltage-dependent L-type calcium channel subunit alpha-1C | CACNA1C | 1:1 | Block_3 | 2.1253 | 12.1755 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O95180 | Voltage-dependent T-type calcium channel subunit alpha-1H | CACNA1H | 1:1 | Block_3 | 1.1137 | 9.7587 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q02641 | Voltage-dependent L-type calcium channel subunit beta-1 | CACNB1 | 1:1 | Block_3 | 3.2552 | 14.0463 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9ULU8 | Calcium-dependent secretion activator 1 | CADPS | 1:1 | Block_3 | 0.48213 | 12.9062 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P22676 | Calretinin | CALB2 | 1:1 | Block_3 | 0.77875 | 8.786 | TRUE | NA | 12.904 | NA | NA | NA | NA | NA | 1.0145
Q9NYX4 | Neuron-specific vesicular protein calcyon | CALY | 1:1 | Block_3 | 2.96 | 11.9782 | FALSE | 10.8158 | NA | NA | NA | 4.4829 | NA | NA | 4.9853
Q16566 | Calcium/calmodulin-dependent protein kinase type IV | CAMK4 | 1:1 | Block_3 | 0.27046 | 12.0776 | TRUE | NA | NA | NA | NA | 0.35253 | NA | NA | 1.8737
Q8N5S9 | Calcium/calmodulin-dependent protein kinase kinase 1 | CAMKK1 | 1:1 | Block_3 | 1.1777 | 10.6679 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8882
P49069 | Guided entry of tail-anchored proteins factor CAMLG | CAMLG | 1:1 | Block_3 | 0.8402 | 12.8611 | TRUE | NA | NA | NA | NA | 0.91567 | NA | NA | 2.4171
Q08AD1 | Calmodulin-regulated spectrin-associated protein 2 | CAMSAP2 | 1:1 | Block_3 | 1.3494 | 12.5235 | FALSE | NA | NA | NA | NA | 2.7791 | NA | NA | 1.942
P40123 | Adenylyl cyclase-associated protein 2 | CAP2 | 1:1 | Block_3 | 0.86008 | 9.235 | FALSE | 3.2646 | 7.2411 | NA | NA | NA | NA | NA | 1.9497
Q14444 | Caprin-1 | CAPRIN1 | 1:1 | Block_3 | -0.18803 | 6.8114 | TRUE | 11.479 | 14.2382 | NA | 0.59374 | 2.0946 | 0.63926 | 1.8312 | 3.2427
Q9Y2V2 | Calcium-regulated heat-stable protein 1 | CARHSP1 | 1:1 | Block_3 | -0.46027 | 6.3014 | FALSE | 20.1242 | 13.0123 | 0.32262 | 1.9687 | 2.8068 | 1.5423 | 2.4569 | 3.5204
O15234 | Protein CASC3 | CASC3 | 1:1 | Block_3 | 1.02 | 6.9248 | TRUE | NA | 17.4422 | NA | 1.1406 | 2.7707 | NA | 2.9677 | 4.959
P29466 | Caspase-1 | CASP1 | 1:1 | Block_3 | 0.012939 | 7.4244 | TRUE | 6.8294 | 5.1383 | NA | NA | 0.8883 | NA | 0.42602 | 5.1097
P49662 | Caspase-4 | CASP4 | 1:1 | Block_3 | -4.1366 | 8.101 | TRUE | 14.4037 | 41.4629 | NA | NA | -1.3921 | NA | -4.1222 | -1.5655
O14958 | Calsequestrin-2 | CASQ2 | 1:1 | Block_3 | 1.9971 | 7.1335 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P20810 | Calpastatin | CAST | 1:1 | Block_3 | 0.10661 | 5.7188 | TRUE | 11.2401 | 10.7583 | 0.10996 | 0.85516 | 1.4838 | 0.26575 | 1.2318 | 2.7016
Q13191 | E3 ubiquitin-protein ligase CBL-B | CBLB | 1:1 | Block_3 | 0.67781 | 12.7526 | FALSE | NA | NA | NA | NA | 0.68742 | NA | NA | 1.4656
P23435 | Cerebellin-1 | CBLN1 | 1:1 | Block_3 | 1.8602 | 10.7039 | TRUE | NA | NA | NA | NA | 2.8094 | NA | NA | 2.086
Q96A19 | Coiled-coil domain-containing protein 102A | CCDC102A | 1:1 | Block_3 | 0.66 | 14.5907 | TRUE | NA | NA | NA | NA | 1.1261 | NA | 0.76704 | 2.3097
Q68D86 | Coiled-coil domain-containing protein 102B | CCDC102B | 1:1 | Block_3 | 0.2058 | 13.3822 | TRUE | NA | 12.1379 | NA | NA | 0.63032 | NA | NA | 0.42305
Q8NCX0 | Coiled-coil domain-containing protein 150 | CCDC150 | 1:1 | Block_3 | -1.4297 | 12.2009 | FALSE | NA | 14.9411 | NA | NA | -0.073932 | NA | NA | 1.8236
Q96PX6 | Coiled-coil domain-containing protein 85A | CCDC85A | 1:1 | Block_3 | -0.24072 | 13.3497 | TRUE | 16.9607 | 11.4378 | NA | 0.32164 | 0.86056 | 0.12525 | 0.94778 | 2.0557
P30279 | G1/S-specific cyclin-D2 | CCND2 | 1:1 | Block_3 | 0.93 | 13.8048 | FALSE | NA | 0.77983 | NA | NA | 1.3222 | NA | NA | 1.4628
P24864 | G1/S-specific cyclin-E1 | CCNE1 | 1:1 | Block_3 | 1.1371 | 10.9664 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P48643 | T-complex protein 1 subunit epsilon | CCT5 | 1:1 | Block_3 | 0.85 | 4.7781 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.6646
Q6UWJ8 | CD164 sialomucin-like 2 protein | CD164L2 | 1:1 | Block_3 | 1.0663 | 9.6385 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P06729 | T-cell surface antigen CD2 | CD2 | 1:1 | Block_3 | 1 | 10.2979 | FALSE | 0.45327 | NA | NA | NA | 1.0603 | NA | NA | 1.1108
P10747 | T-cell-specific surface glycoprotein CD28 | CD28 | 1:1 | Block_3 | 0.45922 | 7.3581 | TRUE | 13.9024 | 7.0386 | NA | NA | 1.2375 | NA | 0.83989 | 2.2146
P16671 | Platelet glycoprotein 4 | CD36 | 1:1 | Block_3 | 0.087256 | 8.9681 | FALSE | 9.4534 | 0.72233 | NA | NA | 0.5046 | NA | 0.55842 | 1.4994
P11049 | Leukocyte antigen CD37 | CD37 | 1:1 | Block_3 | 1.0835 | 2.5491 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P04234 | T-cell surface glycoprotein CD3 delta chain | CD3D | 1:1 | Block_3 | 0.91786 | 10.4859 | TRUE | NA | NA | NA | NA | 0.96901 | NA | NA | NA
P07766 | T-cell surface glycoprotein CD3 epsilon chain | CD3E | 1:1 | Block_3 | 0.92 | 13.6592 | FALSE | NA | 8.3053 | NA | NA | 1.5452 | NA | NA | 3.0663
P09693 | T-cell surface glycoprotein CD3 gamma chain | CD3G | 1:1 | Block_3 | -0.33674 | 10.0595 | TRUE | NA | 6.2894 | NA | NA | NA | NA | NA | -0.11448
P27701 | CD82 antigen | CD82 | 1:1 | Block_3 | -0.53635 | 7.7545 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O75794 | Cell division cycle protein 123 homolog | CDC123 | 1:1 | Block_3 | 1.2925 | 9.2446 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P30304 | M-phase inducer phosphatase 1 | CDC25A | 1:1 | Block_3 | 0.95741 | 11.9386 | TRUE | 5.2983 | 3.9567 | NA | NA | 1.0255 | NA | NA | NA
P30260 | Cell division cycle protein 27 homolog | CDC27 | 1:1 | Block_3 | 2.8529 | 12.7598 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.1778
Q6DT37 | Serine/threonine-protein kinase MRCK gamma | CDC42BPG | 1:1 | Block_3 | 1.4982 | 13.698 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.7525
P55283 | Cadherin-4 | CDH4 | 1:1 | Block_3 | 1.227 | 9.1453 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y2V0 | CDAN1-interacting nuclease 1 | CDIN1 | 1:1 | Block_3 | 1.1601 | 8.813 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UQ88 | Cyclin-dependent kinase 11A | CDK11A | 1:1 | Block_3 | 2.6046 | 10.5478 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96JB5 | CDK5 regulatory subunit-associated protein 3 | CDK5RAP3 | 1:1 | Block_3 | 2.9041 | 10.141 | FALSE | NA | NA | NA | NA | NA | NA | NA | 4.3713
Q86X02 | Cerebellar degeneration-related protein 2-like | CDR2L | 1:1 | Block_3 | -0.1167 | 12.7092 | TRUE | 6.3573 | 10.074 | NA | NA | 2.1164 | NA | 1.9228 | 4.078
A8MTB9 | Carcinoembryonic antigen-related cell adhesion molecule 18 | CEACAM18 | 1:1 | Block_3 | 0.88 | 10.1366 | FALSE | 8.5019 | 0.60661 | NA | NA | 1.062 | NA | NA | NA
Q6UY09 | Carcinoembryonic antigen-related cell adhesion molecule 20 | CEACAM20 | 1:1 | Block_3 | 3.8364 | 12.6958 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P40198 | Carcinoembryonic antigen-related cell adhesion molecule 3 | CEACAM3 | 1:1 | Block_3 | 0.097685 | 8.7157 | TRUE | 20.274 | 7.0808 | NA | 0.19214 | 0.80852 | NA | 0.58875 | 1.6505
P06731 | Carcinoembryonic antigen-related cell adhesion molecule 5 | CEACAM5 | 1:1 | Block_3 | 0.26 | 7.4674 | TRUE | 5.1342 | 13.8191 | NA | 0.6765 | 1.8196 | NA | 0.27456 | 2.1163
P49715 | CCAAT/enhancer-binding protein alpha | CEBPA | 1:1 | Block_3 | 0.75861 | 11.2771 | TRUE | NA | NA | NA | NA | 0.81017 | NA | NA | 1.548
Q9BXF3 | Chromatin remodeling regulator CECR2 | CECR2 | 1:1 | Block_3 | 0.2273 | 11.0687 | FALSE | NA | NA | NA | NA | 0.24339 | NA | NA | 0.23511
Q8N6W0 | CUGBP Elav-like family member 5 | CELF5 | 1:1 | Block_3 | 1 | 10.2152 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96J87 | CUGBP Elav-like family member 6 | CELF6 | 1:1 | Block_3 | 1.2567 | 11.4199 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UHN6 | Cell surface hyaluronidase | CEMIP2 | 1:1 | Block_3 | 0.31 | 10.0036 | FALSE | 7.5912 | 14.4173 | NA | 0.46258 | 1.4072 | NA | NA | 1.1018
Q8N111 | Cell cycle exit and neuronal differentiation protein 1 | CEND1 | 1:1 | Block_3 | -0.92183 | 10.6603 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.22539
P49454 | Centromere protein F | CENPF | 1:1 | Block_3 | 0.42704 | 13.338 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4641
Q9HC77 | Centromere protein J | CENPJ | 1:1 | Block_3 | 1.7409 | 11.9872 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.995
O60308 | Centrosomal protein of 104 kDa | CEP104 | 1:1 | Block_3 | 0.74184 | 11.8236 | TRUE | NA | NA | NA | 1.1147 | 3.2537 | 0.80369 | 2.9222 | 4.8379
Q8N8E3 | Centrosomal protein of 112 kDa | CEP112 | 1:1 | Block_3 | 0.44924 | 13.0779 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.61969
Q8N960 | Centrosomal protein of 120 kDa | CEP120 | 1:1 | Block_3 | -0.34975 | 11.6656 | TRUE | NA | NA | NA | NA | 0.69036 | NA | 0.17512 | 2.6407
Q6ZU80 | Centrosomal protein of 128 kDa | CEP128 | 1:1 | Block_3 | 1.1193 | 11.5957 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8063
O94986 | Centrosomal protein of 152 kDa | CEP152 | 1:1 | Block_3 | 1.2134 | 11.1081 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.2029
Q9UPV0 | Centrosomal protein of 164 kDa | CEP164 | 1:1 | Block_3 | 1.6821 | 10.9309 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.1932
Q9Y4F5 | Centrosomal protein of 170 kDa protein B | CEP170B | 1:1 | Block_3 | 1.0321 | 13.3928 | TRUE | 10.232 | 19.6355 | NA | NA | NA | NA | NA | 1.0499
Q96NB1 | Centrosomal protein 20 | CEP20 | 1:1 | Block_3 | 0.87351 | 12.1442 | FALSE | NA | NA | NA | NA | 2.0608 | NA | 1.9677 | 3.4667
O15078 | Centrosomal protein of 290 kDa | CEP290 | 1:1 | Block_3 | 2.0201 | 11.9503 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5VT06 | Centrosome-associated protein 350 | CEP350 | 1:1 | Block_3 | 1.3219 | 11.7127 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.279
Q9C0F1 | Centrosomal protein of 44 kDa | CEP44 | 1:1 | Block_3 | -0.26 | 12.9601 | TRUE | NA | NA | NA | 1.2436 | 4.0061 | 1.8663 | 4.1648 | 5.9252
Q53EZ4 | Centrosomal protein of 55 kDa | CEP55 | 1:1 | Block_3 | 0.23202 | 13.6015 | FALSE | 1.7944 | 13.8817 | NA | NA | 0.27675 | NA | NA | 2.6825
Q6P2H3 | Centrosomal protein of 85 kDa | CEP85 | 1:1 | Block_3 | 0.98266 | 10.3028 | TRUE | NA | NA | NA | NA | 1.3408 | NA | 1.2138 | 3.0788
Q9Y5P4 | Ceramide transfer protein | CERT1 | 1:1 | Block_3 | 2.5 | NA | NA | NA | NA | NA | NA | NA | NA | NA | 2.9134
P41208 | Centrin-2 | CETN2 | 1:1 | Block_3 | -0.68 | 7.8622 | FALSE | 6.8694 | 3.9743 | NA | 0.53787 | 2.2867 | 0.11085 | 2.2026 | 3.8668
O43822 | Cilia- and flagella-associated protein 410 | CFAP410 | 1:1 | Block_3 | 0.42338 | 12.8222 | TRUE | 14.9189 | 7.933 | NA | NA | 1.9076 | NA | 0.96573 | 3.347
Q9UL16 | Cilia- and flagella-associated protein 45 | CFAP45 | 1:1 | Block_3 | 0.085397 | 13.1057 | TRUE | 15.0952 | 1.4275 | NA | 0.10917 | 1.4102 | NA | 1.118 | 2.5536
P0CG37 | Cryptic protein | CFC1 | 1:1 | Block_3 | -1.1817 | 10.7438 | TRUE | 14.5523 | 15.3329 | -0.61505 | 0.28674 | 1.0081 | 0.041519 | 0.78188 | 1.9326
Q9UEE9 | Craniofacial development protein 1 | CFDP1 | 1:1 | Block_3 | 1.9595 | 13.1856 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.4806
Q9P2M7 | Cingulin | CGN | 1:1 | Block_3 | 1.5912 | 12.1737 | TRUE | 24.7627 | 8.4101 | NA | NA | NA | NA | NA | 1.9512
Q8WUX2 | Glutathione-specific gamma-glutamylcyclotransferase 2 | CHAC2 | 1:1 | Block_3 | -0.17552 | NA | NA | NA | NA | NA | 0.14157 | 1.3705 | NA | 1.2786 | 2.381
Q9NX63 | MICOS complex subunit MIC19 | CHCHD3 | 1:1 | Block_3 | -0.2037 | 11.5391 | TRUE | 19.0162 | 4.1973 | NA | 1.3934 | 3.2706 | 1.6839 | 3.262 | 5.129
Q99653 | Calcineurin B homologous protein 1 | CHP1 | 1:1 | Block_3 | 2.1175 | 9.674 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P11229 | Muscarinic acetylcholine receptor M1 | CHRM1 | 1:1 | Block_3 | -0.29975 | 11.0207 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O15111 | Inhibitor of nuclear factor kappa-B kinase subunit alpha | CHUK | 1:1 | Block_3 | -2.7765 | 6.89 | FALSE | 16.5023 | 15.3947 | -2.2392 | -0.47658 | 2.1326 | 0.028151 | 1.8876 | 3.9276
Q9BW66 | Cyclin-dependent kinase 2-interacting protein | CINP | 1:1 | Block_3 | 2 | 3.3064 | FALSE | 6.0739 | NA | NA | NA | NA | NA | NA | NA
Q7Z460 | CLIP-associating protein 1 | CLASP1 | 1:1 | Block_3 | 1.1873 | 9.8045 | FALSE | NA | NA | NA | NA | 1.629 | NA | 1.5185 | 3.4042
P0C7M8 | C-type lectin domain family 2 member L | CLEC2L | 1:1 | Block_3 | 0.54 | 10.4798 | TRUE | 28.4084 | 20.7308 | NA | NA | 1.5222 | NA | 0.64564 | 2.1294
Q6EIG7 | C-type lectin domain family 6 member A | CLEC6A | 1:1 | Block_3 | -2.5 | 10.5869 | TRUE | 21.1712 | 12.9879 | -0.52184 | 0.2661 | 1.2162 | -0.44587 | 0.70084 | 2.2899
O14967 | Calmegin | CLGN | 1:1 | Block_3 | -0.75 | 10.2323 | TRUE | 30.9361 | 21.4674 | NA | 0.15437 | 1.274 | NA | 0.97442 | 2.4391
Q9NZA1 | Chloride intracellular channel protein 5 | CLIC5 | 1:1 | Block_3 | 2.0778 | 13.5554 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.1829
Q14677 | Clathrin interactor 1 | CLINT1 | 1:1 | Block_3 | 1.0332 | 9.6555 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P54105 | Methylosome subunit pICln | CLNS1A | 1:1 | Block_3 | 2.0111 | 10.8826 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.4171
Q9HAW4 | Claspin | CLSPN | 1:1 | Block_3 | 2.6278 | 14.8762 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5SYC1 | Clavesin-2 | CLVS2 | 1:1 | Block_3 | 0.57969 | 5.7607 | FALSE | NA | NA | NA | NA | 0.61699 | NA | NA | NA
Q8N0X4 | Citramalyl-CoA lyase, mitochondrial | CLYBL | 1:1 | Block_3 | 0.20536 | 10.2105 | FALSE | NA | NA | NA | NA | 1.7846 | NA | 0.54058 | 2.4083
Q9NQW8 | Cyclic nucleotide-gated cation channel beta-3 | CNGB3 | 1:1 | Block_3 | 1.5 | 13.0335 | TRUE | 4.8167 | 8.8669 | NA | NA | NA | NA | NA | NA
Q9Y2B0 | Protein canopy homolog 2 | CNPY2 | 1:1 | Block_3 | 0.70695 | 6.2474 | FALSE | 13.5827 | 11.224 | NA | NA | 0.88792 | NA | NA | 1.0483
Q96F85 | CB1 cannabinoid receptor-interacting protein 1 | CNRIP1 | 1:1 | Block_3 | 0.27 | 7.8132 | FALSE | 12.1611 | 16.0882 | 0.47937 | 1.2458 | 2.1918 | NA | 1.3067 | 2.4069
P26441 | Ciliary neurotrophic factor | CNTF | 1:1 | Block_3 | 2.9071 | 16.215 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q7Z7A1 | Centriolin | CNTRL | 1:1 | Block_3 | 0.36276 | 11.9651 | TRUE | NA | NA | NA | NA | 0.80945 | NA | 0.75827 | 3.1792
P53420 | Collagen alpha-4(IV) chain | COL4A4 | 1:1 | Block_3 | 1.3663 | 9.0492 | TRUE | 15.8722 | 15.2212 | NA | NA | NA | NA | NA | NA
Q14055 | Collagen alpha-2(IX) chain | COL9A2 | 1:1 | Block_3 | 1.6148 | 9.3518 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.0831
O14579 | Coatomer subunit epsilon | COPE | 1:1 | Block_3 | 1.0256 | 10.1961 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6QEF8 | Coronin-6 | CORO6 | 1:1 | Block_3 | 0.40852 | 10.0747 | FALSE | 0.62739 | 8.7115 | NA | NA | 0.64532 | NA | NA | 1.4037
P36551 | Oxygen-dependent coproporphyrinogen-III oxidase, mitochondrial | CPOX | 1:1 | Block_3 | 0.55565 | 6.1877 | FALSE | NA | 1.9167 | NA | NA | 0.99373 | NA | NA | 1.0979
Q5TA50 | Ceramide-1-phosphate transfer protein | CPTP | 1:1 | Block_3 | 1.9745 | 8.1017 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q66K79 | Carboxypeptidase Z | CPZ | 1:1 | Block_3 | -0.8801 | 5.961 | TRUE | 18.0206 | 15.4278 | NA | -0.2695 | 0.39412 | -0.71099 | 0.5646 | 2.1761
Q9BSW2 | EF-hand calcium-binding domain-containing protein 4B | CRACR2A | 1:1 | Block_3 | 0.8748 | 9.4644 | FALSE | NA | NA | NA | NA | 2.565 | NA | 1.735 | 4.9678
O43889 | Cyclic AMP-responsive element-binding protein 3 | CREB3 | 1:1 | Block_3 | 1.25 | 10.7748 | TRUE | 7.5701 | NA | NA | NA | NA | NA | NA | NA
Q9NS37 | CREB/ATF bZIP transcription factor | CREBZF | 1:1 | Block_3 | 1.4868 | 9.4094 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P06850 | Corticoliberin | CRH | 1:1 | Block_3 | -0.27 | 6.6164 | FALSE | 13.3537 | 13.8746 | NA | 0.73391 | 2.2658 | NA | 0.27657 | 2.2478
Q9P021 | Cysteine-rich PDZ-binding protein | CRIPT | 1:1 | Block_3 | 0.88334 | 9.5945 | TRUE | NA | NA | NA | NA | 1.0984 | NA | NA | 1.3854
O75718 | Cartilage-associated protein | CRTAP | 1:1 | Block_3 | 1.2152 | 2.0484 | FALSE | NA | NA | NA | NA | 1.38 | NA | 1.2675 | 2.1462
O43186 | Cone-rod homeobox protein | CRX | 1:1 | Block_3 | -1.6458 | 6.0592 | TRUE | NA | NA | NA | NA | NA | NA | NA | -1.1642
Q14894 | Ketimine reductase mu-crystallin | CRYM | 1:1 | Block_3 | 1.4606 | 5.0266 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.0494
P04141 | Granulocyte-macrophage colony-stimulating factor | CSF2 | 1:1 | Block_3 | 1.7611 | 15.937 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.3044
P09919 | Granulocyte colony-stimulating factor | CSF3 | 1:1 | Block_3 | -2.536 | 1.934 | FALSE | 12.0312 | 7.1052 | -0.87062 | -0.29896 | 0.46691 | -1.2356 | -0.089942 | 1.1371
P48730 | Casein kinase I isoform delta | CSNK1D | 1:1 | Block_3 | 1.7 | 8.3243 | TRUE | NA | 17.0378 | NA | NA | 1.8573 | NA | 1.8376 | 5.9296
O95196 | Chondroitin sulfate proteoglycan 5 | CSPG5 | 1:1 | Block_3 | 0.77 | 11.1648 | TRUE | NA | 4.1699 | NA | 0.99645 | 2.3818 | NA | 0.98154 | 2.0304
P78358 | Cancer/testis antigen 1 | CTAG1A_CTAG1B | 1:1 | Block_3 | 4.8369 | 14.6554 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q16619 | Cardiotrophin-1 | CTF1 | 1:1 | Block_3 | 1.0373 | 5.5688 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.2969
O43310 | CBP80/20-dependent translation initiation factor | CTIF | 1:1 | Block_3 | -0.02424 | 11.2572 | TRUE | 7.9269 | 17.1476 | NA | 2.1955 | 5.2731 | 1.6769 | 5.3681 | 7.3675
P35221 | Catenin alpha-1 | CTNNA1 | 1:1 | Block_3 | 1.0433 | 7.5856 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P17812 | CTP synthase 1 | CTPS1 | 1:1 | Block_3 | 1.76 | 5.1668 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.147
P56202 | Cathepsin W | CTSW | 1:1 | Block_3 | -1.1762 | 9.3998 | FALSE | 12.3536 | 6.8259 | -1.1739 | -0.12766 | 1.7442 | -0.50559 | 0.79643 | 2.0412
Q9NTM9 | Copper homeostasis protein cutC homolog | CUTC | 1:1 | Block_3 | -0.71473 | 7.2415 | FALSE | 14.8603 | 8.3755 | NA | -0.53412 | 1.5881 | NA | 0.19548 | 4.4152
Q86UP6 | CUB and zona pellucida-like domain-containing protein 1 | CUZD1 | 1:1 | Block_3 | 0.72679 | 10.201 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q2TBE0 | CWF19-like protein 2 | CWF19L2 | 1:1 | Block_3 | 0.8694 | 12.377 | TRUE | NA | NA | NA | NA | 0.9016 | NA | NA | 1.291
O95715 | C-X-C motif chemokine 14 | CXCL14 | 1:1 | Block_3 | 1.73 | 5.2592 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P00167 | Cytochrome b5 | CYB5A | 1:1 | Block_3 | 0.47678 | 11.1002 | TRUE | 16.433 | 5.637 | NA | NA | 1.4172 | NA | NA | 2.0391
Q07973 | 1,25-dihydroxyvitamin D(3) 24-hydroxylase, mitochondrial | CYP24A1 | 1:1 | Block_3 | 0.90088 | 8.9603 | FALSE | NA | 2.5803 | NA | NA | 1.1862 | NA | NA | NA
P08684 | Cytochrome P450 3A4 | CYP3A4 | 1:1 | Block_3 | 1.6469 | 10.7627 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O43739 | Cytohesin-3 | CYTH3 | 1:1 | Block_3 | 2.5659 | 13.6429 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86SG3 | Deleted in azoospermia protein 4 | DAZ4 | 1:1 | Block_3 | -0.5225 | 11.8482 | TRUE | 6.3666 | 8.8434 | NA | -0.47605 | -0.03507 | NA | -0.50762 | 0.35537
P43146 | Netrin receptor DCC | DCC | 1:1 | Block_3 | 1.0693 | 6.1106 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
A8MYV0 | Doublecortin domain-containing protein 2C | DCDC2C | 1:1 | Block_3 | 0.13377 | 14.3831 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96SD1 | Protein artemis | DCLRE1C | 1:1 | Block_3 | 1.7824 | 9.5896 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IZD4 | mRNA-decapping enzyme 1B | DCP1B | 1:1 | Block_3 | 0.26024 | 11.8476 | TRUE | 11.543 | 11.5941 | NA | NA | 1.1741 | NA | 0.80015 | 2.2457
Q9UJW0 | Dynactin subunit 4 | DCTN4 | 1:1 | Block_3 | 0.66688 | 9.6934 | TRUE | NA | NA | NA | NA | 2.0804 | NA | 1.7802 | 3.4348
Q96GG9 | DCN1-like protein 1 | DCUN1D1 | 1:1 | Block_3 | 1.3525 | 7.3807 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6PH85 | DCN1-like protein 2 | DCUN1D2 | 1:1 | Block_3 | -1.8242 | 8.6829 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BW61 | DET1- and DDB1-associated protein 1 | DDA1 | 1:1 | Block_3 | 0.66106 | 10.1222 | FALSE | NA | NA | NA | NA | 0.72364 | NA | NA | 1.3007
Q9UMR2 | ATP-dependent RNA helicase DDX19B | DDX19B | 1:1 | Block_3 | 1.9171 | 10.3806 | TRUE | NA | NA | NA | NA | 3.3496 | NA | 3.5359 | 5.1816
Q9BUQ8 | Probable ATP-dependent RNA helicase DDX23 | DDX23 | 1:1 | Block_3 | 0.65019 | 10.9337 | FALSE | NA | NA | NA | NA | 0.81607 | NA | 0.72841 | 2.3076
Q9UHL0 | ATP-dependent RNA helicase DDX25 | DDX25 | 1:1 | Block_3 | 2.82 | 10.1102 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O00148 | ATP-dependent RNA helicase DDX39A | DDX39A | 1:1 | Block_3 | 0.72789 | 9.8297 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.0543
Q9NQI0 | Probable ATP-dependent RNA helicase DDX4 | DDX4 | 1:1 | Block_3 | 1.27 | 14.5467 | TRUE | NA | 7.4221 | NA | NA | 1.317 | NA | NA | 1.5301
Q9Y6V7 | Probable ATP-dependent RNA helicase DDX49 | DDX49 | 1:1 | Block_3 | 0.64474 | 13.794 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.1481
Q86TM3 | Probable ATP-dependent RNA helicase DDX53 | DDX53 | 1:1 | Block_3 | 2.989 | 13.3546 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WTQ1 | Beta-defensin 104 | DEFB104A_DEFB104B | 1:1 | Block_3 | -0.01069 | 9.7718 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q30KQ8 | Beta-defensin 112 | DEFB112 | 1:1 | Block_3 | 0.57659 | 11.0739 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q30KQ4 | Beta-defensin 116 | DEFB116 | 1:1 | Block_3 | 0.65443 | 9.8146 | TRUE | 7.3226 | 23.3448 | NA | NA | NA | NA | NA | 0.81205
Q96PH6 | Defensin beta 118 | DEFB118 | 1:1 | Block_3 | 1.586 | 11.6316 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P78524 | DENN domain-containing protein 2B | DENND2B | 1:1 | Block_3 | 1.3372 | 12.2286 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O43583 | Density-regulated protein | DENR | 1:1 | Block_3 | 1.8057 | 10.9514 | TRUE | 17.8794 | NA | NA | NA | 2.9334 | NA | 2.7233 | 4.5289
Q14129 | Protein DGCR6 | DGCR6 | 1:1 | Block_3 | 1.3 | 11.2711 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.6876
Q9Y6T7 | Diacylglycerol kinase beta | DGKB | 1:1 | Block_3 | 1.3142 | 8.5031 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q13574 | Diacylglycerol kinase zeta | DGKZ | 1:1 | Block_3 | 0.73995 | 8.5086 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.95799
Q02127 | Dihydroorotate dehydrogenase | DHODH | 1:1 | Block_3 | 1.7791 | 8.5558 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y394 | Dehydrogenase/reductase SDR family member 7 | DHRS7 | 1:1 | Block_3 | -1.2877 | 9.0609 | FALSE | 9.4281 | 7.8901 | NA | NA | -0.10416 | NA | -1.275 | 0.20481
Q0P6D2 | Divergent protein kinase domain 1C | DIPK1C | 1:1 | Block_3 | 2.2093 | 7.6691 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q15398 | Disks large-associated protein 5 | DLGAP5 | 1:1 | Block_3 | 1.522 | 12.1789 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86X45 | Dynein axonemal assembly factor 11 | DNAAF11 | 1:1 | Block_3 | -0.39 | 12.9067 | TRUE | 7.2585 | 2.3979 | NA | NA | 0.59426 | NA | NA | 0.80255
Q8TE73 | Dynein axonemal heavy chain 5 | DNAH5 | 1:1 | Block_3 | 0.013146 | 8.2684 | FALSE | 5.1149 | 5.1393 | NA | NA | 0.37086 | NA | NA | 0.50496
P31689 | DnaJ homolog subfamily A member 1 | DNAJA1 | 1:1 | Block_3 | 3.47 | 10.0756 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WWF6 | DnaJ homolog subfamily B member 3 | DNAJB3 | 1:1 | Block_3 | 1.5521 | 14.6486 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96LL9 | DnaJ homolog subfamily C member 30, mitochondrial | DNAJC30 | 1:1 | Block_3 | 1.1405 | 8.8964 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5SXM8 | DNL-type zinc finger protein | DNLZ | 1:1 | Block_3 | -1 | 10.9767 | FALSE | 24.2887 | 19.3068 | NA | -0.46156 | 2.3905 | NA | -0.71989 | 2.3002
Q14184 | Double C2-like domain-containing protein beta | DOC2B | 1:1 | Block_3 | 2.3968 | 14.1974 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.8768
Q9BZ29 | Dedicator of cytokinesis protein 9 | DOCK9 | 1:1 | Block_3 | 0.49636 | 10.1562 | FALSE | 1.7766 | 5.8206 | NA | NA | 0.53961 | NA | NA | 0.67407
O60469 | Cell adhesion molecule DSCAM | DSCAM | 1:1 | Block_3 | 1.0427 | 8.4259 | TRUE | NA | NA | NA | NA | 1.088 | NA | NA | 1.2816
Q86UW9 | Probable E3 ubiquitin-protein ligase DTX2 | DTX2 | 1:1 | Block_3 | 0.76173 | 9.4514 | FALSE | NA | NA | NA | NA | 1.703 | NA | NA | 0.99424
Q8TDB6 | E3 ubiquitin-protein ligase DTX3L | DTX3L | 1:1 | Block_3 | -0.17463 | 10.8904 | TRUE | 0.11471 | 0.074497 | NA | 0.030957 | 1.1985 | 0.50415 | 1.611 | 3.2408
Q68J44 | Dual specificity phosphatase 29 | DUSP29 | 1:1 | Block_3 | 1.9727 | 10.998 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.141
P51452 | Dual specificity protein phosphatase 3 | DUSP3 | 1:1 | Block_3 | 0.7883 | 5.5973 | TRUE | NA | NA | NA | NA | 2.2959 | NA | 1.155 | 2.3498
Q14204 | Cytoplasmic dynein 1 heavy chain 1 | DYNC1H1 | 1:1 | Block_3 | 0.94158 | 11.6864 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.2315
P63172 | Dynein light chain Tctex-type 1 | DYNLT1 | 1:1 | Block_3 | 3.0483 | 9.7182 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O75521 | Enoyl-CoA delta isomerase 2 | ECI2 | 1:1 | Block_3 | 1.2022 | 8.5814 | FALSE | NA | NA | NA | NA | 1.3122 | NA | NA | 1.8045
Q19T08 | Endothelial cell-specific chemotaxis regulator | ECSCR | 1:1 | Block_3 | 0.58214 | 6.8888 | TRUE | NA | NA | NA | NA | 0.58953 | NA | NA | 1.4123
Q9BV94 | ER degradation-enhancing alpha-mannosidase-like protein 2 | EDEM2 | 1:1 | Block_3 | 0.71204 | 6.1878 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5VUJ9 | Dynein regulatory complex protein 8 | EFCAB2 | 1:1 | Block_3 | 1.5915 | 11.3476 | TRUE | NA | NA | NA | NA | 1.7249 | NA | NA | 2.4583
P52799 | Ephrin-B2 | EFNB2 | 1:1 | Block_3 | 0.64082 | 8.8257 | FALSE | NA | 6.6936 | NA | NA | 0.70342 | NA | NA | 1.0019
Q9UHF1 | Epidermal growth factor-like protein 7 | EGFL7 | 1:1 | Block_3 | -0.35032 | NA | NA | 11.0983 | 11.5974 | 0.4016 | 0.98915 | 1.7975 | 0.75939 | 1.5946 | 2.68
Q9H223 | EH domain-containing protein 4 | EHD4 | 1:1 | Block_3 | -0.75682 | 7.6244 | FALSE | 12.8025 | 5.9932 | 0.078802 | 1.3538 | 3.2052 | 1.5416 | 3.2764 | 4.6602
Q9NZJ5 | Eukaryotic translation initiation factor 2-alpha kinase 3 | EIF2AK3 | 1:1 | Block_3 | 1.7373 | 7.0131 | FALSE | NA | NA | NA | NA | 1.8132 | NA | NA | 4.3305
O60573 | Eukaryotic translation initiation factor 4E type 2 | EIF4E2 | 1:1 | Block_3 | 0.87917 | 9.2008 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.5402
P55010 | Eukaryotic translation initiation factor 5 | EIF5 | 1:1 | Block_3 | 1.49 | 11.0943 | FALSE | NA | NA | NA | NA | 2.096 | NA | 1.9988 | 3.4776
P63241 | Eukaryotic translation initiation factor 5A-1 | EIF5A | 1:1 | Block_3 | 0.77202 | 5.0794 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P26378 | ELAV-like protein 4 | ELAVL4 | 1:1 | Block_3 | 0.7113 | 12.7422 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5R3F8 | Protein phosphatase 1 regulatory subunit 29 | ELFN2 | 1:1 | Block_3 | -0.65037 | 10.1086 | TRUE | 19.4642 | 18.0376 | NA | -0.2146 | 0.4155 | NA | 0.043094 | 1.1546
Q15370 | Elongin-B | ELOB | 1:1 | Block_3 | 0.73 | 7.1818 | TRUE | 10.7087 | 21.4125 | NA | NA | 1.8997 | 1.059 | 2.4297 | 3.5391
Q96A84 | EMI domain-containing protein 1 | EMID1 | 1:1 | Block_3 | 0.47898 | 7.6681 | TRUE | NA | NA | NA | NA | 0.50421 | NA | NA | NA
Q8N8S7 | Protein enabled homolog | ENAH | 1:1 | Block_3 | 0.2428 | 8.2646 | TRUE | 7.9194 | 5.9538 | NA | 0.46629 | 1.1493 | NA | 1.0203 | 2.081
Q9UHY7 | Enolase-phosphatase E1 | ENOPH1 | 1:1 | Block_3 | 2 | 6.144 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q07075 | Glutamyl aminopeptidase | ENPEP | 1:1 | Block_3 | 1.9857 | 12.6864 | TRUE | NA | NA | NA | NA | 2.1209 | NA | NA | 2.6756
O43768 | Alpha-endosulfine | ENSA | 1:1 | Block_3 | 2.1401 | 8.6831 | FALSE | NA | NA | NA | NA | 2.5621 | NA | NA | 3.6447
Q9Y5L3 | Ectonucleoside triphosphate diphosphohydrolase 2 | ENTPD2 | 1:1 | Block_3 | -0.82413 | 11.5187 | TRUE | 25.8257 | 18.8842 | -0.64269 | 0.048331 | 0.79782 | -0.69589 | 0.27588 | 1.6142
Q9Y2J2 | Band 4.1-like protein 3 | EPB41L3 | 1:1 | Block_3 | -2.6789 | 7.9461 | TRUE | 15.2262 | 17.482 | -0.46144 | 0.72988 | 3.0284 | 1.3601 | 2.6817 | 4.2572
Q9HCM4 | Band 4.1-like protein 5 | EPB41L5 | 1:1 | Block_3 | 1.7586 | 12.2928 | TRUE | 8.7352 | 11.4806 | NA | NA | NA | NA | NA | NA
Q9H2F5 | Enhancer of polycomb homolog 1 | EPC1 | 1:1 | Block_3 | 1.0742 | 4.6345 | FALSE | NA | NA | NA | NA | 1.3245 | NA | NA | 1.3289
Q9HCE0 | Ectopic P granules protein 5 homolog | EPG5 | 1:1 | Block_3 | 0.33015 | 13.6316 | FALSE | NA | NA | NA | NA | 1.4203 | NA | 1.4707 | 3.2581
Q5JZY3 | Ephrin type-A receptor 10 | EPHA10 | 1:1 | Block_3 | 0.97745 | 10.8019 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P34913 | Bifunctional epoxide hydrolase 2 | EPHX2 | 1:1 | Block_3 | -0.24101 | 8.1278 | FALSE | NA | 8.1486 | NA | 0.82354 | 2.3605 | 0.030969 | 1.7805 | 2.9908
Q9Y6I3 | Epsin-1 | EPN1 | 1:1 | Block_3 | 1.4159 | 8.3404 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O15083 | ERC protein 2 | ERC2 | 1:1 | Block_3 | 2.8099 | 14.1783 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.9238
P07992 | DNA excision repair protein ERCC-1 | ERCC1 | 1:1 | Block_3 | 1.8574 | 10.7881 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P50548 | ETS domain-containing transcription factor ERF | ERF | 1:1 | Block_3 | -0.2946 | 10.035 | FALSE | 15.4 | 12.7898 | NA | 0.52931 | 1.6323 | 0.25392 | 1.7737 | 2.9961
Q8IV48 | 3'-5' exoribonuclease 1 | ERI1 | 1:1 | Block_3 | 0.087481 | 7.0977 | TRUE | NA | NA | NA | NA | 0.14514 | NA | 0.14891 | 2.9149
Q96DN0 | Endoplasmic reticulum resident protein 27 | ERP27 | 1:1 | Block_3 | 1.8539 | 10.821 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
B6SEH8 | Endogenous retrovirus group V member 1 Env polyprotein | ERVV-1 | 1:1 | Block_3 | 1.3918 | 10.1422 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q14674 | Separin | ESPL1 | 1:1 | Block_3 | -0.58937 | 12.8287 | TRUE | NA | NA | NA | NA | 0.55299 | NA | NA | -0.16524
P03372 | Estrogen receptor | ESR1 | 1:1 | Block_3 | 3.2114 | 13.0828 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P34910 | Protein EVI2B | EVI2B | 1:1 | Block_3 | 1.389 | 12.2282 | TRUE | NA | NA | NA | NA | 2.8648 | NA | NA | 1.5105
Q96CN4 | EVI5-like protein | EVI5L | 1:1 | Block_3 | 0.189 | 11.8068 | TRUE | NA | NA | NA | 1.2501 | 4.6111 | 1.2914 | 4.6662 | 6.5644
Q92817 | Envoplakin | EVPL | 1:1 | Block_3 | 2.5365 | 13.2191 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q01780 | Exosome component 10 | EXOSC10 | 1:1 | Block_3 | 1.069 | 8.5767 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.4027
Q9NQT5 | Exosome complex component RRP40 | EXOSC3 | 1:1 | Block_3 | 0.051517 | 11.9117 | FALSE | NA | NA | NA | NA | 0.57815 | NA | 0.23061 | 1.7914
Q92935 | Exostosin-like 1 | EXTL1 | 1:1 | Block_3 | 0.50096 | 8.6821 | TRUE | NA | NA | NA | NA | 0.6358 | NA | NA | NA
Q0Z7S8 | Fatty acid-binding protein 9 | FABP9 | 1:1 | Block_3 | 0.66073 | 10.3337 | FALSE | NA | NA | NA | NA | 1.1367 | NA | NA | 1.3149
Q96CS3 | FAS-associated factor 2 | FAF2 | 1:1 | Block_3 | -2.1183 | 10.3504 | FALSE | NA | NA | NA | NA | 0.34608 | NA | -1.253 | 3.7836
Q8IWE2 | Protein NOXP20 | FAM114A1 | 1:1 | Block_3 | 0.079139 | 3.94 | FALSE | 14.2902 | 12.0289 | NA | 0.16581 | 0.83627 | NA | 0.31095 | 1.173
Q96EK7 | Constitutive coactivator of peroxisome proliferator-activated receptor gamma | FAM120B | 1:1 | Block_3 | -1.4054 | 9.4674 | TRUE | 9.9088 | 9.4484 | -0.35669 | 0.88189 | 2.8491 | 1.0328 | 2.8701 | 4.0908
Q6P995 | Protein FAM171B | FAM171B | 1:1 | Block_3 | 1.6264 | 11.5002 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A4D161 | Protein FAM221A | FAM221A | 1:1 | Block_3 | -0.59758 | 11.8039 | TRUE | 10.577 | 14.3792 | -0.47193 | 1.107 | 3.6556 | 1.5204 | 3.9617 | 5.3775
Q8NEG4 | Protein FAM83F | FAM83F | 1:1 | Block_3 | 1.7524 | 9.4322 | TRUE | NA | NA | NA | NA | 2.1953 | NA | NA | 2.6376
Q9Y285 | Phenylalanine--tRNA ligase alpha subunit | FARSA | 1:1 | Block_3 | 2.522 | 11.8843 | TRUE | NA | NA | NA | NA | NA | NA | 3.088 | 5.2499
Q9NVF7 | F-box only protein 28 | FBXO28 | 1:1 | Block_3 | -0.75697 | 13.6425 | FALSE | 10.8142 | 18.199 | -0.69183 | -0.013867 | 0.66837 | -0.14284 | 0.68475 | 2.0838
Q9UKT4 | F-box only protein 5 | FBXO5 | 1:1 | Block_3 | -0.72901 | 10.9739 | TRUE | NA | 0.22158 | NA | NA | -0.36454 | NA | NA | -0.1434
P12319 | High affinity immunoglobulin epsilon receptor subunit alpha | FCER1A | 1:1 | Block_3 | -1.8 | 7.8853 | TRUE | 22.9777 | 15.8936 | -0.51992 | 0.26721 | 0.87287 | -0.25672 | 0.34796 | 1.0572
P31994 | Low affinity immunoglobulin gamma Fc region receptor II-b | FCGR2B | 1:1 | Block_3 | -1.7477 | 4.8001 | TRUE | 15.6516 | 15.0997 | -1.1766 | 0.41248 | 0.93202 | -1.5087 | 0.31314 | 1.0963
Q0JRZ9 | F-BAR domain only protein 2 | FCHO2 | 1:1 | Block_3 | 1.8043 | 12.9287 | TRUE | NA | NA | NA | 2.2547 | 4.3963 | 1.827 | 3.4685 | 6.1213
Q96P31 | Fc receptor-like protein 3 | FCRL3 | 1:1 | Block_3 | 0.5 | 9.7309 | FALSE | 3.4684 | 12.3152 | NA | NA | 0.79729 | NA | NA | 1.2093
Q6BAA4 | Fc receptor-like B | FCRLB | 1:1 | Block_3 | -0.79 | 8.8549 | TRUE | 16.5863 | 12.6482 | NA | 0.030454 | 0.56955 | -0.45052 | 0.72657 | 2.0055
P10109 | Adrenodoxin, mitochondrial | FDX1 | 1:1 | Block_3 | 1.3 | 7.792 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6P4F2 | Ferredoxin-2, mitochondrial | FDX2 | 1:1 | Block_3 | 6.69 | 16.5512 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P39748 | Flap endonuclease 1 | FEN1 | 1:1 | Block_3 | -0.3748 | 4.5758 | FALSE | 11.8752 | 4.2305 | NA | NA | 0.438 | NA | -0.067851 | 3.0584
P07332 | Tyrosine-protein kinase Fes/Fps | FES | 1:1 | Block_3 | 2.1125 | 4.667 | FALSE | NA | NA | NA | NA | 2.2128 | NA | NA | 3.5893
P61328 | Fibroblast growth factor 12 | FGF12 | 1:1 | Block_3 | 1.7459 | 12.2845 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NP95 | Fibroblast growth factor 20 | FGF20 | 1:1 | Block_3 | 0.93475 | 8.1275 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P12034 | Fibroblast growth factor 5 | FGF5 | 1:1 | Block_3 | -1.89 | 12.5931 | FALSE | 15.4598 | 10.6649 | -0.10832 | 0.75936 | 1.4424 | 0.46991 | 1.3582 | 2.3406
P10767 | Fibroblast growth factor 6 | FGF6 | 1:1 | Block_3 | 1.6536 | 10.1418 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P21781 | Fibroblast growth factor 7 | FGF7 | 1:1 | Block_3 | 0.97534 | 7.4127 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P31371 | Fibroblast growth factor 9 | FGF9 | 1:1 | Block_3 | 2.4215 | 14.9077 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P09769 | Tyrosine-protein kinase Fgr | FGR | 1:1 | Block_3 | -0.1084 | 9.3755 | TRUE | 16.1771 | 15.8366 | NA | 0.7096 | 1.9047 | 0.28521 | 1.4079 | 4.209
P07954 | Fumarate hydratase, mitochondrial | FH | 1:1 | Block_3 | 2.5 | 9.8794 | FALSE | 11.4428 | 9.687 | NA | NA | NA | NA | NA | 2.6438
P49789 | Bis(5'-adenosyl)-triphosphatase | FHIT | 1:1 | Block_3 | 0.001 | 8.337 | TRUE | 2.1852 | 18.5189 | NA | 1.4176 | 4.5385 | 1.2756 | 4.1808 | 5.5612
Q92562 | Polyphosphoinositide phosphatase | FIG4 | 1:1 | Block_3 | 1.0749 | 10.8473 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96AY3 | Peptidyl-prolyl cis-trans isomerase FKBP10 | FKBP10 | 1:1 | Block_3 | -1.5176 | 9.0046 | TRUE | 18.4458 | 13.9744 | -0.28011 | 0.987 | 1.9574 | 0.62005 | 1.6554 | 3.1464
Q5T1M5 | FK506-binding protein 15 | FKBP15 | 1:1 | Block_3 | -1.6377 | 6.5039 | FALSE | 11.7422 | 11.9037 | -0.28603 | 0.30405 | 0.88697 | 0.3455 | 1.1076 | 2.2988
Q9Y680 | Peptidyl-prolyl cis-trans isomerase FKBP7 | FKBP7 | 1:1 | Block_3 | 0.66837 | 10.0881 | FALSE | 8.4448 | 15.3056 | NA | NA | 0.77447 | NA | 0.67826 | 1.8536
Q01543 | Friend leukemia integration 1 transcription factor | FLI1 | 1:1 | Block_3 | 1.6 | 10.8632 | TRUE | NA | NA | NA | NA | 4.2911 | NA | 4.5922 | 7.1916
P36888 | Receptor-type tyrosine-protein kinase FLT3 | FLT3 | 1:1 | Block_3 | -0.15 | 7.7015 | TRUE | NA | NA | NA | NA | 0.4177 | NA | NA | 0.73694
Q96CP2 | FLYWCH family member 2 | FLYWCH2 | 1:1 | Block_3 | 0.6246 | 10.6902 | TRUE | NA | 2.698 | NA | 0.65148 | 1.6508 | 0.65632 | 1.8704 | 5.7332
Q06787 | Fragile X messenger ribonucleoprotein 1 | FMR1 | 1:1 | Block_3 | 1.6957 | 12.7292 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.7485
Q04609 | Glutamate carboxypeptidase 2 | FOLH1 | 1:1 | Block_3 | -0.3632 | 8.803 | FALSE | NA | 17.3875 | NA | NA | 0.75264 | NA | -0.069822 | 1.9613
P01100 | Protein c-Fos | FOS | 1:1 | Block_3 | 1.8646 | 12.9364 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P55318 | Hepatocyte nuclear factor 3-gamma | FOXA3 | 1:1 | Block_3 | 1.5729 | 12.9983 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UPW0 | Forkhead box protein J3 | FOXJ3 | 1:1 | Block_3 | 1.29 | 11.458 | TRUE | NA | NA | NA | NA | NA | NA | 1.3195 | 3.0441
O43524 | Forkhead box protein O3 | FOXO3 | 1:1 | Block_3 | 0.67986 | 8.2749 | FALSE | NA | NA | NA | NA | 1.2973 | NA | 1.1275 | 2.671
Q9Y2L6 | FERM domain-containing protein 4B | FRMD4B | 1:1 | Block_3 | 2.27 | 9.4619 | TRUE | NA | NA | NA | NA | 2.4451 | NA | NA | 4.0172
Q6ZUT3 | FERM domain-containing protein 7 | FRMD7 | 1:1 | Block_3 | 0.75059 | 9.3368 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P35637 | RNA-binding protein FUS | FUS | 1:1 | Block_3 | 0.3961 | 4.4445 | FALSE | 13.9526 | 11.8924 | NA | 0.40188 | 0.88374 | NA | 1.1484 | 2.8719
P19526 | Galactoside alpha-(1,2)-fucosyltransferase 1 | FUT1 | 1:1 | Block_3 | 0.15 | 8.5939 | FALSE | NA | NA | NA | NA | NA | NA | NA | 0.84669
Q96DB9 | FXYD domain-containing ion transport regulator 5 | FXYD5 | 1:1 | Block_3 | 1.7061 | 9.9971 | TRUE | NA | NA | NA | NA | 2.0813 | NA | 2.198 | 3.6736
Q9BQS8 | FYVE and coiled-coil domain-containing protein 1 | FYCO1 | 1:1 | Block_3 | -1.7621 | 10.4337 | TRUE | 31.7134 | 20.4522 | -0.65107 | 0.66953 | 2.7286 | 0.79793 | 2.2359 | 5.3674
Q9ULW2 | Frizzled-10 | FZD10 | 1:1 | Block_3 | 1.0552 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Q9H461 | Frizzled-8 | FZD8 | 1:1 | Block_3 | 0.90323 | 7.8522 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4079
P48169 | Gamma-aminobutyric acid receptor subunit alpha-4 | GABRA4 | 1:1 | Block_3 | 1.7057 | 12.1725 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q99259 | Glutamate decarboxylase 1 | GAD1 | 1:1 | Block_3 | 2.8172 | 11.0087 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q05329 | Glutamate decarboxylase 2 | GAD2 | 1:1 | Block_3 | 1.0686 | 9.958 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O75293 | Growth arrest and DNA damage-inducible protein GADD45 beta | GADD45B | 1:1 | Block_3 | 0.42726 | 7.6183 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6NT46 | G antigen 2A | GAGE2A | 1:1 | Block_3 | -0.23237 | 8.1384 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q14435 | Polypeptide N-acetylgalactosaminyltransferase 3 | GALNT3 | 1:1 | Block_3 | -2.0657 | 8.6125 | TRUE | 10.4774 | 9.5192 | -0.2184 | 0.57436 | 1.9637 | 0.2932 | 1.3125 | 2.2075
Q14353 | Guanidinoacetate N-methyltransferase | GAMT | 1:1 | Block_3 | 1.25 | 4.2903 | FALSE | NA | NA | NA | NA | 1.6712 | NA | NA | 1.9297
O43903 | Growth arrest-specific protein 2 | GAS2 | 1:1 | Block_3 | 0.84263 | 13.2267 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.3689
Q9UFP1 | Golgi-associated kinase 1A | GASK1A | 1:1 | Block_3 | 1.2238 | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1.2774
P23771 | Trans-acting T-cell-specific transcription factor GATA-3 | GATA3 | 1:1 | Block_3 | 1.2454 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
P32455 | Guanylate-binding protein 1 | GBP1 | 1:1 | Block_3 | -1.5502 | NA | NA | 18.1554 | 16.6254 | -1.2102 | -0.33378 | 0.80854 | -0.8551 | 0.49638 | 2.6638
P32456 | Guanylate-binding protein 2 | GBP2 | 1:1 | Block_3 | 0.81529 | 9.3781 | FALSE | NA | NA | NA | 1.2242 | 3.4171 | 1.9462 | 3.5933 | 5.2576
Q96PP9 | Guanylate-binding protein 4 | GBP4 | 1:1 | Block_3 | 1.16 | 5.3436 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.316
Q6ZN66 | Guanylate-binding protein 6 | GBP6 | 1:1 | Block_3 | 1.1617 | 7.3078 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P16383 | Intron Large complex component GCFC2 | GCFC2 | 1:1 | Block_3 | 0.80263 | 8.1793 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.4576
P01275 | Pro-glucagon | GCG | 1:1 | Block_3 | -0.24182 | 10.2037 | TRUE | 15.838 | 24.7158 | NA | 1.0599 | 2.9249 | NA | 1.5172 | 3.3068
P48507 | Glutamate--cysteine ligase regulatory subunit | GCLM | 1:1 | Block_3 | 0.001 | 4.9608 | FALSE | 16.3227 | 18.0854 | NA | 0.87255 | 1.8248 | 0.081234 | 1.8196 | 2.8869
P39905 | Glial cell line-derived neurotrophic factor | GDNF | 1:1 | Block_3 | 0.51244 | NA | NA | NA | 12.3967 | NA | NA | 0.8559 | NA | 0.70889 | 1.5139
O14893 | Gem-associated protein 2 | GEMIN2 | 1:1 | Block_3 | 0.41447 | 10.4973 | FALSE | NA | NA | NA | NA | 1.092 | NA | 0.99246 | 2.9199
Q9NWZ8 | Gem-associated protein 8 | GEMIN8 | 1:1 | Block_3 | -0.21314 | 10.3277 | TRUE | 15.8011 | 10.3418 | NA | 0.012751 | 1.4016 | NA | NA | 1.6548
P14136 | Glial fibrillary acidic protein | GFAP | 1:1 | Block_3 | -0.98033 | 7.3215 | FALSE | 10.8163 | 14.6366 | -0.90506 | 0.072836 | 0.9612 | -0.30223 | 0.97349 | 2.1258
P55789 | FAD-linked sulfhydryl oxidase ALR | GFER | 1:1 | Block_3 | -2.39 | 8.7594 | FALSE | 14.9416 | 8.3151 | -0.6274 | -0.054881 | 1.0898 | 0.082319 | 1.5504 | 2.9433
Q3B7J2 | Glucose-fructose oxidoreductase domain-containing protein 2 | GFOD2 | 1:1 | Block_3 | -0.37477 | NA | NA | NA | NA | NA | NA | 0.9988 | NA | NA | 1.3802
Q9UJY5 | ADP-ribosylation factor-binding protein GGA1 | GGA1 | 1:1 | Block_3 | 2.1626 | 6.6115 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NZ52 | ADP-ribosylation factor-binding protein GGA3 | GGA3 | 1:1 | Block_3 | -0.88063 | 9.9849 | TRUE | 11.998 | 10.8784 | -0.86926 | 1.0122 | 4.0324 | 1.348 | 4.4097 | 5.9971
Q86UU5 | Gametogenetin | GGN | 1:1 | Block_3 | 0.87586 | 10.3488 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P01242 | Growth hormone variant | GH2 | 1:1 | Block_3 | -2.16 | 0.66456 | FALSE | 33.2826 | 35.9253 | NA | -1.8169 | 1.5087 | NA | -1.8231 | 1.0621
Q02643 | Growth hormone-releasing hormone receptor | GHRHR | 1:1 | Block_3 | 1.5692 | 6.7092 | FALSE | 4.2506 | 9.6517 | NA | NA | 2.374 | NA | NA | NA
Q8ND71 | GTPase IMAP family member 8 | GIMAP8 | 1:1 | Block_3 | 0.62138 | 11.1067 | TRUE | NA | NA | NA | NA | 0.91729 | NA | NA | 2.046
P17302 | Gap junction alpha-1 protein | GJA1 | 1:1 | Block_3 | 0.28363 | 11.7463 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.41671
P48165 | Gap junction alpha-8 protein | GJA8 | 1:1 | Block_3 | 1.9612 | 9.9335 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P10070 | Zinc finger protein GLI2 | GLI2 | 1:1 | Block_3 | 1.1104 | 12.1164 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P43220 | Glucagon-like peptide 1 receptor | GLP1R | 1:1 | Block_3 | 1.3379 | 5.4185 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H1C3 | Glycosyltransferase 8 domain-containing protein 2 | GLT8D2 | 1:1 | Block_3 | 0.64 | 6.754 | FALSE | NA | NA | NA | NA | 1.103 | NA | NA | 1.4808
O75496 | Geminin | GMNN | 1:1 | Block_3 | 0.049436 | 13.9953 | TRUE | 9.5554 | 9.9363 | NA | NA | 0.20886 | NA | NA | 1.2747
Q9Y223 | Bifunctional UDP-N-acetylglucosamine 2-epimerase/N-acetylmannosamine kinase | GNE | 1:1 | Block_3 | 1.0393 | 6.491 | FALSE | NA | NA | NA | NA | 1.5495 | NA | 1.6908 | 2.9413
P63211 | Guanine nucleotide-binding protein G(T) subunit gamma-T1 | GNGT1 | 1:1 | Block_3 | 1.77 | 9.9462 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q92805 | Golgin subfamily A member 1 | GOLGA1 | 1:1 | Block_3 | 1.1978 | 11.0983 | TRUE | NA | NA | NA | 1.3873 | 5.0139 | 1.5876 | 4.9843 | 6.763
Q9NYA3 | Golgin subfamily A member 6A | GOLGA6A | 1:1 | Block_3 | -0.79699 | 13.2895 | TRUE | 7.9432 | 4.6455 | NA | NA | -0.5185 | NA | NA | -0.19944
Q08AF8 | Putative golgin subfamily A member 8F/8G | GOLGA8F_GOLGA8G | 1:1 | Block_3 | 1.15 | 11.3053 | TRUE | NA | NA | NA | NA | 1.3017 | NA | NA | 1.8525
Q8NBJ4 | Golgi membrane protein 1 | GOLM1 | 1:1 | Block_3 | -0.64096 | 5.7978 | TRUE | 20.4495 | 8.2829 | -0.23148 | 0.18855 | 0.5713 | -0.34961 | 0.25786 | 0.70473
Q9BRR8 | G patch domain-containing protein 1 | GPATCH1 | 1:1 | Block_3 | 0.42018 | 11.399 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.0425
Q96P66 | Probable G-protein coupled receptor 101 | GPR101 | 1:1 | Block_3 | 0.65216 | 10.0414 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NQ84 | G-protein coupled receptor family C group 5 member C | GPRC5C | 1:1 | Block_3 | 0.60541 | 11.1476 | FALSE | NA | NA | NA | NA | 1.383 | NA | NA | 2.4265
P36969 | Phospholipid hydroperoxide glutathione peroxidase | GPX4 | 1:1 | Block_3 | -1.4014 | NA | NA | NA | 3.5785 | NA | -0.89074 | -0.072097 | NA | 0.15932 | 1.7059
Q13002 | Glutamate receptor ionotropic, kainate 2 | GRIK2 | 1:1 | Block_3 | 0.88495 | 12.046 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q13224 | Glutamate receptor ionotropic, NMDA 2B | GRIN2B | 1:1 | Block_3 | 1.6731 | 9.1687 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9C0E4 | Glutamate receptor-interacting protein 2 | GRIP2 | 1:1 | Block_3 | 1.2855 | 12.0772 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.0354
Q4V328 | GRIP1-associated protein 1 | GRIPAP1 | 1:1 | Block_3 | 0.39228 | 8.7805 | FALSE | NA | 6.4077 | NA | 0.95563 | 3.5938 | 0.96528 | 3.6926 | 5.156
P34947 | G protein-coupled receptor kinase 5 | GRK5 | 1:1 | Block_3 | 0.091623 | 11.6227 | FALSE | NA | 3.4524 | NA | 0.59338 | 2.7772 | NA | 1.477 | 3.76
Q12849 | G-rich sequence factor 1 | GRSF1 | 1:1 | Block_3 | 1.32 | 11.7207 | FALSE | NA | NA | NA | NA | 3.7457 | NA | 2.6653 | 4.9152
A4D1B5 | Gamma-secretase-activating protein | GSAP | 1:1 | Block_3 | -0.64198 | 2.1856 | FALSE | 1.6114 | 4.0265 | NA | NA | -0.6298 | NA | NA | -0.032027
Q16772 | Glutathione S-transferase A3 | GSTA3 | 1:1 | Block_3 | -0.023757 | 9.641 | TRUE | 7.1011 | 7.8632 | NA | NA | 0.57735 | NA | -0.0020025 | 1.4223
Q03013 | Glutathione S-transferase Mu 4 | GSTM4 | 1:1 | Block_3 | -0.60596 | 6.7408 | FALSE | 15.643 | 19.1352 | NA | NA | 0.61106 | NA | NA | 0.98581
P09211 | Glutathione S-transferase P | GSTP1 | 1:1 | Block_3 | 2.3408 | 8.7597 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.2944
P35269 | General transcription factor IIF subunit 1 | GTF2F1 | 1:1 | Block_3 | 0.70351 | 8.4661 | TRUE | NA | NA | NA | NA | 0.95011 | NA | NA | 0.87608
Q9UHL9 | General transcription factor II-I repeat domain-containing protein 1 | GTF2IRD1 | 1:1 | Block_3 | 0.88134 | 7.5456 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BX10 | GTP-binding protein 2 | GTPBP2 | 1:1 | Block_3 | 2.1669 | 11.6977 | TRUE | NA | NA | NA | NA | 3.1282 | NA | 2.4988 | 4.723
P25092 | Guanylyl cyclase C | GUCY2C | 1:1 | Block_3 | 4.2909 | 10.5105 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q16774 | Guanylate kinase | GUK1 | 1:1 | Block_3 | 3.0708 | 8.4964 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q02539 | Histone H1.1 | H1-1 | 1:1 | Block_3 | -0.031471 | 10.4544 | FALSE | 11.214 | 13.8908 | NA | -0.002746 | 0.51538 | NA | 0.346 | 2.4254
O75409 | Huntingtin-interacting protein M | H2AP | 1:1 | Block_3 | 1.9264 | 13.0171 | TRUE | NA | NA | NA | NA | 2.4184 | NA | NA | NA
Q16836 | Hydroxyacyl-coenzyme A dehydrogenase, mitochondrial | HADH | 1:1 | Block_3 | 0.65 | 3.8122 | FALSE | 15.6704 | 8.0542 | NA | NA | 1.0475 | NA | NA | 1.2941
Q9Y450 | HBS1-like protein | HBS1L | 1:1 | Block_3 | 0.057116 | 10.7475 | TRUE | 8.0086 | 20.2648 | NA | 2.1381 | 5.0386 | 1.5324 | 4.5876 | 6.3045
Q9Y5Z7 | Host cell factor 2 | HCFC2 | 1:1 | Block_3 | 0.91215 | 11.7879 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
E2RYF7 | Protein PBMUCL2 | HCG22 | 1:1 | Block_3 | 0.83 | 11.7677 | TRUE | 8.652 | 38.2672 | NA | NA | NA | NA | NA | 1.3476
Q9BY41 | Histone deacetylase 8 | HDAC8 | 1:1 | Block_3 | 1.1846 | 4.6544 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UKV0 | Histone deacetylase 9 | HDAC9 | 1:1 | Block_3 | 1.9837 | 10.7798 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.0364
Q9NRV9 | Heme-binding protein 1 | HEBP1 | 1:1 | Block_3 | 0.34075 | 5.5208 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P62684 | Endogenous retrovirus group K member 113 Gag polyprotein | HERVK-113_GA113 | 1:1 | Block_3 | 1.38 | 11.6256 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4988
A2PYH4 | Probable ATP-dependent DNA helicase HFM1 | HFM1 | 1:1 | Block_3 | 1.63 | 9.756 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O14964 | Hepatocyte growth factor-regulated tyrosine kinase substrate | HGS | 1:1 | Block_3 | 0.80503 | 5.6721 | FALSE | NA | NA | NA | 1.3909 | 3.7907 | 1.2885 | 3.7012 | 5.0578
Q16665 | Hypoxia-inducible factor 1-alpha | HIF1A | 1:1 | Block_3 | 0.064625 | 12.5021 | TRUE | NA | NA | NA | NA | 1.9225 | NA | NA | 1.8769
O00291 | Huntingtin-interacting protein 1 | HIP1 | 1:1 | Block_3 | 0.89384 | 9.4244 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.2313
Q9H2X6 | Homeodomain-interacting protein kinase 2 | HIPK2 | 1:1 | Block_3 | -0.34959 | 9.1116 | TRUE | NA | NA | NA | NA | 0.54868 | NA | NA | -0.11653
P31629 | Transcription factor HIVEP2 | HIVEP2 | 1:1 | Block_3 | 1.7668 | 13.4642 | TRUE | 21.4457 | 16.8664 | NA | NA | NA | NA | NA | 1.9021
P52789 | Hexokinase-2 | HK2 | 1:1 | Block_3 | 2.6597 | 14.5647 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.5866
Q6NT76 | Homeobox-containing protein 1 | HMBOX1 | 1:1 | Block_3 | 1.1433 | 9.6294 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7901
Q01581 | Hydroxymethylglutaryl-CoA synthase, cytoplasmic | HMGCS1 | 1:1 | Block_3 | 1.3411 | 4.504 | FALSE | NA | 3.2435 | NA | NA | NA | NA | NA | NA
O75330 | Hyaluronan mediated motility receptor | HMMR | 1:1 | Block_3 | 2.7084 | 13.3161 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P20823 | Hepatocyte nuclear factor 1-alpha | HNF1A | 1:1 | Block_3 | 1.81 | 11.8913 | FALSE | 13.3146 | 45.0228 | NA | NA | 2.4221 | NA | NA | 2.2108
Q13151 | Heterogeneous nuclear ribonucleoprotein A0 | HNRNPA0 | 1:1 | Block_3 | -0.64162 | 11.3279 | FALSE | 8.484 | 18.0991 | NA | NA | 0.32135 | -0.60653 | 0.52215 | 3.6451
Q86YM7 | Homer protein homolog 1 | HOMER1 | 1:1 | Block_3 | 0.54 | 10.851 | TRUE | 22.6038 | 1.5778 | NA | NA | 0.89872 | NA | NA | 0.92688
Q9UM19 | Hippocalcin-like protein 4 | HPCAL4 | 1:1 | Block_3 | 0.32594 | 8.7196 | TRUE | NA | NA | NA | NA | 0.38412 | NA | NA | 1.172
O60760 | Hematopoietic prostaglandin D synthase | HPGDS | 1:1 | Block_3 | 1.24 | 8.7783 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P01112 | GTPase HRas | HRAS | 1:1 | Block_3 | 1.3385 | 9.9555 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96MM7 | Heparan-sulfate 6-O-sulfotransferase 2 | HS6ST2 | 1:1 | Block_3 | 0.13069 | 5.0682 | TRUE | 13.4366 | 15.2603 | NA | 0.35339 | 1.0135 | NA | 0.70752 | 1.6943
P37058 | 17-beta-hydroxysteroid dehydrogenase type 3 | HSD17B3 | 1:1 | Block_3 | 3.125 | 13.15 | FALSE | NA | NA | NA | NA | 5.2952 | NA | NA | 5.9239
P14625 | Endoplasmin | HSP90B1 | 1:1 | Block_3 | 1.2913 | 10.7297 | FALSE | NA | 26.4651 | NA | NA | NA | NA | NA | 1.3748
P08908 | 5-hydroxytryptamine receptor 1A | HTR1A | 1:1 | Block_3 | -0.13893 | 10.6691 | TRUE | 12.2283 | 15.9787 | NA | 0.070135 | 1.9975 | NA | NA | 0.81824
P28222 | 5-hydroxytryptamine receptor 1B | HTR1B | 1:1 | Block_3 | 2.3026 | 12.3402 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q05084 | Islet cell autoantigen 1 | ICA1 | 1:1 | Block_3 | 1.1644 | 10.451 | FALSE | NA | NA | NA | NA | 2.6401 | NA | 2.8444 | 4.555
P47928 | DNA-binding protein inhibitor ID-4 | ID4 | 1:1 | Block_3 | 1.121 | 10.7739 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BXS1 | Isopentenyl-diphosphate delta-isomerase 2 | IDI2 | 1:1 | Block_3 | 0.73393 | 10.5311 | FALSE | 22.5429 | 9.7929 | NA | NA | 0.92226 | NA | NA | 1.289
P09914 | Interferon-induced protein with tetratricopeptide repeats 1 | IFIT1 | 1:1 | Block_3 | 1.34 | 4.0579 | FALSE | NA | NA | NA | NA | 1.4388 | NA | NA | 2.1386
O14879 | Interferon-induced protein with tetratricopeptide repeats 3 | IFIT3 | 1:1 | Block_3 | 1.6575 | 11.7286 | TRUE | NA | NA | NA | NA | 2.2343 | NA | 2.6786 | 4.7218
P01579 | Interferon gamma | IFNG | 1:1 | Block_3 | 1.16 | 12.6989 | TRUE | NA | 11.3701 | NA | NA | 1.3931 | NA | NA | 2.2769
Q8IU54 | Interferon lambda-1 | IFNL1 | 1:1 | Block_3 | 1.7679 | 13.1491 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.437
Q8IZJ0 | Interferon lambda-2 | IFNL2 | 1:1 | Block_3 | 1.5389 | 4.8564 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P05000 | Interferon omega-1 | IFNW1 | 1:1 | Block_3 | 1.1019 | 11.3394 | FALSE | NA | 12.6326 | NA | NA | NA | NA | NA | NA
Q8IY31 | Intraflagellar transport protein 20 homolog | IFT20 | 1:1 | Block_3 | 0.69312 | 8.1326 | TRUE | NA | NA | NA | NA | 1.6401 | 0.95351 | 1.8127 | 3.0915
Q8IVU1 | Immunoglobulin superfamily DCC subclass member 3 | IGDCC3 | 1:1 | Block_3 | 0.049237 | 7.3802 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O00425 | Insulin-like growth factor 2 mRNA-binding protein 3 | IGF2BP3 | 1:1 | Block_3 | 1.4091 | 9.3762 | FALSE | NA | NA | NA | NA | 2.5772 | NA | 1.705 | 4.039
P38935 | DNA-binding protein SMUBP-2 | IGHMBP2 | 1:1 | Block_3 | 1.7077 | 10.8865 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6NGN9 | IgLON family member 5 | IGLON5 | 1:1 | Block_3 | 0.37946 | 11.4735 | FALSE | 18.598 | 3.308 | NA | NA | 0.40669 | NA | NA | 1.7045
Q9UKS7 | Zinc finger protein Helios | IKZF2 | 1:1 | Block_3 | 1.1625 | 10.1424 | TRUE | NA | NA | NA | NA | 1.1859 | NA | NA | NA
P22301 | Interleukin-10 | IL10 | 1:1 | Block_3 | -0.85376 | 12.6517 | FALSE | 19.5134 | 13.5304 | -0.77503 | -0.010788 | 0.76702 | -0.16813 | 1.0418 | 3.2116
Q13651 | Interleukin-10 receptor subunit alpha | IL10RA | 1:1 | Block_3 | 0.14626 | 12.8934 | FALSE | 3.696 | 5.6892 | NA | NA | 2.0011 | NA | 0.1603 | 1.3573
P20809 | Interleukin-11 | IL11 | 1:1 | Block_3 | 1.0109 | 12.2594 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P42701 | Interleukin-12 receptor subunit beta-1 | IL12RB1 | 1:1 | Block_3 | -1.15 | 10.0715 | TRUE | 35.8616 | 22.7295 | -0.26566 | 0.66431 | 1.3115 | 0.15162 | 1.2183 | 2.5544
Q99665 | Interleukin-12 receptor subunit beta-2 | IL12RB2 | 1:1 | Block_3 | 1.5252 | 8.7965 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P35225 | Interleukin-13 | IL13 | 1:1 | Block_3 | 1.78 | 15.0702 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q14627 | Interleukin-13 receptor subunit alpha-2 | IL13RA2 | 1:1 | Block_3 | 1.08 | 13.1258 | TRUE | 5.3263 | 28.4477 | NA | NA | 1.2345 | NA | NA | 1.3866
Q13261 | Interleukin-15 receptor subunit alpha | IL15RA | 1:1 | Block_3 | 0.45763 | 9.8526 | TRUE | 0.6394 | 1.9575 | NA | NA | 0.98553 | NA | 0.92156 | 2.0699
Q16552 | Interleukin-17A | IL17A | 1:1 | Block_3 | 0.97 | 12.3218 | FALSE | NA | 12.9792 | NA | NA | 1.4833 | NA | NA | 2.7971
Q9P0M4 | Interleukin-17C | IL17C | 1:1 | Block_3 | -2.186 | 9.7875 | FALSE | 8.4772 | 8.279 | -0.98698 | -0.14824 | 0.50196 | -0.90373 | 0.55864 | 2.2361
Q8TAD2 | Interleukin-17D | IL17D | 1:1 | Block_3 | -0.39368 | 10.0932 | TRUE | 8.8062 | 7.694 | -0.15543 | 0.4651 | 1.0385 | 0.072848 | 0.59746 | 1.24
Q96PD4 | Interleukin-17F | IL17F | 1:1 | Block_3 | 0.154 | 9.9587 | TRUE | NA | 3.9622 | NA | NA | 1.1519 | NA | NA | 1.4249
P01583 | Interleukin-1 alpha | IL1A | 1:1 | Block_3 | 1.3733 | 9.8949 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5301
P01584 | Interleukin-1 beta | IL1B | 1:1 | Block_3 | 2.6951 | 11.5379 | TRUE | NA | NA | NA | NA | NA | NA | NA | 4.6593
Q9HB29 | Interleukin-1 receptor-like 2 | IL1RL2 | 1:1 | Block_3 | -1.7928 | 8.0169 | TRUE | 12.8504 | 11.9381 | -0.69302 | 0.45477 | 1.1072 | -0.0020781 | 0.82571 | 1.5139
P60568 | Interleukin-2 | IL2 | 1:1 | Block_3 | 2.2623 | 15.0273 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NYY1 | Interleukin-20 | IL20 | 1:1 | Block_3 | 1.0891 | 10.1139 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UHF4 | Interleukin-20 receptor subunit alpha | IL20RA | 1:1 | Block_3 | 1.3321 | 12.212 | TRUE | NA | NA | NA | NA | 1.8634 | NA | NA | NA
Q6UXL0 | Interleukin-20 receptor subunit beta | IL20RB | 1:1 | Block_3 | 1.6342 | 7.8449 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9HBE5 | Interleukin-21 receptor | IL21R | 1:1 | Block_3 | 1.1645 | 8.3807 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N6P7 | Interleukin-22 receptor subunit alpha-1 | IL22RA1 | 1:1 | Block_3 | 0.012963 | 12.5602 | TRUE | 10.5098 | 7.9221 | NA | 0.097755 | 1.0634 | NA | 0.36317 | 1.3312
Q13007 | Interleukin-24 | IL24 | 1:1 | Block_3 | 0.47592 | 9.7532 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H293 | Interleukin-25 | IL25 | 1:1 | Block_3 | 1.1345 | 12.3968 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P14784 | Interleukin-2 receptor subunit beta | IL2RB | 1:1 | Block_3 | 1.48 | 12.8437 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P08700 | Interleukin-3 | IL3 | 1:1 | Block_3 | 4.5615 | 16.9774 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6EBC2 | Interleukin-31 | IL31 | 1:1 | Block_3 | 1.9715 | 12.2492 | FALSE | NA | NA | NA | NA | 2.4497 | NA | NA | 2.2809
Q8NI17 | Interleukin-31 receptor subunit alpha | IL31RA | 1:1 | Block_3 | 0.42 | 8.9075 | TRUE | 9.7637 | 5.1623 | NA | NA | 0.85785 | NA | 0.47343 | 1.1516
O95760 | Interleukin-33 | IL33 | 1:1 | Block_3 | 1.1255 | 12.8112 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZMJ4 | Interleukin-34 | IL34 | 1:1 | Block_3 | 0.18018 | NA | NA | 9.6317 | 9.5518 | NA | 0.29893 | 1.0285 | NA | 0.37576 | 1.1938
Q9UHA7 | Interleukin-36 alpha | IL36A | 1:1 | Block_3 | 1.3877 | 8.9153 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NZH7 | Interleukin-36 beta | IL36B | 1:1 | Block_3 | 0.38212 | 9.1679 | TRUE | NA | NA | NA | NA | 0.87765 | NA | NA | 0.5255
P26951 | Interleukin-3 receptor subunit alpha | IL3RA | 1:1 | Block_3 | 0.73 | 6.743 | TRUE | 1.0497 | 10.6029 | NA | NA | 1.0073 | NA | NA | 1.0916
P05112 | Interleukin-4 | IL4 | 1:1 | Block_3 | 2.03 | 15.719 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P05113 | Interleukin-5 | IL5 | 1:1 | Block_3 | 1.2 | 13.0713 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P15248 | Interleukin-9 | IL9 | 1:1 | Block_3 | 1.1148 | 10.0349 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H0C8 | Integrin-linked kinase-associated serine/threonine phosphatase 2C | ILKAP | 1:1 | Block_3 | 2.45 | 10.1664 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q17R60 | Interphotoreceptor matrix proteoglycan 1 | IMPG1 | 1:1 | Block_3 | 0.74123 | 10.3039 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UK53 | Inhibitor of growth protein 1 | ING1 | 1:1 | Block_3 | 1.1106 | 10.5013 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.2609
O15327 | Inositol polyphosphate 4-phosphatase type II | INPP4B | 1:1 | Block_3 | 1.4717 | 12.0449 | TRUE | NA | NA | NA | 2.2485 | 5.5521 | 1.4778 | 5.3625 | 7.472
Q15735 | Phosphatidylinositol 4,5-bisphosphate 5-phosphatase A | INPP5J | 1:1 | Block_3 | 2.27 | 15.4141 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.3168
Q4KMZ1 | IQ domain-containing protein C | IQCC | 1:1 | Block_3 | 0.81018 | 11.9343 | TRUE | 10.1927 | 5.7162 | NA | NA | 1.1477 | NA | NA | 1.543
P51617 | Interleukin-1 receptor-associated kinase 1 | IRAK1 | 1:1 | Block_3 | 0.5209 | 8.6487 | FALSE | NA | 2.9109 | NA | 1.0213 | 2.6046 | 1.1005 | 2.4284 | 3.8282
Q96J02 | E3 ubiquitin-protein ligase Itchy homolog | ITCH | 1:1 | Block_3 | 0.53033 | 7.5354 | TRUE | NA | NA | NA | NA | 1.5629 | NA | 1.4183 | 2.5197
P20702 | Integrin alpha-X | ITGAX | 1:1 | Block_3 | 2.9367 | 9.0839 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O14713 | Integrin beta-1-binding protein 1 | ITGB1BP1 | 1:1 | Block_3 | 1.5968 | 13.2341 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.4373
P18564 | Integrin beta-6 | ITGB6 | 1:1 | Block_3 | -0.005333 | 5.6541 | FALSE | 20.2267 | 11.3304 | NA | 0.2571 | 0.79326 | NA | 0.31619 | 1.061
Q9BY32 | Inosine triphosphate pyrophosphatase | ITPA | 1:1 | Block_3 | 3.0359 | 7.8726 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.0363
Q8IWB1 | Inositol 1,4,5-trisphosphate receptor-interacting protein | ITPRIP | 1:1 | Block_3 | 1.54 | 12.0634 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P26440 | Isovaleryl-CoA dehydrogenase, mitochondrial | IVD | 1:1 | Block_3 | 1.5184 | 6.9277 | FALSE | NA | NA | NA | NA | 1.687 | NA | NA | 2.5493
Q96N16 | Janus kinase and microtubule-interacting protein 1 | JAKMIP1 | 1:1 | Block_3 | 0.74777 | 13.0003 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5406
P01591 | Immunoglobulin J chain | JCHAIN | 1:1 | Block_3 | 1.3128 | 5.8588 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q15652 | Probable JmjC domain-containing histone demethylation protein 2C | JMJD1C | 1:1 | Block_3 | 1.7974 | 9.3161 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O76095 | Protein JTB | JTB | 1:1 | Block_3 | -2.18 | 8.6228 | TRUE | 16.6073 | 10.912 | -0.39246 | 0.36141 | 0.87444 | 0.24287 | 1.0099 | 2.2109
P05412 | Transcription factor Jun | JUN | 1:1 | Block_3 | 0.34987 | 10.8437 | FALSE | 16.7382 | 7.5433 | NA | NA | 0.53778 | NA | 0.8299 | 2.3925
Q63ZY3 | KN motif and ankyrin repeat domain-containing protein 2 | KANK2 | 1:1 | Block_3 | -0.56392 | 9.4231 | TRUE | 16.2074 | 9.0825 | -0.1396 | 0.75577 | 1.716 | 0.7984 | 1.6509 | 2.6691
Q92794 | Histone acetyltransferase KAT6A | KAT6A | 1:1 | Block_3 | 0.82714 | 12.6416 | TRUE | NA | NA | NA | NA | 1.4679 | NA | NA | 1.8549
Q9BVA0 | Katanin p80 WD40 repeat-containing subunit B1 | KATNB1 | 1:1 | Block_3 | 0.51799 | 12.0602 | FALSE | NA | NA | NA | NA | 2.5579 | NA | 2.3437 | 3.7085
Q03721 | Potassium voltage-gated channel subfamily C member 4 | KCNC4 | 1:1 | Block_3 | 0.32515 | 13.1856 | TRUE | NA | NA | NA | NA | 0.38505 | NA | NA | NA
Q6PIL6 | Kv channel-interacting protein 4 | KCNIP4 | 1:1 | Block_3 | 1.1112 | 6.7995 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6PI47 | BTB/POZ domain-containing protein KCTD18 | KCTD18 | 1:1 | Block_3 | 1.2657 | 10.6009 | TRUE | NA | NA | NA | NA | 1.76 | NA | NA | 2.0558
Q9Y4C1 | Lysine-specific demethylase 3A | KDM3A | 1:1 | Block_3 | 1.571 | 8.0663 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q587J8 | KH domain-containing protein 3 | KHDC3L | 1:1 | Block_3 | 0.001 | 9.5211 | TRUE | NA | NA | NA | NA | 1.217 | NA | NA | 0.50725
Q8IZA0 | Dyslexia-associated protein KIAA0319-like protein | KIAA0319L | 1:1 | Block_3 | -0.48753 | 10.7984 | FALSE | 17.7514 | 17.162 | 0.059969 | 0.48368 | 1.9068 | -0.053255 | 0.59647 | 1.3315
O60268 | Uncharacterized protein KIAA0513 | KIAA0513 | 1:1 | Block_3 | -0.073929 | 7.9067 | FALSE | 11.5132 | 15.6202 | NA | 1.7157 | 6.8718 | 1.6039 | 6.8147 | 8.1691
Q96AT1 | Uncharacterized protein KIAA1143 | KIAA1143 | 1:1 | Block_3 | -0.2351 | 10.1628 | TRUE | 10.345 | 6.1924 | NA | 0.27548 | 2.0884 | 0.48393 | 2.2392 | 4.988
Q86T90 | Protein hinderin | KIAA1328 | 1:1 | Block_3 | 1.4182 | 13.565 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7853
Q9HCM3 | UPF0606 protein KIAA1549 | KIAA1549 | 1:1 | Block_3 | 1.3852 | 9.8168 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZVL6 | UPF0606 protein KIAA1549L | KIAA1549L | 1:1 | Block_3 | 1.8953 | 13.2954 | TRUE | NA | NA | NA | NA | 2.1937 | NA | NA | NA
O43896 | Kinesin-like protein KIF1C | KIF1C | 1:1 | Block_3 | 1 | 9.7046 | TRUE | NA | 20.1008 | NA | NA | 1.2087 | NA | 1.5195 | 3.1631
Q96Q89 | Kinesin-like protein KIF20B | KIF20B | 1:1 | Block_3 | 1.7879 | 11.3011 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86VH2 | Kinesin-like protein KIF27 | KIF27 | 1:1 | Block_3 | 0.85889 | 10.6071 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y496 | Kinesin-like protein KIF3A | KIF3A | 1:1 | Block_3 | 1.6581 | 11.0499 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O95239 | Chromosome-associated kinesin KIF4A | KIF4A | 1:1 | Block_3 | 0.17224 | 10.1787 | TRUE | NA | NA | NA | NA | 1.082 | NA | 0.37833 | 2.4814
Q92845 | Kinesin-associated protein 3 | KIFAP3 | 1:1 | Block_3 | -0.031245 | 11.441 | TRUE | 9.9864 | 10.6834 | NA | NA | 0.42698 | NA | NA | 0.57462
Q9BW19 | Kinesin-like protein KIFC1 | KIFC1 | 1:1 | Block_3 | -0.62227 | 14.2884 | TRUE | 15.7465 | 20.9444 | NA | -0.57967 | 0.70694 | -0.13846 | 0.99299 | 3.3522
P43629 | Killer cell immunoglobulin-like receptor 3DL1 | KIR3DL1 | 1:1 | Block_3 | -1.4093 | 10.5539 | TRUE | 14.8053 | 13.0626 | NA | -0.60898 | 0.80365 | NA | 0.15144 | 1.7285
P43630 | Killer cell immunoglobulin-like receptor 3DL2 | KIR3DL2 | 1:1 | Block_3 | 0.39798 | 9.079 | TRUE | 15.4507 | 17.8457 | NA | NA | 0.84065 | NA | NA | NA
Q96J84 | Kin of IRRE-like protein 1 | KIRREL1 | 1:1 | Block_3 | -0.084349 | 7.2206 | FALSE | 7.1282 | 19.2775 | NA | NA | 0.23109 | NA | NA | 1.2003
Q6UWL6 | Kin of IRRE-like protein 2 | KIRREL2 | 1:1 | Block_3 | -0.10502 | 7.9976 | TRUE | 10.5559 | 18.8308 | NA | 0.27224 | 0.78332 | NA | 0.4073 | 1.2589
O43474 | Krueppel-like factor 4 | KLF4 | 1:1 | Block_3 | 0.8003 | 14.0614 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.3198
O60662 | Kelch-like protein 41 | KLHL41 | 1:1 | Block_3 | 0.9805 | 9.5023 | FALSE | NA | NA | NA | NA | NA | NA | NA | 0.99417
Q9UKR0 | Kallikrein-12 | KLK12 | 1:1 | Block_3 | -2 | 7.3509 | TRUE | 15.9438 | 13.2072 | NA | -0.81525 | 0.71292 | NA | -0.15439 | 1.5238
Q9UKQ9 | Kallikrein-9 | KLK9 | 1:1 | Block_3 | 0.30523 | 7.4365 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P26715 | NKG2-A/NKG2-B type II integral membrane protein | KLRC1 | 1:1 | Block_3 | 1.241 | 10.7442 | TRUE | NA | NA | NA | NA | 1.9571 | NA | NA | NA
Q9NZS2 | Killer cell lectin-like receptor subfamily F member 1 | KLRF1 | 1:1 | Block_3 | -0.75 | 9.5795 | TRUE | 15.3155 | 23.4674 | -0.32931 | 0.51834 | 1.3835 | -0.05712 | 0.92199 | 2.115
P02533 | Keratin, type I cytoskeletal 14 | KRT14 | 1:1 | Block_3 | 1.3217 | 9.2674 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q04695 | Keratin, type I cytoskeletal 17 | KRT17 | 1:1 | Block_3 | 2.53 | 8.1825 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P13647 | Keratin, type II cytoskeletal 5 | KRT5 | 1:1 | Block_3 | -1.68 | 7.8716 | FALSE | 12.188 | 15.7662 | -1.6139 | 0.11396 | 1.1599 | -0.96497 | 0.37793 | 2.3376
P05787 | Keratin, type II cytoskeletal 8 | KRT8 | 1:1 | Block_3 | 1.2141 | 9.6683 | FALSE | 9.5474 | 9.897 | NA | NA | NA | NA | NA | 2.4243
Q9UN81 | LINE-1 retrotransposable element ORF1 protein | L1RE1 | 1:1 | Block_3 | 1.5079 | 13.6103 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.9767
P25391 | Laminin subunit alpha-1 | LAMA1 | 1:1 | Block_3 | -0.27678 | 7.3921 | TRUE | NA | NA | NA | NA | 0.36064 | NA | NA | NA
P28838 | Cytosol aminopeptidase | LAP3 | 1:1 | Block_3 | 0.45535 | 5.0523 | FALSE | NA | 1.9439 | NA | NA | 0.71916 | NA | 0.96124 | 2.6985
Q8N3X6 | Ligand-dependent nuclear receptor corepressor-like protein | LCORL | 1:1 | Block_3 | 0.76927 | 8.1662 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZMV7 | Protein LEKR1 | LEKR1 | 1:1 | Block_3 | 0.67162 | 12.4479 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.80893
Q5T871 | Late cornified envelope-like proline-rich protein 1 | LELP1 | 1:1 | Block_3 | 0.86 | 14.6218 | TRUE | NA | 3.6263 | NA | NA | 0.99009 | NA | NA | 1.195
Q8WVC0 | RNA polymerase-associated protein LEO1 | LEO1 | 1:1 | Block_3 | 0.11464 | 8.8757 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.796
P01229 | Lutropin subunit beta | LHB | 1:1 | Block_3 | -2.5 | 6.3086 | FALSE | 25.8361 | 12.2096 | NA | NA | -0.60532 | NA | NA | -0.39179
P15018 | Leukemia inhibitory factor | LIF | 1:1 | Block_3 | 0.8395 | 11.1366 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P49916 | DNA ligase 3 | LIG3 | 1:1 | Block_3 | 1.421 | 7.1704 | FALSE | NA | NA | NA | NA | 1.7392 | NA | 1.9525 | 3.1162
P59901 | Leukocyte immunoglobulin-like receptor subfamily A member 4 | LILRA4 | 1:1 | Block_3 | 0.39072 | 5.1068 | FALSE | 12.8574 | 22.8329 | NA | NA | 1.1426 | NA | NA | 1.2067
Q8NHJ6 | Leukocyte immunoglobulin-like receptor subfamily B member 4 | LILRB4 | 1:1 | Block_3 | -0.71371 | 8.9968 | TRUE | 18.0949 | 14.7254 | 0.070466 | 1.4029 | 2.5207 | 0.90665 | 2.0619 | 3.73
Q6P5Q4 | Leiomodin-2 | LMOD2 | 1:1 | Block_3 | 2.2532 | 9.2848 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O60711 | Leupaxin | LPXN | 1:1 | Block_3 | -0.93598 | 6.7661 | TRUE | 9.0434 | 9.8771 | NA | NA | -0.68908 | NA | -0.81291 | -0.011475
P50851 | Lipopolysaccharide-responsive and beige-like anchor protein | LRBA | 1:1 | Block_3 | -0.12923 | 9.2341 | TRUE | 2.3551 | 13.2348 | NA | 0.56599 | 4.331 | 1.1574 | 4.2723 | 5.7919
Q9ULH4 | Leucine-rich repeat and fibronectin type-III domain-containing protein 2 | LRFN2 | 1:1 | Block_3 | -0.56 | 8.6449 | TRUE | 17.8606 | 9.2998 | NA | NA | 0.68965 | NA | NA | 0.90755
Q6UXM1 | Leucine-rich repeats and immunoglobulin-like domains protein 3 | LRIG3 | 1:1 | Block_3 | 1.6741 | 8.8771 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P98164 | Low-density lipoprotein receptor-related protein 2 | LRP2 | 1:1 | Block_3 | 0.79838 | 8.0091 | TRUE | 15.2287 | 0.59941 | NA | NA | 0.93742 | NA | NA | 0.81732
Q9P2M1 | LRP2-binding protein | LRP2BP | 1:1 | Block_3 | 1 | NA | NA | NA | NA | NA | NA | NA | NA | NA | 1.25
Q5VT99 | Leucine-rich repeat-containing protein 38 | LRRC38 | 1:1 | Block_3 | -0.00062245 | 10.5763 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96CN5 | Leucine-rich repeat-containing protein 45 | LRRC45 | 1:1 | Block_3 | 0.96622 | 9.4514 | TRUE | NA | NA | NA | NA | 1.375 | NA | 1.0105 | 2.1181
Q32MZ4 | Leucine-rich repeat flightless-interacting protein 1 | LRRFIP1 | 1:1 | Block_3 | 0.99606 | 6.9726 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.137
Q6UXK5 | Leucine-rich repeat neuronal protein 1 | LRRN1 | 1:1 | Block_3 | -3.6523 | 6.5678 | TRUE | 9.3653 | 12.6423 | -0.57882 | 0.25401 | 0.85704 | -0.70464 | 0.22303 | 0.98339
Q9HBL6 | Leucine-rich repeat and transmembrane domain-containing protein 1 | LRTM1 | 1:1 | Block_3 | 1.7464 | 11.6666 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H089 | Large subunit GTPase 1 homolog | LSG1 | 1:1 | Block_3 | 0.6955 | 11.2759 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O15116 | U6 snRNA-associated Sm-like protein LSm1 | LSM1 | 1:1 | Block_3 | 0.69452 | 10.5216 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.71945
O95777 | U6 snRNA-associated Sm-like protein LSm8 | LSM8 | 1:1 | Block_3 | 0.78622 | 4.7838 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NS15 | Latent-transforming growth factor beta-binding protein 3 | LTBP3 | 1:1 | Block_3 | -1.4044 | 8.6951 | TRUE | 10.7497 | 14.6461 | -0.54916 | 0.38008 | 1.1802 | -0.073153 | 0.83817 | 2.1419
Q8WV07 | Protein LTO1 homolog | LTO1 | 1:1 | Block_3 | 1.2829 | 10.9599 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NQ29 | Putative RNA-binding protein Luc7-like 1 | LUC7L | 1:1 | Block_3 | -0.22298 | 9.4623 | TRUE | 9.3344 | 7.3916 | NA | 0.10356 | 0.51994 | 0.27695 | 1.0264 | 2.7697
O95232 | Luc7-like protein 3 | LUC7L3 | 1:1 | Block_3 | -0.2164 | 8.3257 | TRUE | 12.6888 | 13.8047 | NA | -0.14388 | 0.47084 | -0.040316 | 0.8674 | 2.3257
Q9NX58 | Cell growth-regulating nucleolar protein | LYAR | 1:1 | Block_3 | 1.6167 | 9.4011 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.2771
Q8N2G4 | Ly6/PLAUR domain-containing protein 1 | LYPD1 | 1:1 | Block_3 | 0.73138 | 8.449 | TRUE | NA | NA | NA | NA | 1.0022 | NA | NA | NA
Q6UX82 | Ly6/PLAUR domain-containing protein 8 | LYPD8 | 1:1 | Block_3 | -2 | 10.0387 | TRUE | 24.97 | 36.2024 | -1.3897 | -0.085641 | 1.1675 | -0.88979 | 0.36198 | 1.4607
O95372 | Acyl-protein thioesterase 2 | LYPLA2 | 1:1 | Block_3 | -0.038765 | 5.1303 | FALSE | 6.0818 | 11.8119 | NA | NA | 1.9974 | NA | NA | 1.717
Q7Z3D4 | LysM and putative peptidoglycan-binding domain-containing protein 3 | LYSMD3 | 1:1 | Block_3 | 0.15964 | 9.8162 | TRUE | 9.6713 | 33.3873 | NA | NA | 1.6958 | NA | 1.0183 | 2.8514
Q7Z4W2 | Lysozyme-like protein 2 | LYZL2 | 1:1 | Block_3 | 0.025735 | 10.0672 | TRUE | 10.2321 | 10.6643 | NA | 0.079879 | 0.45198 | NA | 0.27641 | 0.987
Q9UPN3 | Microtubule-actin cross-linking factor 1, isoforms 1/2/3/4/5 | MACF1 | 1:1 | Block_3 | -2.8606 | 6.3115 | TRUE | 8.2596 | 4.45 | -2.7805 | -1.2218 | 2.0727 | -2.2603 | -0.34808 | 1.4191
Q7L5Y9 | E3 ubiquitin-protein transferase MAEA | MAEA | 1:1 | Block_3 | 0.95 | 11.8287 | TRUE | NA | NA | NA | NA | 1.7702 | NA | 1.2333 | 2.878
P43357_P43360 | Melanoma-associated antigen 3/6 | MAGEA3_MAGEA6 | 1:1 | Block_3 | 1.9117 | 12.7343 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O15479 | Melanoma-associated antigen B2 | MAGEB2 | 1:1 | Block_3 | 0.46204 | 9.6801 | TRUE | NA | NA | NA | NA | 0.9184 | NA | 0.77884 | 3.264
Q9UBF1 | Melanoma-associated antigen C2 | MAGEC2 | 1:1 | Block_3 | 2.509 | 10.9547 | FALSE | NA | NA | NA | NA | 2.9796 | NA | NA | 2.7632
Q9Y5V3 | Melanoma-associated antigen D1 | MAGED1 | 1:1 | Block_3 | 1.5571 | 14.1757 | TRUE | NA | NA | NA | NA | 2.3302 | NA | 2.0905 | 3.5301
Q6UXC1 | Apical endosomal glycoprotein | MAMDC4 | 1:1 | Block_3 | 0.001 | 8.2619 | TRUE | 9.2886 | 11.6087 | NA | 0.43627 | 1.5686 | NA | NA | 1.587
Q5VSG8 | Glycoprotein endo-alpha-1,2-mannosidase-like protein | MANEAL | 1:1 | Block_3 | 1.2653 | 6.9763 | FALSE | 10.9666 | 37.9216 | NA | NA | NA | NA | NA | NA
A6NHS7 | MANSC domain-containing protein 4 | MANSC4 | 1:1 | Block_3 | 0.8 | 11.6362 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.0099
Q9H492 | Microtubule-associated proteins 1A/1B light chain 3A | MAP1LC3A | 1:1 | Block_3 | 3.6792 | 11.2979 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6NCE7 | Microtubule-associated proteins 1A/1B light chain 3 beta 2 | MAP1LC3B2 | 1:1 | Block_3 | 3.1304 | 10.5672 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q02750 | Dual specificity mitogen-activated protein kinase kinase 1 | MAP2K1 | 1:1 | Block_3 | 1.4108 | 9.1172 | FALSE | NA | NA | NA | NA | 2.8662 | 1.7264 | 3.961 | 6.3434
Q56UN5 | Mitogen-activated protein kinase kinase kinase 19 | MAP3K19 | 1:1 | Block_3 | 2 | 14.4186 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q99683 | Mitogen-activated protein kinase kinase kinase 5 | MAP3K5 | 1:1 | Block_3 | 1.9052 | 4.8297 | FALSE | NA | NA | NA | NA | 3.4972 | NA | 2.0978 | 5.8669
Q49MG5 | Microtubule-associated protein 9 | MAP9 | 1:1 | Block_3 | 0.10979 | 14.1813 | TRUE | NA | 7.4655 | NA | NA | 1.7973 | NA | 1.0836 | 3.0935
O15264 | Mitogen-activated protein kinase 13 | MAPK13 | 1:1 | Block_3 | 1.825 | 6.5357 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P49137 | MAP kinase-activated protein kinase 2 | MAPKAPK2 | 1:1 | Block_3 | 0.91744 | 11.6694 | FALSE | NA | NA | NA | 0.96892 | 3.011 | NA | 3.1155 | 4.316
Q9UPY8 | Microtubule-associated protein RP/EB family member 3 | MAPRE3 | 1:1 | Block_3 | 2.0136 | 9.5165 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P10636 | Microtubule-associated protein tau | MAPT | 1:1 | Block_3 | 4.19 | 13.6561 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96A59 | MARVEL domain-containing protein 3 | MARVELD3 | 1:1 | Block_3 | 0.40633 | 11.9963 | TRUE | 14.9563 | 8.5059 | NA | 0.55621 | 1.9854 | NA | 0.89169 | 2.2965
Q7Z434 | Mitochondrial antiviral-signaling protein | MAVS | 1:1 | Block_3 | -2.6382 | 6.5986 | FALSE | 8.407 | 9.281 | -0.97543 | 0.33198 | 3.5628 | 0.91871 | 2.8496 | 5.1543
P61244 | Protein max | MAX | 1:1 | Block_3 | 1.4842 | 7.4162 | TRUE | NA | NA | NA | NA | 3.6471 | NA | 3.4044 | 5.2038
Q7L590 | Protein MCM10 homolog | MCM10 | 1:1 | Block_3 | 1.461 | 10.4379 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TC05 | Nuclear protein MDM1 | MDM1 | 1:1 | Block_3 | 1.55 | 11.1718 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7471
Q13503 | Mediator of RNA polymerase II transcription subunit 21 | MED21 | 1:1 | Block_3 | 1.42 | 10.4849 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7381
Q96HR3 | Mediator of RNA polymerase II transcription subunit 30 | MED30 | 1:1 | Block_3 | 1.7383 | 6.2581 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6BM72 | Multiple epidermal growth factor-like domains protein 11 | MEGF11 | 1:1 | Block_3 | 1.2145 | 10.0953 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.2388
Q99687 | Homeobox protein Meis3 | MEIS3 | 1:1 | Block_3 | 2.84 | 11.4486 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P50221 | Homeobox protein MOX-1 | MEOX1 | 1:1 | Block_3 | 0.83754 | 12.0331 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.87102
P53582 | Methionine aminopeptidase 1 | METAP1 | 1:1 | Block_3 | 1.8753 | 12.1346 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5U5Q3 | RNA-binding E3 ubiquitin-protein ligase MEX3C | MEX3C | 1:1 | Block_3 | 1.5442 | 10.2918 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O75121 | Microfibrillar-associated protein 3-like | MFAP3L | 1:1 | Block_3 | 0.037706 | 8.6498 | FALSE | NA | NA | NA | NA | 0.70487 | NA | NA | 0.31298
Q96AX9 | E3 ubiquitin-protein ligase MIB2 | MIB2 | 1:1 | Block_3 | 0.80887 | 8.6579 | FALSE | NA | 6.9696 | NA | NA | 1.3232 | NA | NA | 0.89216
Q8IY33 | MICAL-like protein 2 | MICALL2 | 1:1 | Block_3 | 0.17018 | 10.5537 | TRUE | NA | 2.552 | NA | NA | 2.1796 | NA | 1.6115 | 3.7516
Q7Z6M3 | Allergin-1 | MILR1 | 1:1 | Block_3 | -1.33 | 10.5349 | TRUE | 15.5318 | 15.885 | -0.45948 | 0.386 | 1.631 | -0.31635 | 0.87539 | 2.1304
Q8NBR6 | Ubiquitin carboxyl-terminal hydrolase MINDY-2 | MINDY2 | 1:1 | Block_3 | -2.4812 | 8.9608 | FALSE | 7.0829 | 8.6205 | -0.17514 | 0.40306 | 1.0887 | 0.32991 | 1.2098 | 2.38
P46013 | Proliferation marker protein Ki-67 | MKI67 | 1:1 | Block_3 | 1.1212 | 11.3683 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.069
Q03111 | Protein ENL | MLLT1 | 1:1 | Block_3 | 1.2311 | 11.4393 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96EY8 | Corrinoid adenosyltransferase MMAB | MMAB | 1:1 | Block_3 | -0.46 | 9.3138 | TRUE | 7.2201 | NA | NA | 0.071266 | 2.1867 | NA | 0.66722 | 2.5242
P45452 | Collagenase 3 | MMP13 | 1:1 | Block_3 | -0.17026 | 10.9824 | FALSE | 2.0344 | 4.3381 | NA | NA | 0.54843 | NA | 0.14083 | 1.4137
P22033 | Methylmalonyl-CoA mutase, mitochondrial | MMUT | 1:1 | Block_3 | 2.5 | 11.1264 | TRUE | NA | NA | NA | NA | 2.7872 | NA | NA | 3.6192
Q10571 | Transcriptional activator MN1 | MN1 | 1:1 | Block_3 | 2.2245 | 14.5121 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.6223
P51948 | CDK-activating kinase assembly factor MAT1 | MNAT1 | 1:1 | Block_3 | 0.85779 | 8.8278 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UQ07 | MAPK/MAK/MRK overlapping kinase | MOK | 1:1 | Block_3 | -1.301 | 15.3303 | FALSE | 18.0047 | 5.3756 | NA | NA | -0.4509 | NA | NA | -0.55129
Q9UBU8 | Mortality factor 4-like protein 1 | MORF4L1 | 1:1 | Block_3 | 0.98915 | 7.9878 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.447
Q8NDC4 | MORN repeat-containing protein 4 | MORN4 | 1:1 | Block_3 | 1.7501 | 12.4781 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q99550 | M-phase phosphoprotein 9 | MPHOSPH9 | 1:1 | Block_3 | 1.7147 | 9.9661 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P25325 | 3-mercaptopyruvate sulfurtransferase | MPST | 1:1 | Block_3 | -2.5542 | 6.838 | FALSE | 12.0486 | 10.7337 | 0.23665 | 2.2417 | 3.9112 | 1.6173 | 3.8284 | 5.3357
Q8TCY5 | Melanocortin-2 receptor accessory protein | MRAP | 1:1 | Block_3 | 2.2084 | 13.1392 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BV20 | Methylthioribose-1-phosphate isomerase | MRI1 | 1:1 | Block_3 | 0.6 | 6.9152 | FALSE | 25.2093 | 13.6481 | NA | 0.95683 | 2.0447 | NA | 1.5142 | 3.0877
Q9H2W6 | 39S ribosomal protein L46, mitochondrial | MRPL46 | 1:1 | Block_3 | 0.92977 | 7.9833 | FALSE | NA | NA | NA | NA | 1.5808 | NA | NA | 2.1324
Q86TS9 | 39S ribosomal protein L52, mitochondrial | MRPL52 | 1:1 | Block_3 | 0.768 | 3.3319 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y3D3 | 28S ribosomal protein S16, mitochondrial | MRPS16 | 1:1 | Block_3 | 1.305 | 7.3603 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y676 | 28S ribosomal protein S18b, mitochondrial | MRPS18B | 1:1 | Block_3 | 1.4859 | 6.379 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96E11 | Ribosome-recycling factor, mitochondrial | MRRF | 1:1 | Block_3 | 0.15726 | 10.0071 | FALSE | NA | NA | NA | NA | 2.0821 | NA | 1.0285 | 2.4535
Q969V6 | Myocardin-related transcription factor A | MRTFA | 1:1 | Block_3 | 0.20886 | 10.3949 | FALSE | 10.5601 | 9.4153 | NA | 1.1521 | 2.9801 | 0.59477 | 3.0676 | 4.7362
Q96KJ4 | Mesothelin-like protein | MSLNL | 1:1 | Block_3 | 1.0523 | 7.555 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P11586 | C-1-tetrahydrofolate synthase, cytoplasmic | MTHFD1 | 1:1 | Block_3 | 0.80071 | 7.3755 | FALSE | 8.1715 | 7.5444 | NA | NA | 1.2356 | NA | 1.3248 | 2.2993
Q2M296 | Methenyltetrahydrofolate synthase domain-containing protein | MTHFSD | 1:1 | Block_3 | 1.8445 | 11.2229 | FALSE | NA | NA | NA | NA | 2.6299 | NA | 2.062 | 4.2497
Q99707 | Methionine synthase | MTR | 1:1 | Block_3 | 1.7473 | 8.5037 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9ULD2 | Microtubule-associated tumor suppressor 1 | MTUS1 | 1:1 | Block_3 | 1.341 | 12.8203 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4396
Q5JR59 | Microtubule-associated tumor suppressor candidate 2 | MTUS2 | 1:1 | Block_3 | 0.62288 | 12.9457 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H3R2 | Mucin-13 | MUC13 | 1:1 | Block_3 | 0.52906 | 8.2855 | TRUE | 4.1679 | 4.0467 | NA | NA | 0.78519 | NA | NA | 1.7412
Q8N307 | Mucin-20 | MUC20 | 1:1 | Block_3 | -0.21496 | 5.8002 | TRUE | 15.2352 | 10.2757 | NA | NA | 0.97575 | NA | NA | 1.5581
Q3MIW9 | Mucin-like protein 3 | MUCL3 | 1:1 | Block_3 | 0.9199 | 10.3308 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q03426 | Mevalonate kinase | MVK | 1:1 | Block_3 | 0.58082 | 6.9633 | FALSE | 14.3216 | 5.9426 | NA | NA | 1.3574 | NA | 1.1836 | 2.8731
Q9Y623 | Myosin-4 | MYH4 | 1:1 | Block_3 | 1.2232 | 12.9985 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A7E2Y1 | Myosin-7B | MYH7B | 1:1 | Block_3 | 1.1732 | 11.4391 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P05976 | Myosin light chain 1/3, skeletal muscle isoform | MYL1 | 1:1 | Block_3 | 0.87 | 6.9374 | TRUE | 18.8493 | 10.5863 | NA | NA | 1.3403 | NA | NA | 2.1128
Q96A32 | Myosin regulatory light chain 11 | MYL11 | 1:1 | Block_3 | 0.78422 | 6.1495 | FALSE | 3.45 | 2.4621 | NA | NA | 1.1737 | NA | 0.81805 | 3.1288
P08590 | Myosin light chain 3 | MYL3 | 1:1 | Block_3 | -2.5801 | 6.1319 | TRUE | 8.117 | 15.1946 | -2.0719 | -0.68585 | 1.5659 | -1.2709 | 0.77347 | 2.9899
P14649 | Myosin light chain 6B | MYL6B | 1:1 | Block_3 | 1.9083 | 10.3294 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.2544
Q9UM54 | Unconventional myosin-VI | MYO6 | 1:1 | Block_3 | 2.3196 | 15.3406 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q13459 | Unconventional myosin-IXb | MYO9B | 1:1 | Block_3 | 0.80555 | 7.8375 | FALSE | NA | NA | NA | NA | 1.0884 | NA | NA | 2.318
P52179 | Myomesin-1 | MYOM1 | 1:1 | Block_3 | 0.57097 | 11.3726 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P54296 | Myomesin-2 | MYOM2 | 1:1 | Block_3 | 1.584 | 8.6152 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7719
Q08AG7 | Mitotic-spindle organizing protein 1 | MZT1 | 1:1 | Block_3 | 0.18517 | 11.939 | FALSE | 5.9481 | 3.5315 | NA | 0.28627 | 2.5149 | 0.41523 | 1.549 | 3.256
Q6NZ67 | Mitotic-spindle organizing protein 2B | MZT2B | 1:1 | Block_3 | 0.84545 | 10.9733 | FALSE | NA | NA | NA | NA | 2.5649 | NA | 2.393 | 3.6957
O15049 | NEDD4-binding protein 3 | N4BP3 | 1:1 | Block_3 | 0.95148 | 12.9164 | TRUE | 12.146 | NA | NA | NA | NA | NA | NA | NA
Q86UY6 | N-alpha-acetyltransferase 40 | NAA40 | 1:1 | Block_3 | 0.84853 | 6.6639 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q93015 | N-alpha-acetyltransferase 80 | NAA80 | 1:1 | Block_3 | 1.5655 | 10.4698 | FALSE | NA | NA | NA | NA | 2.8422 | NA | 3.4243 | 5.5014
O15069 | NAC-alpha domain-containing protein 1 | NACAD | 1:1 | Block_3 | 1.6096 | 13.9588 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8226
Q96HR8 | H/ACA ribonucleoprotein complex non-core subunit NAF1 | NAF1 | 1:1 | Block_3 | 1.4517 | 11.8467 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P43490 | Nicotinamide phosphoribosyltransferase | NAMPT | 1:1 | Block_3 | -0.4191 | 4.2378 | FALSE | 3.0633 | 3.6662 | NA | NA | 0.2509 | NA | 0.20914 | 3.6551
O60934 | Nibrin | NBN | 1:1 | Block_3 | 0.24772 | 12.1752 | FALSE | NA | NA | NA | NA | NA | NA | 0.28891 | 2.0311
Q9BPX3 | Condensin complex subunit 3 | NCAPG | 1:1 | Block_3 | 0.21833 | 11.212 | TRUE | 17.6489 | 9.5876 | NA | 0.64932 | 1.4554 | 0.30606 | 1.3207 | 3.0242
Q6ZVX7 | F-box only protein 50 | NCCRP1 | 1:1 | Block_3 | -0.44009 | 8.6715 | TRUE | 15.9941 | 9.448 | -0.38027 | 0.23979 | 0.90502 | -0.43347 | 0.4885 | 1.5727
Q9NZQ3 | NCK-interacting protein with SH3 domain | NCKIPSD | 1:1 | Block_3 | 0.53981 | 5.6089 | FALSE | NA | 8.8659 | NA | NA | 2.762 | NA | 0.8756 | 1.985
Q969V3 | Nicalin | NCLN | 1:1 | Block_3 | 2.116 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Q8N183 | NADH dehydrogenase [ubiquinone] 1 alpha subcomplex assembly factor 2 | NDUFAF2 | 1:1 | Block_3 | 0.40163 | 11.2503 | TRUE | 1.2859 | NA | NA | 1.2571 | 4.9973 | 1.7184 | 3.6746 | 5.5431
O96000 | NADH dehydrogenase [ubiquinone] 1 beta subcomplex subunit 10 | NDUFB10 | 1:1 | Block_3 | 0.74268 | 10.704 | TRUE | NA | 25.548 | NA | NA | 0.95749 | NA | NA | 1.8568
O75380 | NADH dehydrogenase [ubiquinone] iron-sulfur protein 6, mitochondrial | NDUFS6 | 1:1 | Block_3 | 0.48697 | 11.3215 | TRUE | 10.3955 | 18.5925 | NA | 1.9463 | 5.9068 | 2.419 | 4.2243 | 6.2507
P20929 | Nebulin | NEB | 1:1 | Block_3 | 0.69532 | 12.0388 | FALSE | 12.8992 | 15.0213 | NA | NA | 0.92017 | NA | NA | 1.5473
Q8N987 | N-terminal EF-hand calcium-binding protein 1 | NECAB1 | 1:1 | Block_3 | 1.0922 | 10.9902 | TRUE | NA | NA | NA | NA | 1.1924 | NA | NA | 1.9731
Q7Z6G3 | N-terminal EF-hand calcium-binding protein 2 | NECAB2 | 1:1 | Block_3 | -1.1716 | 11.0069 | TRUE | 12.0164 | 11.5564 | -0.58424 | 0.091895 | 0.67477 | -0.083753 | 0.59568 | 1.5921
Q9NVZ3 | Adaptin ear-binding coat-associated protein 2 | NECAP2 | 1:1 | Block_3 | 1.0702 | 6.5744 | FALSE | NA | NA | NA | NA | 2.3211 | NA | 2.5967 | 3.9857
Q8NHV4 | Protein NEDD1 | NEDD1 | 1:1 | Block_3 | -0.78614 | 10.5622 | TRUE | 8.5358 | 12.9209 | -0.76517 | 0.38832 | 2.2172 | 0.18482 | 1.7838 | 3.5121
P07196 | Neurofilament light polypeptide | NEFL | 1:1 | Block_3 | 1.5 | 11.5106 | FALSE | NA | NA | NA | NA | 1.5557 | NA | NA | 3.1121
P51956 | Serine/threonine-protein kinase Nek3 | NEK3 | 1:1 | Block_3 | -0.71579 | 13.0218 | TRUE | 8.6877 | 17.626 | NA | -0.70518 | 2.0256 | NA | 1.0197 | 3.2358
Q8TDX7 | Serine/threonine-protein kinase Nek7 | NEK7 | 1:1 | Block_3 | 0.62639 | 9.9415 | FALSE | NA | 23.8179 | NA | NA | 2.1255 | NA | 1.3682 | 3.2658
Q9UMX5 | Neudesin | NENF | 1:1 | Block_3 | 1.1014 | 8.9193 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4318
Q12968 | Nuclear factor of activated T-cells, cytoplasmic 3 | NFATC3 | 1:1 | Block_3 | 1 | 10.4196 | FALSE | NA | NA | NA | NA | 1.0318 | NA | NA | 2.3585
P08651 | Nuclear factor 1 C-type | NFIC | 1:1 | Block_3 | 0.9824 | 5.4287 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P19838 | Nuclear factor NF-kappa-B p105 subunit | NFKB1 | 1:1 | Block_3 | -0.28668 | 3.701 | FALSE | 6.5879 | 12.3259 | NA | 0.041722 | 1.9279 | 0.11019 | 2.0419 | 3.4092
Q00653 | Nuclear factor NF-kappa-B p100 subunit | NFKB2 | 1:1 | Block_3 | 1.4642 | 11.1141 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O00221 | NF-kappa-B inhibitor epsilon | NFKBIE | 1:1 | Block_3 | 1.3 | 10.4873 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.3427
Q12986 | Transcriptional repressor NF-X1 | NFX1 | 1:1 | Block_3 | -1.7146 | 10.7672 | TRUE | NA | NA | NA | NA | 0.13049 | NA | -0.67016 | 0.88534
P23511 | Nuclear transcription factor Y subunit alpha | NFYA | 1:1 | Block_3 | -0.27 | 4.0908 | FALSE | 17.2217 | 20.8068 | NA | NA | 0.12227 | NA | -0.0050004 | 2.0399
Q9NX24 | H/ACA ribonucleoprotein complex subunit 2 | NHP2 | 1:1 | Block_3 | 0.40824 | 8.7928 | FALSE | 4.4442 | 3.6427 | NA | NA | NA | NA | NA | 1.4199
Q92982 | Ninjurin-1 | NINJ1 | 1:1 | Block_3 | 1.6736 | 11.338 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NYS0 | NF-kappa-B inhibitor-interacting Ras-like protein 1 | NKIRAS1 | 1:1 | Block_3 | 1.9388 | 10.6171 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N2Q7 | Neuroligin-1 | NLGN1 | 1:1 | Block_3 | 1.9682 | 11.6145 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.1585
Q8NFZ4 | Neuroligin-2 | NLGN2 | 1:1 | Block_3 | -1.6561 | 8.2766 | FALSE | NA | NA | NA | NA | -1.6373 | NA | NA | -1.1038
Q7RTR0 | NACHT, LRR and PYD domains-containing protein 9 | NLRP9 | 1:1 | Block_3 | -0.08704 | 7.1835 | FALSE | NA | NA | NA | NA | -0.065314 | NA | NA | NA
P15531 | Nucleoside diphosphate kinase A | NME1 | 1:1 | Block_3 | 1.2381 | 6.6905 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q13253 | Noggin | NOG | 1:1 | Block_3 | -1.65 | 9.6748 | TRUE | 27.5779 | 18.2414 | 0.047011 | 2.7044 | 4.2236 | 1.0756 | 3.5883 | 5.739
O00567 | Nucleolar protein 56 | NOP56 | 1:1 | Block_3 | 1.7081 | 11.76 | TRUE | 17.227 | 18.1994 | NA | NA | 1.8988 | NA | NA | 1.9553
P29475 | Nitric oxide synthase, brain | NOS1 | 1:1 | Block_3 | -0.08157 | 12.0709 | TRUE | 14.1387 | 13.1073 | NA | 0.051676 | 0.94236 | NA | 0.9352 | 2.5891
P35228 | Nitric oxide synthase, inducible | NOS2 | 1:1 | Block_3 | 1.1592 | 4.6074 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P29474 | Nitric oxide synthase, endothelial | NOS3 | 1:1 | Block_3 | 0.83102 | 12.6343 | TRUE | 8.1196 | 2.1409 | NA | 1.4038 | 1.9808 | NA | 1.5877 | 2.7398
Q9NP85 | Podocin | NPHS2 | 1:1 | Block_3 | 0.42798 | 11.879 | TRUE | NA | NA | NA | NA | 2.0991 | NA | NA | 0.81259
P16860 | Natriuretic peptides B | NPPB | 1:1 | Block_3 | 0.8034 | 12.7867 | TRUE | 0.73477 | 6.0029 | NA | 1.4983 | 3.0603 | NA | 1.8989 | 6.5031
P16066 | Atrial natriuretic peptide receptor 1 | NPR1 | 1:1 | Block_3 | 1.6683 | 4.9848 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y639 | Neuroplastin | NPTN | 1:1 | Block_3 | -0.48 | 5.4234 | FALSE | 29.0295 | 33.3148 | NA | -0.068702 | 1.2985 | NA | -0.062169 | 1.0872
P15559 | NAD(P)H dehydrogenase [quinone] 1 | NQO1 | 1:1 | Block_3 | 0.83043 | 10.3148 | FALSE | 6.5731 | 3.1824 | NA | NA | 1.7314 | NA | NA | 2.1025
Q9NPD7 | Neuritin | NRN1 | 1:1 | Block_3 | 1.014 | 6.1026 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q99748 | Neurturin | NRTN | 1:1 | Block_3 | 1.1906 | 11.5764 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P58400 | Neurexin-1-beta | NRXN1 | 1:1 | Block_3 | -1.184 | 9.4941 | FALSE | 7.7157 | 9.732 | NA | -0.87751 | -0.33943 | NA | -0.61593 | 0.26068
Q9Y4C0 | Neurexin-3 | NRXN3 | 1:1 | Block_3 | 1.1739 | 8.6113 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q15738 | Sterol-4-alpha-carboxylate 3-dehydrogenase, decarboxylating | NSDHL | 1:1 | Block_3 | 0.46902 | 6.4286 | FALSE | NA | NA | NA | NA | NA | NA | NA | 0.77296
Q96MF7 | E3 SUMO-protein ligase NSE2 | NSMCE2 | 1:1 | Block_3 | 1.8889 | 9.9814 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P34130 | Neurotrophin-4 | NTF4 | 1:1 | Block_3 | 0.87679 | 11.8045 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96CW9 | Netrin-G2 | NTNG2 | 1:1 | Block_3 | 1.3644 | 8.0589 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NV35 | Nucleotide triphosphate diphosphatase NUDT15 | NUDT15 | 1:1 | Block_3 | 1.2935 | 10.411 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7502
Q9UKX7 | Nuclear pore complex protein Nup50 | NUP50 | 1:1 | Block_3 | 0.7 | 5.8912 | FALSE | 27.7688 | 10.0539 | NA | NA | 0.79825 | NA | NA | 0.97828
Q86Y26 | NUT family member 1 | NUTM1 | 1:1 | Block_3 | 1.8877 | 10.5857 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6UWF7 | NXPE family member 4 | NXPE4 | 1:1 | Block_3 | 1.2087 | 7.5545 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P58417 | Neurexophilin-1 | NXPH1 | 1:1 | Block_3 | 0.31214 | 7.565 | TRUE | 6.8174 | NA | NA | NA | 0.49413 | NA | 0.31485 | 0.87032
Q5W0B1 | ORC ubiquitin ligase 1 | OBI1 | 1:1 | Block_3 | 0.39105 | 12.9572 | TRUE | NA | NA | NA | NA | 0.44672 | NA | NA | 1.8632
A1E959 | Odontogenic ameloblast-associated protein | ODAM | 1:1 | Block_3 | 0.93885 | 9.1039 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O75665 | Centriole and centriolar satellite protein OFD1 | OFD1 | 1:1 | Block_3 | 1.7632 | 12.9255 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8632
O15294 | UDP-N-acetylglucosamine--peptide N-acetylglucosaminyltransferase 110 kDa subunit | OGT | 1:1 | Block_3 | 2.2854 | 7.4014 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P61366 | Osteocrin | OSTN | 1:1 | Block_3 | 1.0964 | 9.8173 | TRUE | NA | NA | NA | NA | 1.5788 | NA | NA | NA
Q7RTW8 | Otoancorin | OTOA | 1:1 | Block_3 | 0.79025 | 8.9621 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5VV17 | OTU domain-containing protein 1 | OTUD1 | 1:1 | Block_3 | -0.038789 | 12.2198 | TRUE | 10.9065 | 8.1166 | NA | NA | 0.42158 | NA | 0.12893 | 2.2529
Q8WZ82 | Esterase OVCA2 | OVCA2 | 1:1 | Block_3 | 0.81324 | 7.8787 | FALSE | NA | NA | NA | NA | 1.4331 | NA | 1.627 | 2.4397
Q8IVL6 | Prolyl 3-hydroxylase 3 | P3H3 | 1:1 | Block_3 | 0.37694 | 8.6125 | TRUE | NA | NA | NA | NA | 0.38992 | NA | NA | NA
Q9Y2J8 | Protein-arginine deiminase type-2 | PADI2 | 1:1 | Block_3 | 1.2731 | 12.9859 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.8142
Q9UM07 | Protein-arginine deiminase type-4 | PADI4 | 1:1 | Block_3 | -0.6273 | 5.8595 | FALSE | 10.5527 | 5.6621 | NA | NA | 0.16627 | NA | -0.47581 | 1.5577
Q99487 | Platelet-activating factor acetylhydrolase 2, cytoplasmic | PAFAH2 | 1:1 | Block_3 | 0.48 | 9.1126 | FALSE | NA | 15.3249 | NA | 0.681 | 1.5794 | NA | 1.577 | 3.069
Q9H074 | Polyadenylate-binding protein-interacting protein 1 | PAIP1 | 1:1 | Block_3 | 0.69908 | 7.5705 | TRUE | NA | NA | NA | NA | 1.2431 | NA | 1.085 | 2.3328
Q9ULR5 | Polyadenylate-binding protein-interacting protein 2B | PAIP2B | 1:1 | Block_3 | 1.1048 | 8.7717 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5192
O96013 | Serine/threonine-protein kinase PAK 4 | PAK4 | 1:1 | Block_3 | -1.5715 | 8.3165 | FALSE | 8.1961 | 14.2031 | NA | -1.3999 | 0.57159 | NA | 0.227 | 1.9185
Q8WX93 | Palladin | PALLD | 1:1 | Block_3 | 1.5013 | 9.5282 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q13219 | Pappalysin-1 | PAPPA | 1:1 | Block_3 | 0.81845 | 6.6382 | FALSE | 12.7856 | 13.1525 | NA | NA | 1.2529 | NA | NA | 1.0707
Q8TEW0 | Partitioning defective 3 homolog | PARD3 | 1:1 | Block_3 | 1.7075 | 11.1344 | FALSE | 2.6812 | 3.1057 | NA | NA | 3.2984 | NA | 2.1193 | 3.9289
P09874 | Poly [ADP-ribose] polymerase 1 | PARP1 | 1:1 | Block_3 | 0.0042425 | 10.8618 | FALSE | NA | NA | NA | NA | 0.37124 | NA | 0.34995 | 3.9319
P30039 | Phenazine biosynthesis-like domain-containing protein | PBLD | 1:1 | Block_3 | 0.51978 | 6.7052 | FALSE | NA | NA | NA | NA | 0.70038 | NA | NA | 1.3713
A6NGG8 | Photoreceptor cilium actin regulator | PCARE | 1:1 | Block_3 | 1.1565 | 11.7077 | TRUE | 4.0086 | 15.5931 | NA | NA | 1.1869 | NA | NA | 1.2395
Q9Y5G1 | Protocadherin gamma-B3 | PCDHGB3 | 1:1 | Block_3 | 1.0212 | 7.4599 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O95613 | Pericentrin | PCNT | 1:1 | Block_3 | 1.7907 | 7.4508 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N4E4 | Phosducin-like protein 2 | PDCL2 | 1:1 | Block_3 | 0.95353 | 12.4715 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.3735
Q9Y233 | cAMP and cAMP-inhibited cGMP 3',5'-cyclic phosphodiesterase 10A | PDE10A | 1:1 | Block_3 | -1.9177 | 10.5838 | FALSE | NA | NA | NA | NA | NA | NA | NA | -1.6148
Q14123 | Dual specificity calcium/calmodulin-dependent 3',5'-cyclic nucleotide phosphodiesterase 1C | PDE1C | 1:1 | Block_3 | 0.65722 | 10.4975 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q08499 | cAMP-specific 3',5'-cyclic phosphodiesterase 4D | PDE4D | 1:1 | Block_3 | 1.11 | 3.1009 | FALSE | NA | NA | NA | NA | 1.8663 | NA | NA | NA
Q9NRA1 | Platelet-derived growth factor C | PDGFC | 1:1 | Block_3 | 0.21999 | 6.7108 | FALSE | 11.9385 | 7.9448 | NA | 0.53485 | 0.96639 | NA | 0.72395 | 1.2775
Q13087 | Protein disulfide-isomerase A2 | PDIA2 | 1:1 | Block_3 | 1.2365 | 8.5645 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q14554 | Protein disulfide-isomerase A5 | PDIA5 | 1:1 | Block_3 | 1.5895 | 11.3454 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P0J1 | [Pyruvate dehydrogenase [acetyl-transferring]]-phosphatase 1, mitochondrial | PDP1 | 1:1 | Block_3 | 2.4576 | 10.1774 | FALSE | NA | NA | NA | NA | 3.122 | NA | 2.5476 | 4.5871
Q9NUG6 | p53 and DNA damage-regulated protein 1 | PDRG1 | 1:1 | Block_3 | 1.6493 | 9.3962 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.452
Q6P996 | Pyridoxal-dependent decarboxylase domain-containing protein 1 | PDXDC1 | 1:1 | Block_3 | 1.2996 | 4.7076 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.3299
O15018 | PDZ domain-containing protein 2 | PDZD2 | 1:1 | Block_3 | 1.0424 | 12.8219 | FALSE | NA | NA | NA | NA | 1.5072 | NA | NA | 2.8364
Q96S96 | Phosphatidylethanolamine-binding protein 4 | PEBP4 | 1:1 | Block_3 | -3.63 | 6.1309 | TRUE | 12.2263 | 9.5092 | -2.6323 | -0.58626 | 0.43856 | NA | -0.41121 | 1.1949
Q9BRX2 | Protein pelota homolog | PELO | 1:1 | Block_3 | 1.3775 | 11.4066 | TRUE | NA | NA | NA | 1.9605 | 3.8064 | NA | 3.8432 | 5.7951
Q9NQP4 | Prefoldin subunit 4 | PFDN4 | 1:1 | Block_3 | 1.1322 | 5.0978 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O15212 | Prefoldin subunit 6 | PFDN6 | 1:1 | Block_3 | -0.11 | 10.593 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96LB8 | Peptidoglycan recognition protein 4 | PGLYRP4 | 1:1 | Block_3 | 0.57578 | 12.1283 | TRUE | 29.9276 | 15.4457 | NA | NA | 1.1001 | NA | NA | 1.7877
Q96G03 | Phosphopentomutase | PGM2 | 1:1 | Block_3 | 1.8879 | 12.6739 | FALSE | 5.8887 | 1.1802 | NA | NA | 1.8905 | NA | NA | 2.6947
P06401 | Progesterone receptor | PGR | 1:1 | Block_3 | 1.1842 | 8.7048 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BVI0 | PHD finger protein 20 | PHF20 | 1:1 | Block_3 | 0.20085 | 13.0103 | TRUE | 1.1443 | 12.7277 | NA | NA | 0.25838 | NA | NA | 0.59175
Q86SQ0 | Pleckstrin homology-like domain family B member 2 | PHLDB2 | 1:1 | Block_3 | 1.1229 | 12.1851 | FALSE | 11.6868 | 24.32 | NA | NA | 1.495 | NA | NA | 1.6915
Q9Y2I7 | 1-phosphatidylinositol 3-phosphate 5-kinase | PIKFYVE | 1:1 | Block_3 | 2.7455 | 12.7095 | TRUE | NA | NA | NA | NA | NA | NA | NA | 4.9595
Q9Y237 | Peptidyl-prolyl cis-trans isomerase NIMA-interacting 4 | PIN4 | 1:1 | Block_3 | 0.60783 | 12.2327 | TRUE | 12.2305 | 25.6925 | NA | NA | 1.8109 | NA | 0.71516 | 2.4056
Q9GZP4 | PITH domain-containing protein 1 | PITHD1 | 1:1 | Block_3 | 0.53067 | 7.6494 | TRUE | NA | NA | NA | NA | 0.56065 | NA | NA | 0.74628
O43164 | E3 ubiquitin-protein ligase Praja-2 | PJA2 | 1:1 | Block_3 | 1.0337 | 8.4313 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5463
Q6P5Z2 | Serine/threonine-protein kinase N3 | PKN3 | 1:1 | Block_3 | -1.0356 | 8.4895 | FALSE | NA | NA | NA | NA | NA | NA | NA | -0.53836
Q9UNK4 | Group IID secretory phospholipase A2 | PLA2G2D | 1:1 | Block_3 | 1.1025 | 5.9244 | FALSE | NA | NA | NA | NA | 1.2471 | NA | NA | NA
Q96KN8 | Phospholipase A and acyltransferase 5 | PLAAT5 | 1:1 | Block_3 | 2.2918 | 13.6605 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NQ66 | 1-phosphatidylinositol 4,5-bisphosphate phosphodiesterase beta-1 | PLCB1 | 1:1 | Block_3 | 0.95616 | 7.9965 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9ULL1 | Pleckstrin homology domain-containing family G member 1 | PLEKHG1 | 1:1 | Block_3 | 0.53779 | 9.4109 | FALSE | 19.9795 | 8.5755 | NA | NA | NA | NA | NA | 1.0141
O60240 | Perilipin-1 | PLIN1 | 1:1 | Block_3 | -0.3625 | 8.6926 | FALSE | 11.0221 | 12.9677 | NA | 0.47433 | 1.1892 | -0.2477 | 0.48513 | 1.6294
Q8IUK5 | Plexin domain-containing protein 1 | PLXDC1 | 1:1 | Block_3 | 0.48176 | 8.8605 | TRUE | 6.7754 | 8.3534 | NA | 0.58399 | 1.0144 | NA | NA | 1.1461
O15305 | Phosphomannomutase 2 | PMM2 | 1:1 | Block_3 | 1.2535 | 6.772 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P54277 | PMS1 protein homolog 1 | PMS1 | 1:1 | Block_3 | 2 | 10.2954 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q15126 | Phosphomevalonate kinase | PMVK | 1:1 | Block_3 | 0.83318 | 6.4663 | FALSE | NA | NA | NA | 2.2408 | 5.2576 | NA | 5.476 | 7.1063
Q96T60 | Bifunctional polynucleotide phosphatase/kinase | PNKP | 1:1 | Block_3 | 2.1537 | 11.407 | FALSE | NA | NA | NA | NA | 3.102 | NA | NA | NA
Q8ND90 | Paraneoplastic antigen Ma1 | PNMA1 | 1:1 | Block_3 | 2.4172 | 8.0534 | FALSE | NA | NA | NA | NA | 2.977 | NA | NA | 4.9142
Q9UL42 | Paraneoplastic antigen Ma2 | PNMA2 | 1:1 | Block_3 | 1.605 | 5.3827 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TCS8 | Polyribonucleotide nucleotidyltransferase 1, mitochondrial | PNPT1 | 1:1 | Block_3 | 1.4966 | 12.2582 | FALSE | NA | NA | NA | NA | 3.3861 | NA | 2.886 | 4.7334
Q9H1D9 | DNA-directed RNA polymerase III subunit RPC6 | POLR3F | 1:1 | Block_3 | 1.2644 | 12.3253 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.883
O75334 | Liprin-alpha-2 | PPFIA2 | 1:1 | Block_3 | 0.56756 | 11.3237 | FALSE | NA | 1.6117 | NA | NA | 0.60363 | NA | NA | 0.99681
Q9UNP9 | Peptidyl-prolyl cis-trans isomerase E | PPIE | 1:1 | Block_3 | 1.3707 | 12.3609 | FALSE | NA | 12.3229 | NA | NA | 1.663 | NA | NA | 1.5971
P35813 | Protein phosphatase 1A | PPM1A | 1:1 | Block_3 | -1.0731 | 6.84 | TRUE | NA | NA | NA | NA | NA | NA | NA | -0.6515
O60237 | Protein phosphatase 1 regulatory subunit 12B | PPP1R12B | 1:1 | Block_3 | 0.95176 | 12.0813 | TRUE | 4.1045 | 14.4465 | NA | 0.95434 | 2.5368 | NA | 2.0816 | 5.2967
Q9UD71 | Protein phosphatase 1 regulatory subunit 1B | PPP1R1B | 1:1 | Block_3 | 1.2808 | 6.7128 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86WC6 | Protein phosphatase 1 regulatory subunit 27 | PPP1R27 | 1:1 | Block_3 | -2.36 | 6.4887 | FALSE | 15.8288 | 11.6124 | -0.76312 | 0.40129 | 1.2713 | -1.1174 | -0.088198 | 1.0124
Q6NXS1 | Protein phosphatase inhibitor 2 family member B | PPP1R2B | 1:1 | Block_3 | 0.965 | 10.7702 | TRUE | NA | NA | NA | NA | 2.6538 | NA | 3.3824 | 4.9398
P50897 | Palmitoyl-protein thioesterase 1 | PPT1 | 1:1 | Block_3 | -1.1443 | 8.1643 | FALSE | NA | NA | NA | NA | -0.21908 | NA | NA | -0.060012
O60828 | Polyglutamine-binding protein 1 | PQBP1 | 1:1 | Block_3 | 1.3001 | 7.4622 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P78395 | Melanoma antigen preferentially expressed in tumors | PRAME | 1:1 | Block_3 | 1.2031 | 6.6627 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O43663 | Protein regulator of cytokinesis 1 | PRC1 | 1:1 | Block_3 | 2.95 | 13.7371 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P30048 | Thioredoxin-dependent peroxide reductase, mitochondrial | PRDX3 | 1:1 | Block_3 | 1.1155 | 3.3409 | FALSE | NA | NA | NA | NA | 3.7906 | NA | 2.4355 | 3.9048
Q9HCU5 | Prolactin regulatory element-binding protein | PREB | 1:1 | Block_3 | 0.76174 | 5.7989 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.2036
Q9Y478 | 5'-AMP-activated protein kinase subunit beta-1 | PRKAB1 | 1:1 | Block_3 | 0.67485 | 10.3931 | TRUE | 14.5117 | 14.5255 | NA | NA | 0.96111 | NA | 1.1408 | 2.6278
Q9UGI9 | 5'-AMP-activated protein kinase subunit gamma-3 | PRKAG3 | 1:1 | Block_3 | 2.137 | 9.8502 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P31321 | cAMP-dependent protein kinase type I-beta regulatory subunit | PRKAR1B | 1:1 | Block_3 | 0.80134 | 10.7906 | TRUE | NA | NA | NA | NA | 2.0246 | NA | 1.4947 | 2.6612
Q04759 | Protein kinase C theta type | PRKCQ | 1:1 | Block_3 | 1.2492 | 7.8438 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.379
Q9BZL6 | Serine/threonine-protein kinase D2 | PRKD2 | 1:1 | Block_3 | 1.34 | 8.0064 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O75569 | Interferon-inducible double-stranded RNA-dependent protein kinase activator A | PRKRA | 1:1 | Block_3 | 1.0257 | 8.5566 | FALSE | NA | NA | NA | NA | 1.4778 | NA | 1.7021 | 3.0959
Q9UNN8 | Endothelial protein C receptor | PROCR | 1:1 | Block_3 | 1.9253 | 6.2442 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O43490 | Prominin-1 | PROM1 | 1:1 | Block_3 | 1.1378 | 10.8774 | FALSE | NA | NA | NA | NA | 1.3367 | NA | NA | 2.5584
P41219 | Peripherin | PRPH | 1:1 | Block_3 | 1.2065 | 13.1385 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.4059
Q8IV56 | Proline-rich protein 15 | PRR15 | 1:1 | Block_3 | 0.79983 | 9.0104 | FALSE | NA | NA | NA | NA | 1.0917 | NA | NA | 1.0168
P86496 | Proline-rich protein 20A | PRR20A | 1:1 | Block_3 | 0.41658 | 8.8788 | TRUE | 4.6684 | NA | NA | NA | 0.47598 | NA | NA | 0.46812
P86481 | Proline-rich protein 20B | PRR20B | 1:1 | Block_3 | 0.23313 | 9.6718 | FALSE | NA | 1.3666 | NA | NA | 0.47534 | NA | NA | 0.59436
A6NEV1 | Proline-rich protein 23A | PRR23A | 1:1 | Block_3 | 1.1076 | 12.646 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P85299 | Proline-rich protein 5 | PRR5 | 1:1 | Block_3 | 0.95431 | 7.8655 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96M27 | Protein PRRC1 | PRRC1 | 1:1 | Block_3 | -0.56912 | 4.0076 | FALSE | 17.6462 | 10.6867 | -0.10694 | 0.83872 | 1.607 | 0.41501 | 1.304 | 1.9216
Q9GZN4 | Brain-specific serine protease 4 | PRSS22 | 1:1 | Block_3 | 0.24117 | 10.164 | TRUE | 9.3149 | 8.9491 | NA | 0.38218 | 1.1997 | NA | 0.64378 | 1.6297
Q9NRG1 | Phosphoribosyltransferase domain-containing protein 1 | PRTFDC1 | 1:1 | Block_3 | 1.0315 | 9.8936 | FALSE | NA | NA | NA | 1.382 | 5.1868 | NA | 4.5579 | 7.3868
Q8WUY3 | Protein prune homolog 2 | PRUNE2 | 1:1 | Block_3 | 2.051 | 12.0043 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P25786 | Proteasome subunit alpha type-1 | PSMA1 | 1:1 | Block_3 | 0.001 | 4.8552 | FALSE | NA | 2.8839 | NA | NA | NA | NA | NA | NA
P17980 | 26S proteasome regulatory subunit 6A | PSMC3 | 1:1 | Block_3 | 2.101 | 9.2341 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q16401 | 26S proteasome non-ATPase regulatory subunit 5 | PSMD5 | 1:1 | Block_3 | 0.82477 | 5.7917 | FALSE | NA | NA | NA | NA | 1.1227 | NA | NA | 1.3315
O95456 | Proteasome assembly chaperone 1 | PSMG1 | 1:1 | Block_3 | -2.6347 | 9.7719 | FALSE | 13.1767 | 12.2114 | -0.96681 | 1.1562 | 2.7437 | 0.51367 | 2.5975 | 3.8424
Q6PGN9 | Proline/serine-rich coiled-coil protein 1 | PSRC1 | 1:1 | Block_3 | 1.5873 | 9.9345 | FALSE | NA | NA | NA | NA | 1.6663 | NA | NA | 2.8156
P60484 | Phosphatidylinositol 3,4,5-trisphosphate 3-phosphatase and dual-specificity protein phosphatase PTEN | PTEN | 1:1 | Block_3 | 2.2988 | 9.3323 | FALSE | NA | NA | NA | NA | 3.1215 | NA | 2.6528 | 4.2859
E9PB15 | Putative protein PTGES3L | PTGES3L | 1:1 | Block_3 | 1.1589 | 10.2688 | FALSE | NA | NA | NA | NA | 2.749 | NA | 2.0873 | 4.4096
O75365 | Protein tyrosine phosphatase type IVA 3 | PTP4A3 | 1:1 | Block_3 | -0.33188 | 8.4266 | FALSE | NA | 3.8078 | NA | NA | 0.66003 | -0.29676 | 0.6229 | 1.4943
Q8WUK0 | Phosphatidylglycerophosphatase and protein-tyrosine phosphatase 1 | PTPMT1 | 1:1 | Block_3 | 0.2701 | 11.327 | TRUE | 8.6021 | 10.7611 | NA | 0.36378 | 1.0541 | NA | NA | 1.4108
P43378 | Tyrosine-protein phosphatase non-receptor type 9 | PTPN9 | 1:1 | Block_3 | 1.6278 | 9.9788 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q03393 | 6-pyruvoyl tetrahydrobiopterin synthase | PTS | 1:1 | Block_3 | -0.61 | 9.2765 | TRUE | 12.3198 | 11.696 | NA | -0.41331 | 0.93066 | -0.46699 | 0.62136 | 2.4559
O95997 | Securin | PTTG1 | 1:1 | Block_3 | 3.4323 | 13.8553 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P49023 | Paxillin | PXN | 1:1 | Block_3 | -3 | 7.4971 | TRUE | 8.0142 | 13.4266 | -1.2136 | -0.36938 | 0.38721 | -0.82153 | 0.4338 | 3.9573
P10082 | Peptide YY | PYY | 1:1 | Block_3 | 0.001 | 8.858 | TRUE | 19.3155 | 9.0223 | NA | 0.42066 | 2.4372 | NA | 1.5461 | 3.8519
Q14088 | Ras-related protein Rab-33A | RAB33A | 1:1 | Block_3 | 1 | 8.1923 | FALSE | NA | NA | NA | NA | 1.5398 | NA | NA | 2.1972
O95755 | Ras-related protein Rab-36 | RAB36 | 1:1 | Block_3 | 1.8808 | 10.1578 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96AX2 | Ras-related protein Rab-37 | RAB37 | 1:1 | Block_3 | -3.8384 | 7.6653 | TRUE | 16.0198 | 7.9976 | -3.8376 | -0.096254 | 1.8032 | -1.7595 | 1.8534 | 3.0698
Q96DA2 | Ras-related protein Rab-39B | RAB39B | 1:1 | Block_3 | -1.0926 | 9.3538 | FALSE | NA | NA | NA | NA | NA | NA | NA | -0.30854
Q15042 | Rab3 GTPase-activating protein catalytic subunit | RAB3GAP1 | 1:1 | Block_3 | 2.3995 | 11.5414 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q7Z6P3 | Ras-related protein Rab-44 | RAB44 | 1:1 | Block_3 | -0.10818 | 7.8276 | FALSE | NA | 21.1016 | NA | NA | -0.089644 | NA | NA | 4.3022
Q9NRW1 | Ras-related protein Rab-6B | RAB6B | 1:1 | Block_3 | 0.36309 | 4.7942 | TRUE | NA | NA | NA | NA | 1.0946 | NA | NA | 1.0154
Q7Z6M1 | Rab9 effector protein with kelch motifs | RABEPK | 1:1 | Block_3 | 1 | NA | NA | NA | NA | NA | NA | 1.0003 | NA | NA | 1.6036
Q5R372 | Rab GTPase-activating protein 1-like | RABGAP1L | 1:1 | Block_3 | 0.33356 | 9.4757 | TRUE | 14.9961 | 14.7399 | NA | 0.75841 | 2.0845 | 0.39977 | 2.3279 | 3.6825
P60763 | Ras-related C3 botulinum toxin substrate 3 | RAC3 | 1:1 | Block_3 | 1.2295 | 7.8168 | FALSE | NA | NA | NA | NA | 1.4331 | NA | NA | NA
P54727 | UV excision repair protein RAD23 homolog B | RAD23B | 1:1 | Block_3 | -0.25704 | 5.3714 | FALSE | 12.7745 | 13.1784 | -0.083509 | 0.79538 | 2.2203 | 0.99213 | 2.156 | 2.9033
Q06609 | DNA repair protein RAD51 homolog 1 | RAD51 | 1:1 | Block_3 | 0.76929 | 12.364 | TRUE | 8.495 | 15.2501 | NA | NA | NA | NA | NA | NA
Q5JS13 | Ras-specific guanine nucleotide-releasing factor RalGPS1 | RALGPS1 | 1:1 | Block_3 | 0.41986 | 14.5883 | TRUE | 10.3887 | 11.6905 | NA | NA | 0.98033 | NA | NA | 1.4137
O60895 | Receptor activity-modifying protein 2 | RAMP2 | 1:1 | Block_3 | -4 | 6.9435 | TRUE | 10.2729 | 6.1175 | -0.23105 | 0.16576 | 0.57447 | 0.36818 | 1.0916 | 2.1603
P49792 | E3 SUMO-protein ligase RanBP2 | RANBP2 | 1:1 | Block_3 | 0.59902 | 8.3236 | FALSE | NA | NA | NA | NA | 0.63852 | NA | NA | 1.5544
P46060 | Ran GTPase-activating protein 1 | RANGAP1 | 1:1 | Block_3 | 1 | 6.2825 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P62834 | Ras-related protein Rap-1A | RAP1A | 1:1 | Block_3 | 1.5384 | 7.3972 | FALSE | 9.6507 | 11.5392 | NA | NA | NA | NA | NA | NA
P47736 | Rap1 GTPase-activating protein 1 | RAP1GAP | 1:1 | Block_3 | 0.40524 | 10.5624 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.76379
P52306 | Rap1 GTPase-GDP dissociation stimulator 1 | RAP1GDS1 | 1:1 | Block_3 | -0.61079 | 10.6974 | FALSE | 3.9176 | 2.6239 | NA | NA | 0.53205 | NA | NA | 0.84068
P49788 | Retinoic acid receptor responder protein 1 | RARRES1 | 1:1 | Block_3 | 0.1222 | 9.9718 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P20936 | Ras GTPase-activating protein 1 | RASA1 | 1:1 | Block_3 | 0.94758 | 6.4452 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q13972 | Ras-specific guanine nucleotide-releasing factor 1 | RASGRF1 | 1:1 | Block_3 | 1.8765 | 12.457 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P50749 | Ras association domain-containing protein 2 | RASSF2 | 1:1 | Block_3 | 0.64446 | 10.3362 | FALSE | NA | 14.1011 | NA | NA | 0.65505 | NA | NA | 4.4059
A6NFN3 | RNA binding protein fox-1 homolog 3 | RBFOX3 | 1:1 | Block_3 | -1.1867 | 10.591 | TRUE | 11.3061 | 14.964 | -0.72015 | -0.012036 | 0.64621 | -0.029572 | 1.0902 | 3.2941
Q9Y4C8 | Probable RNA-binding protein 19 | RBM19 | 1:1 | Block_3 | 1.0499 | 10.0522 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P49756 | RNA-binding protein 25 | RBM25 | 1:1 | Block_3 | 2.0893 | 10.4574 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P98179 | RNA-binding protein 3 | RBM3 | 1:1 | Block_3 | 0.93786 | 10.8945 | TRUE | 11.7673 | 8.9897 | NA | 1.8863 | 3.6535 | 2.0206 | 3.5971 | 5.1722
P09455 | Retinol-binding protein 1 | RBP1 | 1:1 | Block_3 | 1.5901 | 10.1276 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.8817
Q5TC82 | Roquin-1 | RC3H1 | 1:1 | Block_3 | 0.85994 | 9.7542 | TRUE | 8.0454 | 13.0803 | NA | 1.1974 | 1.8528 | 0.919 | 2.0648 | 3.886
P18754 | Regulator of chromosome condensation | RCC1 | 1:1 | Block_3 | 0.1645 | 8.5033 | TRUE | NA | NA | NA | NA | 0.83736 | NA | 0.23443 | 1.6531
Q9UKL0 | REST corepressor 1 | RCOR1 | 1:1 | Block_3 | 0.49846 | 12.8671 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.2657
Q01201 | Transcription factor RelB | RELB | 1:1 | Block_3 | 1.5027 | 9.3964 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P51606 | N-acylglucosamine 2-epimerase | RENBP | 1:1 | Block_3 | 1.078 | 9.0839 | FALSE | NA | NA | NA | 1.4213 | 4.2639 | 1.1793 | 4.3715 | 5.9717
Q96D71 | RalBP1-associated Eps domain-containing protein 1 | REPS1 | 1:1 | Block_3 | 4.0054 | 7.013 | FALSE | 17.1504 | NA | NA | NA | 4.0288 | NA | NA | 6.7123
Q13127 | RE1-silencing transcription factor | REST | 1:1 | Block_3 | 0.59629 | 11.5723 | FALSE | NA | 9.6059 | NA | NA | NA | NA | NA | 0.92243
Q9Y3B8 | Oligoribonuclease, mitochondrial | REXO2 | 1:1 | Block_3 | 2.1819 | 8.9724 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WZ73 | E3 ubiquitin-protein ligase rififylin | RFFL | 1:1 | Block_3 | 1.0668 | 7.8608 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P48378 | DNA-binding protein RFX2 | RFX2 | 1:1 | Block_3 | 0.64 | 10.5624 | FALSE | 4.5316 | 8.3424 | NA | NA | 1.1083 | NA | 1.181 | 2.3662
O15211 | Ral guanine nucleotide dissociation stimulator-like 2 | RGL2 | 1:1 | Block_3 | 1.8115 | 8.5669 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P57771 | Regulator of G-protein signaling 8 | RGS8 | 1:1 | Block_3 | -0.1013 | 8.334 | FALSE | 13.825 | 20.374 | NA | NA | 1.0579 | NA | NA | 0.76676
Q6R327 | Rapamycin-insensitive companion of mTOR | RICTOR | 1:1 | Block_3 | 1.5366 | 7.3812 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O95786 | Antiviral innate immune response receptor RIG-I | RIGI | 1:1 | Block_3 | 0.39065 | 8.8795 | FALSE | NA | NA | NA | 0.44235 | 1.7024 | 0.44907 | 2.0262 | 4.2562
Q5EBL4 | RILP-like protein 1 | RILPL1 | 1:1 | Block_3 | 1.2988 | 9.8268 | TRUE | NA | NA | NA | NA | 2.5376 | NA | 2.7439 | 4.2579
A6NDU8 | RAB7A-interacting MON1-CCZ1 complex subunit 1 | RIMOC1 | 1:1 | Block_3 | 1.3611 | 10.4022 | FALSE | NA | NA | NA | NA | 2.8481 | NA | 2.4159 | 4.0893
P57078 | Receptor-interacting serine/threonine-protein kinase 4 | RIPK4 | 1:1 | Block_3 | 0.15544 | 7.9238 | FALSE | 18.9527 | 19.2528 | NA | NA | 0.85621 | NA | NA | 2.7256
Q8NCN4 | E3 ubiquitin-protein ligase RNF169 | RNF169 | 1:1 | Block_3 | 1.0444 | 13.9655 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.9433
Q96BH1 | E3 ubiquitin-protein ligase RNF25 | RNF25 | 1:1 | Block_3 | 0.91958 | 10.9666 | TRUE | NA | NA | NA | NA | 1.3773 | NA | 1.1026 | 2.5735
Q96EP0 | E3 ubiquitin-protein ligase RNF31 | RNF31 | 1:1 | Block_3 | 3 | 9.4373 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P78317 | E3 ubiquitin-protein ligase RNF4 | RNF4 | 1:1 | Block_3 | 0.96721 | 9.4706 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.6972
O75150 | E3 ubiquitin-protein ligase BRE1B | RNF40 | 1:1 | Block_3 | 1.1917 | 8.9634 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H4P4 | E3 ubiquitin-protein ligase NRDP1 | RNF41 | 1:1 | Block_3 | -0.085259 | 6.3266 | FALSE | 17.9262 | 10.4802 | NA | 0.66553 | 2.1215 | 0.52227 | 1.9944 | 3.5534
Q68DV7 | E3 ubiquitin-protein ligase RNF43 | RNF43 | 1:1 | Block_3 | 1.22 | 8.5337 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q92753 | Nuclear receptor ROR-beta | RORB | 1:1 | Block_3 | 0.40425 | 6.4313 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P15927 | Replication protein A 32 kDa subunit | RPA2 | 1:1 | Block_3 | 2.17 | 10.522 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.2219
Q13156 | Replication protein A 30 kDa subunit | RPA4 | 1:1 | Block_3 | -1.9822 | 10.5466 | FALSE | 11.3138 | 18.7723 | NA | -1.8082 | 0.21793 | NA | -1.8381 | 0.59109
Q9H6T3 | RNA polymerase II-associated protein 3 | RPAP3 | 1:1 | Block_3 | 1.0974 | 8.7288 | FALSE | NA | NA | NA | NA | 3.4612 | NA | 3.4942 | 4.9787
Q96AT9 | Ribulose-phosphate 3-epimerase | RPE | 1:1 | Block_3 | 1.2319 | 7.1678 | TRUE | NA | NA | NA | NA | 2.9086 | NA | 2.4575 | 4.1583
Q92834 | X-linked retinitis pigmentosa GTPase regulator | RPGR | 1:1 | Block_3 | 1.66 | 10.7158 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y2J0 | Rabphilin-3A | RPH3A | 1:1 | Block_3 | -1.87 | 10.4825 | TRUE | 19.6413 | 17.9548 | -1.5582 | -0.66349 | 0.3217 | -0.92812 | 0.0064328 | 3.3596
Q9NQG5 | Regulation of nuclear pre-mRNA domain-containing protein 1B | RPRD1B | 1:1 | Block_3 | 0.55018 | 5.4734 | FALSE | NA | NA | NA | NA | 1.1589 | NA | 1.0227 | 2.0109
P46783 | 40S ribosomal protein S10 | RPS10 | 1:1 | Block_3 | 1.624 | 2.136 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P10301 | Ras-related protein R-Ras | RRAS | 1:1 | Block_3 | 1.1545 | 14.097 | FALSE | NA | NA | NA | NA | 2.8115 | NA | 2.4774 | 6.0821
P31350 | Ribonucleoside-diphosphate reductase subunit M2 | RRM2 | 1:1 | Block_3 | -0.15 | 6.3284 | FALSE | 14.824 | 14.896 | NA | 0.072642 | 0.78881 | NA | 0.75059 | 2.14
Q7LG56 | Ribonucleoside-diphosphate reductase subunit M2 B | RRM2B | 1:1 | Block_3 | 2.1292 | 14.6121 | FALSE | NA | NA | NA | NA | 2.4568 | NA | 2.2283 | 3.7912
P56182 | Ribosomal RNA processing protein 1 homolog A | RRP1 | 1:1 | Block_3 | -0.22024 | 8.8077 | FALSE | 7.6995 | 5.6373 | NA | NA | 1.3837 | NA | NA | 0.9224
Q9Y3B9 | RRP15-like protein | RRP15 | 1:1 | Block_3 | 1.55 | 12.9091 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IZC4 | Rhotekin-2 | RTKN2 | 1:1 | Block_3 | 1.6228 | 8.1846 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y265 | RuvB-like 1 | RUVBL1 | 1:1 | Block_3 | -0.78 | 1.4251 | FALSE | NA | NA | NA | NA | 0.36047 | NA | NA | 1.0529
Q8N488 | RING1 and YY1-binding protein | RYBP | 1:1 | Block_3 | 0.96056 | 10.4255 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.1605
P21817 | Ryanodine receptor 1 | RYR1 | 1:1 | Block_3 | 1.6472 | 9.1099 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9HCY8 | Protein S100-A14 | S100A14 | 1:1 | Block_3 | 0.80644 | 9.0611 | TRUE | NA | NA | NA | NA | 1.6314 | NA | NA | 1.3907
P33764 | Protein S100-A3 | S100A3 | 1:1 | Block_3 | 0.93439 | 11.0231 | TRUE | 30.637 | 20.0002 | NA | NA | 1.069 | NA | NA | 1.9841
P10523 | S-arrestin | SAG | 1:1 | Block_3 | 1.2179 | 9.6583 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IZD0 | Sterile alpha motif domain-containing protein 14 | SAMD14 | 1:1 | Block_3 | 1.1557 | 9.8327 | FALSE | 21.607 | 11.1241 | NA | 1.313 | 4.918 | NA | 4.3675 | 6.8174
O00422 | Histone deacetylase complex subunit SAP18 | SAP18 | 1:1 | Block_3 | 1.4058 | 9.9692 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O43290 | U4/U6.U5 tri-snRNP-associated protein 1 | SART1 | 1:1 | Block_3 | 0.12247 | 10.9539 | TRUE | 5.4646 | 2.2574 | NA | NA | 0.18814 | NA | 0.37111 | 2.8386
Q6UVJ0 | Spindle assembly abnormal protein 6 homolog | SASS6 | 1:1 | Block_3 | 0.6703 | 13.2264 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.2054
P21673 | Diamine acetyltransferase 1 | SAT1 | 1:1 | Block_3 | 1.0417 | 9.451 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96F10 | Thialysine N-epsilon-acetyltransferase | SAT2 | 1:1 | Block_3 | -0.13912 | 10.7377 | TRUE | 8.1735 | 21.5496 | NA | 0.38551 | 1.632 | -0.0020627 | 1.3705 | 3.1075
Q9H4B6 | Protein salvador homolog 1 | SAV1 | 1:1 | Block_3 | 0.10017 | 11.0483 | FALSE | NA | NA | NA | NA | 0.55348 | NA | 0.10886 | 2.0948
Q8IVN8 | Somatomedin-B and thrombospondin type-1 domain-containing protein | SBSPON | 1:1 | Block_3 | -0.59495 | 7.6476 | FALSE | 7.2988 | 12.7687 | -0.58174 | 0.54267 | 1.2956 | -0.53003 | 0.58211 | 1.8866
Q8WTV0 | Scavenger receptor class B member 1 | SCARB1 | 1:1 | Block_3 | 1.2632 | 11.9546 | FALSE | 11.5027 | 20.2968 | NA | NA | 3.2335 | NA | NA | 2.6426
Q13296 | Mammaglobin-A | SCGB2A2 | 1:1 | Block_3 | 1.66 | 9.1094 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O76038 | Secretagogin | SCGN | 1:1 | Block_3 | -0.067254 | 7.0572 | FALSE | 10.4822 | 6.3583 | NA | 0.22202 | 0.722 | NA | 0.38112 | 1.5231
Q9Y6U3 | Scinderin | SCIN | 1:1 | Block_3 | 1.0664 | 10.7032 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O60939 | Sodium channel subunit beta-2 | SCN2B | 1:1 | Block_3 | 1.47 | 9.476 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7483
Q9NY46 | Sodium channel protein type 3 subunit alpha | SCN3A | 1:1 | Block_3 | 1.54 | 10.3473 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NY72 | Sodium channel subunit beta-3 | SCN3B | 1:1 | Block_3 | 0.9348 | 8.3037 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P22307 | Sterol carrier protein 2 | SCP2 | 1:1 | Block_3 | 0.36705 | 9.9305 | FALSE | NA | NA | NA | NA | 1.0457 | NA | 0.82815 | 2.1592
P09683 | Secretin | SCT | 1:1 | Block_3 | 0.53 | 5.627 | FALSE | 26.1912 | 26.914 | NA | 0.57002 | 1.4545 | NA | NA | 1.9852
Q8IZQ5 | Selenoprotein H | SELENOH | 1:1 | Block_3 | 0.5 | 8.0433 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.892
Q9BVL4 | Protein adenylyltransferase SelO, mitochondrial | SELENOO | 1:1 | Block_3 | 0.45118 | 9.8796 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.46164
Q16181 | Septin-7 | SEPTIN7 | 1:1 | Block_3 | 1.7753 | 11.4553 | TRUE | NA | NA | NA | NA | 3.0954 | NA | 2.0665 | 4.0673
P36952 | Serpin B5 | SERPINB5 | 1:1 | Block_3 | 1.8335 | 4.7303 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P50454 | Serpin H1 | SERPINH1 | 1:1 | Block_3 | 1.9758 | 4.3617 | FALSE | NA | NA | NA | NA | 4.8665 | NA | 4.2745 | 6.7821
Q8IWL1 | Pulmonary surfactant-associated protein A2 | SFTPA2 | 1:1 | Block_3 | -0.18809 | 7.6245 | TRUE | 8.4599 | 12.157 | NA | 0.39877 | 2.2417 | NA | 0.19159 | 1.312
Q96ES7 | SAGA-associated factor 29 | SGF29 | 1:1 | Block_3 | 1.2753 | 11.2841 | FALSE | NA | NA | NA | NA | 1.4509 | NA | 1.3771 | 2.6347
Q9UQQ2 | SH2B adapter protein 3 | SH2B3 | 1:1 | Block_3 | 1.1684 | 8.8798 | FALSE | NA | NA | NA | NA | 2.4906 | NA | 2.0069 | 3.8785
Q9UJC5 | SH3 domain-binding glutamic acid-rich-like protein 2 | SH3BGRL2 | 1:1 | Block_3 | 0.92302 | 9.8112 | TRUE | 16.8743 | 9.6845 | NA | NA | 2.4394 | NA | 1.7308 | 3.331
O60239 | SH3 domain-binding protein 5 | SH3BP5 | 1:1 | Block_3 | -0.30443 | 11.2349 | TRUE | 13.5725 | 8.9223 | NA | 0.11857 | 2.6881 | NA | 2.1716 | 4.5775
Q7L8J4 | SH3 domain-binding protein 5-like | SH3BP5L | 1:1 | Block_3 | 1.2073 | 11.3938 | TRUE | NA | NA | NA | NA | 2.9517 | NA | 2.9105 | 4.4898
Q5HYK7 | SH3 domain-containing protein 19 | SH3D19 | 1:1 | Block_3 | 1.0099 | 9.744 | TRUE | NA | 4.03 | NA | NA | NA | NA | NA | 1.056
Q99963 | Endophilin-A3 | SH3GL3 | 1:1 | Block_3 | 1.53 | 10.2561 | FALSE | 23.2169 | 32.6277 | NA | NA | NA | NA | NA | 1.7135
P29353 | SHC-transforming protein 1 | SHC1 | 1:1 | Block_3 | 0.94638 | 7.6221 | FALSE | 7.2899 | 2.1846 | NA | NA | NA | NA | NA | NA
Q92529 | SHC-transforming protein 3 | SHC3 | 1:1 | Block_3 | 1.6107 | 10.6539 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96IW2 | SH2 domain-containing adapter protein D | SHD | 1:1 | Block_3 | 1.0578 | 9.7748 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UHJ6 | Sedoheptulokinase | SHPK | 1:1 | Block_3 | 0.77433 | 6.0452 | FALSE | NA | 9.1715 | NA | NA | 0.79063 | NA | NA | NA
Q149N8 | E3 ubiquitin-protein ligase SHPRH | SHPRH | 1:1 | Block_3 | 0.50828 | 11.2428 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.60222
Q6ZMC9 | Sialic acid-binding Ig-like lectin 15 | SIGLEC15 | 1:1 | Block_3 | 0.88 | 9.9684 | TRUE | 0.64042 | 8.5736 | NA | NA | 1.0495 | NA | NA | 0.95461
Q96FS4 | Signal-induced proliferation-associated protein 1 | SIPA1 | 1:1 | Block_3 | -0.21113 | 10.7744 | TRUE | 25.8913 | 14.4643 | NA | 0.26626 | 1.8098 | 0.34946 | 2.0765 | 3.8631
Q96EB6 | NAD-dependent protein deacetylase sirtuin-1 | SIRT1 | 1:1 | Block_3 | 1.707 | 10.7658 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7475
Q9NXA8 | NAD-dependent protein deacylase sirtuin-5, mitochondrial | SIRT5 | 1:1 | Block_3 | 2.5 | 7.6423 | FALSE | 8.1116 | 23.4517 | NA | NA | NA | NA | NA | 3.4583
Q9Y3P8 | Signaling threshold-regulating transmembrane adapter 1 | SIT1 | 1:1 | Block_3 | 0.39565 | 10.066 | TRUE | 22.0066 | 11.5521 | NA | 0.71901 | 1.551 | NA | 0.59163 | 1.8099
Q15477 | Helicase SKI2W | SKIC2 | 1:1 | Block_3 | 1.5884 | 8.6522 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q13291 | Signaling lymphocytic activation molecule | SLAMF1 | 1:1 | Block_3 | -0.53138 | 7.6184 | FALSE | 15.0134 | 20.5177 | -0.42461 | 0.16473 | 0.86171 | -0.40559 | 0.3962 | 1.505
Q96DU3 | SLAM family member 6 | SLAMF6 | 1:1 | Block_3 | 2.0514 | 7.4073 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P0V8 | SLAM family member 8 | SLAMF8 | 1:1 | Block_3 | -2.39 | 9.8086 | TRUE | 19.7987 | 14.4564 | -0.58653 | 0.16648 | 0.97659 | -0.42087 | 0.7687 | 2.0153
P55011 | Solute carrier family 12 member 2 | SLC12A2 | 1:1 | Block_3 | -0.28048 | 7.7618 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.6293
P53985 | Monocarboxylate transporter 1 | SLC16A1 | 1:1 | Block_3 | -1.0935 | 8.8201 | FALSE | 7.8509 | 11.4328 | 0.13094 | 0.68909 | 1.4251 | 0.63446 | 1.5175 | 2.7094
P43007 | Neutral amino acid transporter A | SLC1A4 | 1:1 | Block_3 | 2 | 8.7265 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6P1M0 | Long-chain fatty acid transport protein 4 | SLC27A4 | 1:1 | Block_3 | 0.93346 | 11.385 | TRUE | NA | NA | NA | NA | 2.479 | NA | 2.4347 | 4.8225
O00337 | Sodium/nucleoside cotransporter 1 | SLC28A1 | 1:1 | Block_3 | 1.1315 | 12.2808 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7049
Q8N130 | Sodium-dependent phosphate transport protein 2C | SLC34A3 | 1:1 | Block_3 | 1.9413 | 11.6428 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q15043 | Metal cation symporter ZIP14 | SLC39A14 | 1:1 | Block_3 | -1.0347 | 5.3398 | FALSE | 21.9418 | 12.0555 | -0.41083 | 0.38663 | 1.2448 | 0.08413 | 0.87993 | 2.0992
Q9NP59 | Solute carrier family 40 member 1 | SLC40A1 | 1:1 | Block_3 | 0.68106 | 12.5449 | TRUE | 2.9427 | NA | NA | NA | NA | NA | NA | 1.2277
Q53GD3 | Choline transporter-like protein 4 | SLC44A4 | 1:1 | Block_3 | 1.1649 | 8.7881 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86UW2 | Organic solute transporter subunit beta | SLC51B | 1:1 | Block_3 | -2.1753 | 9.236 | TRUE | 11.7777 | 12.2381 | -1.17 | -0.28095 | 1.0874 | -1.1493 | -0.14314 | 1.0018
Q9GZT3 | SRA stem-loop-interacting RNA-binding protein, mitochondrial | SLIRP | 1:1 | Block_3 | 0.96348 | 7.7003 | TRUE | 11.5021 | 5.9478 | NA | NA | 2.0208 | NA | NA | 2.0468
O94813 | Slit homolog 2 protein | SLIT2 | 1:1 | Block_3 | -0.53489 | 8.2749 | TRUE | 14.4519 | 10.9382 | -0.46019 | 0.14313 | 0.82121 | -0.34474 | 0.50736 | 1.3221
Q9H2G2 | STE20-like serine/threonine-protein kinase | SLK | 1:1 | Block_3 | 1.96 | 10.726 | TRUE | 4.2832 | 5.6107 | NA | NA | NA | NA | NA | NA
Q15797 | Mothers against decapentaplegic homolog 1 | SMAD1 | 1:1 | Block_3 | 1.32 | 12.4763 | FALSE | NA | NA | NA | NA | 3.5032 | NA | 3.5402 | 5.4966
Q15796 | Mothers against decapentaplegic homolog 2 | SMAD2 | 1:1 | Block_3 | 2 | 7.0232 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.3057
P51531 | Probable global transcription activator SNF2L2 | SMARCA2 | 1:1 | Block_3 | 1.1155 | 9.827 | TRUE | NA | NA | NA | NA | 1.1692 | NA | 1.171 | 2.5036
Q9UQE7 | Structural maintenance of chromosomes protein 3 | SMC3 | 1:1 | Block_3 | 1.5898 | 6.9963 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P49901 | Sperm mitochondrial-associated cysteine-rich protein | SMCP | 1:1 | Block_3 | 0.65 | 14.1918 | TRUE | NA | NA | NA | NA | 1.0485 | NA | NA | NA
Q9NY59 | Sphingomyelin phosphodiesterase 3 | SMPD3 | 1:1 | Block_3 | 1.02 | 9.4915 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.2391
P52788 | Spermine synthase | SMS | 1:1 | Block_3 | 0.5625 | 7.4731 | FALSE | NA | NA | NA | NA | 1.0815 | NA | NA | 0.68272
P60880 | Synaptosomal-associated protein 25 | SNAP25 | 1:1 | Block_3 | 1.1733 | 12.394 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.6458
P55769 | NHP2-like protein 1 | SNU13 | 1:1 | Block_3 | -0.20209 | 9.0903 | FALSE | NA | NA | NA | NA | -0.12987 | NA | NA | 3.7239
Q9BX66 | Sorbin and SH3 domain-containing protein 1 | SORBS1 | 1:1 | Block_3 | 1.2713 | 8.0639 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.3171
Q96PQ0 | VPS10 domain-containing receptor SorCS2 | SORCS2 | 1:1 | Block_3 | 1.1386 | 9.1809 | FALSE | 12.3083 | 17.0625 | NA | 1.1743 | 1.9516 | NA | 1.6988 | 3.0691
Q6X4U4 | Sclerostin domain-containing protein 1 | SOSTDC1 | 1:1 | Block_3 | 1.397 | 9.0052 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q2M3V2 | Ankyrin repeat domain-containing protein SOWAHA | SOWAHA | 1:1 | Block_3 | 2.3342 | 13.8526 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O95416 | Transcription factor SOX-14 | SOX14 | 1:1 | Block_3 | -0.27106 | 8.3392 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P48431 | Transcription factor SOX-2 | SOX2 | 1:1 | Block_3 | 1.9535 | 14.7722 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P48436 | Transcription factor SOX-9 | SOX9 | 1:1 | Block_3 | 1.04 | 12.5212 | TRUE | NA | 5.6681 | NA | NA | 1.3547 | NA | NA | 1.7812
Q15506 | Sperm surface protein Sp17 | SPA17 | 1:1 | Block_3 | -0.90853 | 12.2111 | FALSE | NA | NA | NA | -0.63316 | 1.3924 | -0.20792 | 1.4576 | 2.9889
Q96QH8 | Sperm acrosome-associated protein 5 | SPACA5 | 1:1 | Block_3 | 0.61109 | 9.4191 | TRUE | 1.4568 | 1.7746 | NA | NA | 0.82731 | NA | NA | 1.2246
Q07617 | Sperm-associated antigen 1 | SPAG1 | 1:1 | Block_3 | 0.63915 | 12.3352 | FALSE | NA | NA | NA | NA | 0.98561 | NA | 0.7952 | 2.1972
Q9UBP0 | Spastin | SPAST | 1:1 | Block_3 | 0.40255 | 9.8881 | TRUE | NA | NA | NA | 0.86315 | 3.0971 | 0.52756 | 3.0436 | 5.4091
Q8TC71 | Mitochondria-eating protein | SPATA18 | 1:1 | Block_3 | 0.33867 | 11.9323 | TRUE | NA | NA | NA | NA | 1.1349 | NA | NA | 1.1189
Q8IUW3 | Spermatogenesis-associated protein 2-like protein | SPATA2L | 1:1 | Block_3 | -2.76 | 6.9287 | FALSE | 35.1724 | 19.3503 | -1.8488 | -0.64316 | 0.46756 | -1.6681 | -0.20181 | 1.6678
P0C7L1 | Serine protease inhibitor Kazal-type 8 | SPINK8 | 1:1 | Block_3 | 1.5 | 10.8502 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WWL2 | Protein spire homolog 2 | SPIRE2 | 1:1 | Block_3 | 0.40887 | 6.6101 | TRUE | 11.4836 | 4.2879 | NA | NA | 1.2627 | NA | NA | 1.7433
Q7Z698 | Sprouty-related, EVH1 domain-containing protein 2 | SPRED2 | 1:1 | Block_3 | 2.7029 | 10.768 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.6593
Q9H741 | SREBP regulating gene protein | SPRING1 | 1:1 | Block_3 | 0.001 | 7.6263 | FALSE | 35.8528 | 10.7322 | NA | 0.28263 | 0.91604 | NA | 0.66596 | 1.6078
P22528 | Cornifin-B | SPRR1B | 1:1 | Block_3 | 0.80349 | 7.6245 | TRUE | NA | NA | NA | NA | 2.1671 | NA | NA | NA
O43597 | Protein sprouty homolog 2 | SPRY2 | 1:1 | Block_3 | 0.57903 | 9.0376 | FALSE | NA | NA | NA | NA | 2.1179 | NA | 1.5918 | 3.2483
O15020 | Spectrin beta chain, non-erythrocytic 2 | SPTBN2 | 1:1 | Block_3 | 0.8628 | 10.9058 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5842
Q9C0H9 | SRC kinase signaling inhibitor 1 | SRCIN1 | 1:1 | Block_3 | 0.33726 | 9.9159 | TRUE | 8.4 | 2.9021 | NA | NA | NA | NA | NA | 0.79279
O75044 | SLIT-ROBO Rho GTPase-activating protein 2 | SRGAP2 | 1:1 | Block_3 | 1.7538 | 10.3687 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P05455 | Lupus La protein | SSB | 1:1 | Block_3 | 0.5 | 8.46 | FALSE | NA | NA | NA | NA | 1.4045 | NA | 0.63399 | 2.6429
Q04837 | Single-stranded DNA-binding protein, mitochondrial | SSBP1 | 1:1 | Block_3 | 1.2898 | NA | NA | 9.2195 | 4.4855 | NA | NA | 1.3649 | NA | NA | NA
Q8TE77 | Protein phosphatase Slingshot homolog 3 | SSH3 | 1:1 | Block_3 | 0.55228 | 7.8133 | FALSE | NA | 0.94554 | NA | 0.6275 | 1.1977 | NA | 1.1246 | 2.0678
Q92185 | Alpha-N-acetylneuraminide alpha-2,8-sialyltransferase | ST8SIA1 | 1:1 | Block_3 | 1.0675 | 11.6501 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O95793 | Double-stranded RNA-binding protein Staufen homolog 1 | STAU1 | 1:1 | Block_3 | 0.19 | 10.8548 | FALSE | NA | 4.5749 | NA | NA | 2.6142 | NA | 1.1824 | 3.3126
Q687X5 | Metalloreductase STEAP4 | STEAP4 | 1:1 | Block_3 | 0.95009 | 12.7114 | TRUE | 12.3814 | NA | NA | NA | 1.4503 | NA | NA | 1.9894
Q9Y6E0 | Serine/threonine-protein kinase 24 | STK24 | 1:1 | Block_3 | -2.7144 | 7.0579 | TRUE | NA | NA | NA | NA | -1.8111 | NA | -1.877 | -0.48364
Q9UJZ1 | Stomatin-like protein 2, mitochondrial | STOML2 | 1:1 | Block_3 | 1.6973 | 11.1557 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O14662 | Syntaxin-16 | STX16 | 1:1 | Block_3 | 1.95 | 10.2709 | FALSE | 12.327 | 8.4093 | NA | NA | NA | NA | NA | 2.5782
P61266 | Syntaxin-1B | STX1B | 1:1 | Block_3 | -3.3117 | 8.0228 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q13190 | Syntaxin-5 | STX5 | 1:1 | Block_3 | 1.871 | 3.0334 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.3592
O43752 | Syntaxin-6 | STX6 | 1:1 | Block_3 | 1.8726 | 12.4952 | FALSE | NA | NA | NA | NA | 1.9575 | NA | NA | 3.33
O00186 | Syntaxin-binding protein 3 | STXBP3 | 1:1 | Block_3 | 1.8479 | 8.717 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.337
Q9Y2K9 | Syntaxin-binding protein 5-like | STXBP5L | 1:1 | Block_3 | -2.0016 | 7.4 | FALSE | NA | NA | NA | -1.9804 | 0.28346 | NA | NA | -1.1097
Q9Y2Z0 | Protein SGT1 homolog | SUGT1 | 1:1 | Block_3 | 0.61514 | 5.6093 | FALSE | NA | NA | NA | 1.4704 | 4.6686 | 1.539 | 4.851 | 6.0197
Q8NBK3 | Formylglycine-generating enzyme | SUMF1 | 1:1 | Block_3 | 1.7782 | 6.1728 | FALSE | NA | NA | NA | NA | 2.2789 | NA | NA | NA
Q15527 | Surfeit locus protein 2 | SURF2 | 1:1 | Block_3 | 0.52472 | 12.1672 | TRUE | NA | NA | NA | NA | 1.2382 | NA | NA | 2.1449
Q6UWL2 | Sushi domain-containing protein 1 | SUSD1 | 1:1 | Block_3 | 0.26144 | 8.3544 | TRUE | 18.2567 | 15.4212 | NA | 0.86306 | 2.0938 | 0.39413 | 1.6699 | 2.9866
A8MT33 | Synaptonemal complex central element protein 1-like | SYCE1L | 1:1 | Block_3 | 0.83451 | 12.8368 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.84696
O95926 | Pre-mRNA-splicing factor SYF2 | SYF2 | 1:1 | Block_3 | 0.67611 | 13.6609 | FALSE | NA | 1.2353 | NA | NA | 1.5056 | NA | 1.3922 | 3.7596
Q96PV0 | Ras/Rap GTPase-activating protein SynGAP | SYNGAP1 | 1:1 | Block_3 | 1.3083 | 14.1252 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O15061 | Synemin | SYNM | 1:1 | Block_3 | -0.46035 | 10.7537 | FALSE | 7.5068 | 19.6976 | -0.44336 | 0.6094 | 1.493 | 0.52178 | 1.2844 | 2.7536
Q9NYJ8 | TGF-beta-activated kinase 1 and MAP3K7-binding protein 2 | TAB2 | 1:1 | Block_3 | 0.76725 | 7.4694 | FALSE | NA | NA | NA | NA | 3.3621 | NA | 2.3028 | 4.4615
O75529 | TAF5-like RNA polymerase II p300/CBP-associated factor-associated factor 65 kDa subunit 5L | TAF5L | 1:1 | Block_3 | 1.4807 | 8.9202 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UI15 | Transgelin-3 | TAGLN3 | 1:1 | Block_3 | 1.0259 | 6.8307 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q92844 | TRAF family member-associated NF-kappa-B activator | TANK | 1:1 | Block_3 | 2.825 | 12.3869 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.0176
Q03518 | Antigen peptide transporter 1 | TAP1 | 1:1 | Block_3 | 2.3969 | 8.9915 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q15633 | RISC-loading complex subunit TARBP2 | TARBP2 | 1:1 | Block_3 | 0.28206 | 9.7649 | TRUE | NA | NA | NA | 0.49061 | 2.7122 | 1.0589 | 2.6585 | 4.7977
B6A8C7 | T-cell-interacting, activating receptor on myeloid cells protein 1 | TARM1 | 1:1 | Block_3 | 0.29 | 10.481 | FALSE | 15.2096 | 15.8285 | NA | NA | 0.89208 | NA | NA | 0.75936
Q9HA65 | TBC1 domain family member 17 | TBC1D17 | 1:1 | Block_3 | 0.68053 | 8.6082 | TRUE | NA | NA | NA | 0.85719 | 1.7384 | NA | 1.2796 | 2.2574
Q9BYX2 | TBC1 domain family member 2A | TBC1D2 | 1:1 | Block_3 | 1.2994 | 10.7626 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.3594
O60907 | F-box-like/WD repeat-containing protein TBL1X | TBL1X | 1:1 | Block_3 | 0.17 | 9.8713 | TRUE | 5.7017 | 21.9887 | NA | 1.2316 | 3.0679 | 0.74685 | 3.2776 | 4.6626
Q16650 | T-box brain protein 1 | TBR1 | 1:1 | Block_3 | 1.071 | 11.8437 | FALSE | NA | 13.4821 | NA | NA | 1.1051 | NA | NA | 1.1965
Q96MN5 | Transcription elongation factor A N-terminal and central domain-containing protein 2 | TCEANC2 | 1:1 | Block_3 | 0.44775 | 3.6691 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WWU5 | T-complex protein 11 homolog | TCP11 | 1:1 | Block_3 | 1.0853 | 9.2043 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P13385 | Teratocarcinoma-derived growth factor 1 | TDGF1 | 1:1 | Block_3 | -0.77385 | 10.1805 | TRUE | 15.3949 | 17.2985 | NA | NA | 3.5345 | NA | -0.24203 | 5.2143
P48775 | Tryptophan 2,3-dioxygenase | TDO2 | 1:1 | Block_3 | 1.2132 | 10.9605 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y2W6 | Tudor and KH domain-containing protein | TDRKH | 1:1 | Block_3 | 1.6081 | 12.2032 | TRUE | NA | NA | NA | NA | 3.3661 | NA | 2.6466 | 4.6907
O75443 | Alpha-tectorin | TECTA | 1:1 | Block_3 | 1.7898 | 11.5968 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q10587 | Thyrotroph embryonic factor | TEF | 1:1 | Block_3 | 0.001 | 12.391 | FALSE | 25.7813 | 25.5115 | NA | 0.366 | 1.4023 | NA | 0.38436 | 1.3418
Q8NDF8 | Terminal nucleotidyltransferase 4B | TENT4B | 1:1 | Block_3 | 1.3154 | 10.7294 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P54274 | Telomeric repeat-binding factor 1 | TERF1 | 1:1 | Block_3 | 1.9065 | 11.8998 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.1521
Q9NYB0 | Telomeric repeat-binding factor 2-interacting protein 1 | TERF2IP | 1:1 | Block_3 | 0.001 | 9.4417 | TRUE | NA | NA | NA | 0.13193 | 1.8424 | 0.046542 | 2.1079 | 3.6006
Q9UGI8 | Testin | TES | 1:1 | Block_3 | 0.12226 | 10.0714 | TRUE | 10.0706 | 8.6505 | NA | 0.2 | 0.7784 | NA | 0.63152 | 4.4732
Q6N021 | Methylcytosine dioxygenase TET2 | TET2 | 1:1 | Block_3 | 1.1988 | 12.271 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.6557
O43247 | Testis-expressed protein 33 | TEX33 | 1:1 | Block_3 | 1.8341 | 14.7451 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5T0J7 | Testis-expressed protein 35 | TEX35 | 1:1 | Block_3 | 1.1717 | 15.5415 | TRUE | 11.9395 | 19.4314 | NA | NA | 1.3964 | NA | NA | 1.8563
Q00059 | Transcription factor A, mitochondrial | TFAM | 1:1 | Block_3 | 0.79549 | 2.8369 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P05549 | Transcription factor AP-2-alpha | TFAP2A | 1:1 | Block_3 | 2.6295 | 10.9796 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O43548 | Protein-glutamine gamma-glutamyltransferase 5 | TGM5 | 1:1 | Block_3 | 0.73986 | 11.7718 | TRUE | NA | NA | NA | NA | 2.1933 | NA | NA | NA
Q9P2Z0 | THAP domain-containing protein 10 | THAP10 | 1:1 | Block_3 | 0.26004 | 11.7214 | FALSE | NA | NA | NA | NA | 0.37421 | NA | NA | 0.29143
O43422 | 52 kDa repressor of the inhibitor of the protein kinase | THAP12 | 1:1 | Block_3 | 0.77016 | 10.4566 | TRUE | NA | NA | NA | NA | 0.80593 | NA | NA | NA
Q8WTV1 | THAP domain-containing protein 3 | THAP3 | 1:1 | Block_3 | 1.9576 | 15.2539 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P0DJG4 | Testicular haploid expressed gene protein-like | THEGL | 1:1 | Block_3 | -0.74418 | 8.4335 | TRUE | 3.717 | 5.5861 | NA | NA | -0.1713 | NA | NA | -0.39947
Q9NS62 | Thrombospondin type-1 domain-containing protein 1 | THSD1 | 1:1 | Block_3 | 1.6047 | 9.6341 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.9025
Q9P016 | Thymocyte nuclear protein 1 | THYN1 | 1:1 | Block_3 | 1.8135 | 5.0096 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P31483 | Cytotoxic granule associated RNA binding protein TIA1 | TIA1 | 1:1 | Block_3 | 2.1299 | 4.7344 | FALSE | NA | NA | NA | NA | 4.4922 | NA | 4.4926 | 7.165
Q495A1 | T-cell immunoreceptor with Ig and ITIM domains | TIGIT | 1:1 | Block_3 | 1.1571 | 10.4702 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.6537
O95049 | Tight junction protein ZO-3 | TJP3 | 1:1 | Block_3 | 0.9 | 12.1055 | FALSE | NA | NA | NA | NA | 1.3818 | NA | NA | 1.339
P04183 | Thymidine kinase, cytosolic | TK1 | 1:1 | Block_3 | 0.9 | 6.8143 | FALSE | NA | 5.1604 | NA | NA | 0.94958 | NA | NA | 1.2805
O60603 | Toll-like receptor 2 | TLR2 | 1:1 | Block_3 | 0.93362 | 10.7416 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O00206 | Toll-like receptor 4 | TLR4 | 1:1 | Block_3 | 0.76451 | 6.9701 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.0035
Q96EY4 | Translation machinery-associated protein 16 | TMA16 | 1:1 | Block_3 | 0.36214 | 13.0962 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.2228
Q8N6Q1 | Transmembrane and coiled-coil domain-containing protein 5A | TMCO5A | 1:1 | Block_3 | 0.7272 | 12.6187 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.0134
Q13445 | Transmembrane emp24 domain-containing protein 1 | TMED1 | 1:1 | Block_3 | 2.6662 | 7.8377 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P49755 | Transmembrane emp24 domain-containing protein 10 | TMED10 | 1:1 | Block_3 | 2.83 | 11.0536 | TRUE | NA | NA | NA | NA | 2.9099 | NA | NA | NA
Q7Z7H5 | Transmembrane emp24 domain-containing protein 4 | TMED4 | 1:1 | Block_3 | 0.87053 | 6.6653 | FALSE | NA | 8.7 | NA | NA | NA | NA | NA | NA
Q6UXF1 | Transmembrane protein 108 | TMEM108 | 1:1 | Block_3 | 0.23847 | 11.058 | FALSE | NA | NA | NA | NA | 0.41143 | NA | NA | 0.33391
Q9NZQ9 | Tropomodulin-4 | TMOD4 | 1:1 | Block_3 | 0.8 | 10.7245 | FALSE | 12.7164 | 12.9328 | NA | NA | 1.1365 | NA | NA | 1.5234
Q86T26 | Transmembrane protease serine 11B | TMPRSS11B | 1:1 | Block_3 | 0.69371 | 9.9293 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P01375 | Tumor necrosis factor | TNF | 1:1 | Block_3 | 0.14095 | 12.581 | FALSE | 16.937 | 12.2658 | NA | 0.74113 | 1.6564 | NA | 1.1203 | 2.334
O95379 | Tumor necrosis factor alpha-induced protein 8 | TNFAIP8 | 1:1 | Block_3 | 1.1127 | 8.8427 | FALSE | NA | NA | NA | NA | 1.1874 | NA | NA | 2.2906
Q96RJ3 | Tumor necrosis factor receptor superfamily member 13C | TNFRSF13C | 1:1 | Block_3 | -0.56 | 8.1001 | TRUE | 16.3777 | 13.001 | -0.39519 | 0.31015 | 1.1756 | -0.52375 | 0.73584 | 1.6217
O95150 | Tumor necrosis factor ligand superfamily member 15 | TNFSF15 | 1:1 | Block_3 | 0.3278 | 8.3354 | TRUE | 17.7159 | 22.5655 | NA | 0.98004 | 1.5927 | NA | 1.0602 | 2.2316
P41273 | Tumor necrosis factor ligand superfamily member 9 | TNFSF9 | 1:1 | Block_3 | 1.92 | 10.6723 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P19429 | Troponin I, cardiac muscle | TNNI3 | 1:1 | Block_3 | 1.795 | 7.8805 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O15417 | Trinucleotide repeat-containing gene 18 protein | TNRC18 | 1:1 | Block_3 | 1.0911 | 9.4089 | TRUE | NA | NA | NA | NA | 1.2343 | NA | NA | 1.4184
Q6ZUX3 | TOG array regulator of axonemal microtubules protein 2 | TOGARAM2 | 1:1 | Block_3 | 0.42704 | 10.5975 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O75674 | TOM1-like protein 1 | TOM1L1 | 1:1 | Block_3 | 1.2315 | 9.3107 | TRUE | NA | NA | NA | NA | 2.15 | NA | 1.3268 | 2.5006
Q969P6 | DNA topoisomerase I, mitochondrial | TOP1MT | 1:1 | Block_3 | 1.191 | 10.0847 | FALSE | 8.1433 | 7.2248 | NA | NA | NA | NA | NA | 1.3974
O94900 | Thymocyte selection-associated high mobility group box protein TOX | TOX | 1:1 | Block_3 | 0.6585 | 11.0883 | TRUE | NA | NA | NA | NA | 1.5558 | NA | 1.9952 | 3.8222
O15405 | TOX high mobility group box family member 3 | TOX3 | 1:1 | Block_3 | 0.87238 | 10.226 | TRUE | NA | NA | NA | NA | 0.8916 | NA | NA | 2.004
P04637 | Cellular tumor antigen p53 | TP53 | 1:1 | Block_3 | 2.1767 | 10.5051 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.9851
Q12888 | TP53-binding protein 1 | TP53BP1 | 1:1 | Block_3 | 0.841 | 8.8751 | FALSE | 11.3672 | 10.2717 | NA | NA | 1.0754 | NA | NA | 2.6467
Q8IXH6 | Tumor protein p53-inducible nuclear protein 2 | TP53INP2 | 1:1 | Block_3 | 0.16061 | 10.342 | TRUE | 2.2901 | 2.0327 | NA | NA | 0.65261 | NA | NA | 0.93114
O15350 | Tumor protein p73 | TP73 | 1:1 | Block_3 | 2.8949 | 13.0705 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P0DKB5 | Trophoblast glycoprotein-like | TPBGL | 1:1 | Block_3 | 1.9152 | 3.6166 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P06753 | Tropomyosin alpha-3 chain | TPM3 | 1:1 | Block_3 | 1.3666 | 10.4257 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.4729
P51580 | Thiopurine S-methyltransferase | TPMT | 1:1 | Block_3 | 1.2219 | 9.1034 | TRUE | NA | NA | NA | NA | 2.4674 | NA | 2.5249 | 4.3591
P59282 | Tubulin polymerization-promoting protein family member 2 | TPPP2 | 1:1 | Block_3 | 1.1401 | 10.9125 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BW30 | Tubulin polymerization-promoting protein family member 3 | TPPP3 | 1:1 | Block_3 | -1.23 | 8.2633 | TRUE | 20.2423 | 14.7071 | -1.2276 | -0.11561 | 0.75123 | -0.90899 | 0.092259 | 1.443
P12270 | Nucleoprotein TPR | TPR | 1:1 | Block_3 | 0.43608 | 9.6433 | TRUE | NA | NA | NA | NA | 0.78811 | NA | NA | 1.9996
Q9Y3C4 | EKC/KEOPS complex subunit TPRKB | TPRKB | 1:1 | Block_3 | 1.1871 | 8.4033 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q4KMQ1 | Taperin | TPRN | 1:1 | Block_3 | 2.0182 | 12.013 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BZJ3 | Tryptase delta | TPSD1 | 1:1 | Block_3 | 0.34339 | 10.5478 | TRUE | 13.4499 | 9.9184 | NA | NA | 0.62257 | NA | NA | 1.3628
P13693 | Translationally-controlled tumor protein | TPT1 | 1:1 | Block_3 | 2.5571 | 10.1846 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.4338
Q13114 | TNF receptor-associated factor 3 | TRAF3 | 1:1 | Block_3 | 2.8316 | 8.2078 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O43734 | E3 ubiquitin ligase TRAF3IP2 | TRAF3IP2 | 1:1 | Block_3 | 1.0161 | 11.0704 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UPV9 | Trafficking kinesin-binding protein 1 | TRAK1 | 1:1 | Block_3 | 0.77384 | 12.665 | FALSE | NA | NA | NA | NA | 1.0334 | NA | 1.0445 | 3.3548
Q9Y296 | Trafficking protein particle complex subunit 4 | TRAPPC4 | 1:1 | Block_3 | 1.9472 | 7.9442 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O14717 | tRNA (cytosine(38)-C(5))-methyltransferase | TRDMT1 | 1:1 | Block_3 | -0.30982 | 8.8302 | FALSE | NA | NA | NA | NA | -0.021443 | NA | NA | -0.033147
O15164 | Transcription intermediary factor 1-alpha | TRIM24 | 1:1 | Block_3 | 1.2652 | 10.7315 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q12899 | Tripartite motif-containing protein 26 | TRIM26 | 1:1 | Block_3 | 1.9154 | 9.9878 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.1144
Q9HCM9 | E3 ubiquitin-protein ligase TRIM39 | TRIM39 | 1:1 | Block_3 | 2.0085 | 12.9441 | TRUE | NA | NA | NA | NA | 3.2097 | NA | NA | 3.0565
Q6P9F5 | E3 ubiquitin ligase TRIM40 | TRIM40 | 1:1 | Block_3 | 0.74397 | 7.7381 | FALSE | NA | NA | NA | NA | NA | NA | NA | 0.93689
Q86WT6 | E3 ubiquitin-protein ligase TRIM69 | TRIM69 | 1:1 | Block_3 | 0.79745 | 8.9206 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NET8 | Transient receptor potential cation channel subfamily V member 3 | TRPV3 | 1:1 | Block_3 | 0.87946 | 11.5527 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q92574 | Hamartin | TSC1 | 1:1 | Block_3 | 0.9 | 4.3654 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.5826
Q969D9 | Thymic stromal lymphopoietin | TSLP | 1:1 | Block_3 | 0.45645 | 13.1239 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O95858 | Tetraspanin-15 | TSPAN15 | 1:1 | Block_3 | 2.0143 | 10.7654 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P41732 | Tetraspanin-7 | TSPAN7 | 1:1 | Block_3 | 2.7963 | 8.2755 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q16762 | Thiosulfate sulfurtransferase | TST | 1:1 | Block_3 | 0.84996 | 3.7083 | FALSE | NA | NA | NA | NA | 1.3914 | NA | 0.93665 | 2.1727
Q5TCY1 | Tau-tubulin kinase 1 | TTBK1 | 1:1 | Block_3 | -0.94476 | 8.4372 | FALSE | NA | 5.2873 | NA | NA | -0.48656 | NA | NA | 0.65833
Q99614 | Tetratricopeptide repeat protein 1 | TTC1 | 1:1 | Block_3 | 0.84889 | 6.965 | FALSE | NA | NA | NA | NA | 1.1445 | NA | 0.96763 | 2.2499
Q5I0X7 | Tetratricopeptide repeat protein 32 | TTC32 | 1:1 | Block_3 | 1.1755 | 10.6296 | TRUE | 4.309 | 11.7773 | NA | NA | 1.8771 | NA | NA | 5.0806
Q9UNY4 | Transcription termination factor 2 | TTF2 | 1:1 | Block_3 | 1.9151 | 12.0261 | TRUE | NA | NA | NA | NA | 2.364 | NA | 2.3552 | 4.1964
Q8WZ42 | Titin | TTN | 1:1 | Block_3 | 1.5909 | 7.5765 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q13509 | Tubulin beta-3 chain | TUBB3 | 1:1 | Block_3 | 2.2025 | 6.9385 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q3ZCM7 | Tubulin beta-8 chain | TUBB8 | 1:1 | Block_3 | 1.9081 | 8.1595 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P42681 | Tyrosine-protein kinase TXK | TXK | 1:1 | Block_3 | -0.68362 | 8.1168 | FALSE | 8.8193 | 3.6495 | NA | NA | NA | NA | NA | -0.28729
Q8N3L3 | Beta-taxilin | TXLNB | 1:1 | Block_3 | 0.3952 | 11.208 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O43396 | Thioredoxin-like protein 1 | TXNL1 | 1:1 | Block_3 | 1.0446 | 7.4075 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P04818 | Thymidylate synthase | TYMS | 1:1 | Block_3 | -0.36944 | 10.1373 | FALSE | NA | 7.3881 | NA | NA | 0.02182 | NA | 0.017799 | 1.9423
Q9H832 | Ubiquitin-conjugating enzyme E2 Z | UBE2Z | 1:1 | Block_3 | 1.312 | 12.325 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.5595
Q8N7F7 | Ubiquitin-like protein 4B | UBL4B | 1:1 | Block_3 | 1.1101 | 11.4733 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.2799
Q9H347 | Ubiquilin-3 | UBQLN3 | 1:1 | Block_3 | 1.0605 | 8.7739 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.3008
Q8WUN7 | Ubiquitin domain-containing protein 2 | UBTD2 | 1:1 | Block_3 | 0.48458 | 11.3274 | TRUE | 10.8474 | 10.5482 | NA | NA | 0.8138 | NA | 0.84886 | 1.8324
P17480 | Nucleolar transcription factor 1 | UBTF | 1:1 | Block_3 | 0.75493 | 10.7084 | TRUE | NA | NA | NA | NA | 0.99488 | NA | NA | 1.2873
P0CB47 | Upstream-binding factor 1-like protein 1 | UBTFL1 | 1:1 | Block_3 | 0.19637 | 9.5681 | TRUE | NA | NA | NA | NA | 0.34803 | NA | NA | NA
O60701 | UDP-glucose 6-dehydrogenase | UGDH | 1:1 | Block_3 | 1.3241 | 5.1937 | FALSE | NA | 28.6032 | NA | NA | 1.3971 | NA | NA | 1.5522
Q9NYU2 | UDP-glucose:glycoprotein glucosyltransferase 1 | UGGT1 | 1:1 | Block_3 | 0.61708 | 8.0681 | FALSE | NA | NA | NA | NA | NA | NA | NA | 0.62425
C9J7I0 | UBAP1-MVB12-associated | UMAD1 | 1:1 | Block_3 | 0.22682 | 11.8232 | TRUE | 13.4335 | 23.3111 | 0.23173 | 1.6405 | 3.4271 | 1.9944 | 3.9503 | 5.6516
Q6UXZ4 | Netrin receptor UNC5D | UNC5D | 1:1 | Block_3 | 1.086 | 7.6163 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P2D8 | Protein unc-79 homolog | UNC79 | 1:1 | Block_3 | 1.3697 | 11.6609 | FALSE | NA | NA | NA | NA | 1.4187 | NA | NA | 2.5276
Q9H1J1 | Regulator of nonsense transcripts 3A | UPF3A | 1:1 | Block_3 | 2.0135 | 10.147 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.1845
B0FP48 | Uroplakin-3b-like protein 1 | UPK3BL1 | 1:1 | Block_3 | -1.3735 | 14.7264 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96K76 | Ubiquitin carboxyl-terminal hydrolase 47 | USP47 | 1:1 | Block_3 | -0.34137 | 11.4343 | FALSE | 21.292 | 11.1151 | NA | NA | -0.21755 | NA | NA | 0.89995
O95183 | Vesicle-associated membrane protein 5 | VAMP5 | 1:1 | Block_3 | 1.0415 | 13.9243 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.3638
Q9H867 | Protein N-lysine methyltransferase METTL21D | VCPKMT | 1:1 | Block_3 | 0.8 | 6.9804 | FALSE | 10.7975 | 26.0846 | NA | NA | 1.513 | NA | NA | 1.408
P32241 | Vasoactive intestinal polypeptide receptor 1 | VIPR1 | 1:1 | Block_3 | 2.2354 | 12.249 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UK41 | Vacuolar protein sorting-associated protein 28 homolog | VPS28 | 1:1 | Block_3 | 1.3207 | 9.4078 | FALSE | NA | NA | NA | NA | 1.9049 | NA | 1.5116 | 2.9134
Q6UX27 | V-set and transmembrane domain-containing protein 1 | VSTM1 | 1:1 | Block_3 | 0.43214 | 6.9166 | TRUE | 1.5632 | NA | NA | NA | 0.68543 | NA | 0.57959 | 1.7511
A6NLU5 | V-set and transmembrane domain-containing protein 2B | VSTM2B | 1:1 | Block_3 | 0.81 | 6.8226 | TRUE | 10.0945 | 16.0197 | NA | NA | 0.97906 | NA | NA | NA
Q96N03 | V-set and transmembrane domain-containing protein 2-like protein | VSTM2L | 1:1 | Block_3 | 0.54214 | 8.2903 | TRUE | 1.1937 | 1.664 | NA | NA | 1.2206 | NA | NA | 1.0972
Q7Z7D3 | V-set domain-containing T-cell activation inhibitor 1 | VTCN1 | 1:1 | Block_3 | 1.6086 | 13.5029 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O00534 | von Willebrand factor A domain-containing protein 5A | VWA5A | 1:1 | Block_3 | 1.0762 | 10.201 | FALSE | NA | NA | NA | NA | 1.1412 | NA | NA | 1.5372
P42768 | Actin nucleation-promoting factor WAS | WAS | 1:1 | Block_3 | 2.4619 | 10.8971 | TRUE | NA | NA | NA | NA | 4.2254 | NA | NA | 5.7232
O00401 | Actin nucleation-promoting factor WASL | WASL | 1:1 | Block_3 | 1.2 | 8.7594 | FALSE | 26.9248 | 28.6083 | NA | NA | NA | NA | NA | 3.9215
Q8TF30 | WASP homolog-associated protein with actin, membranes and microtubules | WHAMM | 1:1 | Block_3 | 0.527 | 12.2259 | TRUE | 21.5204 | 21.0744 | NA | NA | 1.2279 | NA | 0.83881 | 2.2415
O14904 | Protein Wnt-9a | WNT9A | 1:1 | Block_3 | 0.59765 | 11.7904 | TRUE | 7.8966 | 10.104 | NA | NA | 1.0614 | NA | 0.80273 | 1.6857
Q2TBF2 | WSC domain-containing protein 2 | WSCD2 | 1:1 | Block_3 | 2.0981 | 9.4178 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P98170 | E3 ubiquitin-protein ligase XIAP | XIAP | 1:1 | Block_3 | 1.1686 | 6.5865 | FALSE | 4.9404 | NA | NA | NA | 2.48 | NA | 2.0502 | 3.88
Q13426 | DNA repair protein XRCC4 | XRCC4 | 1:1 | Block_3 | 1.3382 | 8.5058 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.9472
Q8N7P7 | Uncharacterized protein FLJ40521 | YH007 | 1:1 | Block_3 | 0.5 | 14.1911 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q7Z739 | YTH domain-containing family protein 3 | YTHDF3 | 1:1 | Block_3 | -0.73552 | 6.8838 | TRUE | 9.4734 | 11.3288 | -0.73506 | 0.76298 | 4.3671 | 1.4398 | 4.5078 | 5.9214
P25490 | Transcriptional repressor protein YY1 | YY1 | 1:1 | Block_3 | 3.851 | 8.935 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P86452 | Zinc finger BED domain-containing protein 6 | ZBED6 | 1:1 | Block_3 | 0.91492 | 11.6303 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q05516 | Zinc finger and BTB domain-containing protein 16 | ZBTB16 | 1:1 | Block_3 | 1.8618 | 12.5836 | TRUE | NA | NA | NA | NA | 2.3413 | NA | 2.2784 | 4.4469
Q96GY0 | Zinc finger C2HC domain-containing protein 1A | ZC2HC1A | 1:1 | Block_3 | -0.1 | 11.3829 | TRUE | NA | NA | NA | NA | 0.84669 | NA | 0.6035 | 1.6409
Q5D1E8 | Endoribonuclease ZC3H12A | ZC3H12A | 1:1 | Block_3 | 0.89565 | 11.8629 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.5958
Q6NZY4 | Zinc finger CCHC domain-containing protein 8 | ZCCHC8 | 1:1 | Block_3 | 1.5098 | 9.0519 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6FIF0 | AN1-type zinc finger protein 6 | ZFAND6 | 1:1 | Block_3 | 1.3304 | 10.1694 | TRUE | NA | NA | NA | NA | 1.4744 | NA | NA | 2.3333
Q9BRR0 | Zinc finger protein with KRAB and SCAN domains 3 | ZKSCAN3 | 1:1 | Block_3 | 0.45828 | 10.1732 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H091 | Zinc finger MYND domain-containing protein 15 | ZMYND15 | 1:1 | Block_3 | -2.4718 | 9.5895 | FALSE | NA | 5.6835 | NA | NA | -2.2377 | NA | NA | -1.4452
O43670 | BUB3-interacting and GLEBS motif-containing protein ZNF207 | ZNF207 | 1:1 | Block_3 | 1.1941 | 11.1413 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.0002
Q96GC6 | Neurotrophin receptor-interacting factor homolog | ZNF274 | 1:1 | Block_3 | -0.45171 | 8.9359 | TRUE | 8.037 | 9.1099 | NA | -0.4377 | 1.1709 | NA | NA | 1.2857
Q2M1K9 | Zinc finger protein 423 | ZNF423 | 1:1 | Block_3 | -0.31112 | 12.6032 | TRUE | 11.8542 | 4.7521 | NA | NA | 0.35532 | NA | -0.24005 | 1.2283
Q7Z3I7 | Zinc finger protein 572 | ZNF572 | 1:1 | Block_3 | -1.6252 | 10.8961 | TRUE | 6.2769 | 5.4381 | NA | NA | -1.2163 | NA | NA | -0.7288
Q9BS31 | Zinc finger protein 649 | ZNF649 | 1:1 | Block_3 | 0.5647 | 8.1048 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P51815 | Zinc finger protein 75D | ZNF75D | 1:1 | Block_3 | 1.5929 | 12.4208 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96NB3 | Zinc finger protein 830 | ZNF830 | 1:1 | Block_3 | 1.0849 | 9.4568 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.4434
Q8WWF5 | E3 ubiquitin-protein ligase ZNRF4 | ZNRF4 | 1:1 | Block_3 | 4.6564 | 13.3316 | FALSE | NA | NA | NA | NA | 5.6302 | NA | NA | 5.7352
Q12836 | Zona pellucida sperm-binding protein 4 | ZP4 | 1:1 | Block_3 | 0.31352 | 10.0189 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O75312 | Zinc finger protein ZPR1 | ZPR1 | 1:1 | Block_3 | 1.5 | 10.0159 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NRG9 | Aladin | AAAS | 1:1 | Block_2 | 1.8414 | 11.3379 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6PD74 | Alpha- and gamma-adaptin-binding protein p34 | AAGAB | 1:1 | Block_2 | 1.2264 | 12.5683 | TRUE | NA | NA | NA | NA | 3.5161 | 1.2408 | 3.692 | 5.5117
Q4L235 | Beta-alanine-activating enzyme | AASDH | 1:1 | Block_2 | 2.1906 | 10.0867 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.3436
Q2M3G0 | ATP-binding cassette sub-family B member 5 | ABCB5 | 1:1 | Block_2 | 1.2714 | 12.5841 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4674
O95255 | ATP-binding cassette sub-family C member 6 | ABCC6 | 1:1 | Block_2 | 0.45799 | 16.6378 | TRUE | NA | NA | NA | NA | 2.7137 | NA | NA | 1.7359
Q09428 | ATP-binding cassette sub-family C member 8 | ABCC8 | 1:1 | Block_2 | 1.7409 | 13.1903 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.6896
Q9UG63 | ATP-binding cassette sub-family F member 2 | ABCF2 | 1:1 | Block_2 | 1.3494 | 12.6172 | FALSE | 19.6387 | NA | NA | NA | NA | NA | 1.4721 | 2.751
O94929 | Actin-binding LIM protein 3 | ABLIM3 | 1:1 | Block_2 | 1.104 | 12.0459 | TRUE | NA | NA | NA | 1.2331 | 4.3351 | 1.3075 | 3.6243 | 5.7471
Q8N0Z2 | Actin-binding Rho-activating protein | ABRA | 1:1 | Block_2 | 1.969 | 11.9009 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N6N7 | Acyl-CoA-binding domain-containing protein 7 | ACBD7 | 1:1 | Block_2 | 0.17997 | 7.4247 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UKV3 | Apoptotic chromatin condensation inducer in the nucleus | ACIN1 | 1:1 | Block_2 | 1.3562 | 6.9037 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WYK0 | Acetyl-coenzyme A thioesterase | ACOT12 | 1:1 | Block_2 | 1.9246 | 13.5527 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q08AH1 | Acyl-coenzyme A synthetase ACSM1, mitochondrial | ACSM1 | 1:1 | Block_2 | 1.2401 | 10.3811 | FALSE | NA | NA | NA | NA | 1.4686 | NA | NA | 2.0676
Q9H9F9 | Actin-related protein 5 | ACTR5 | 1:1 | Block_2 | 1.7565 | 9.5016 | FALSE | 5.4821 | NA | NA | NA | NA | NA | NA | NA
O60266 | Adenylate cyclase type 3 | ADCY3 | 1:1 | Block_2 | 2.2485 | 11.276 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.4303
O95622 | Adenylate cyclase type 5 | ADCY5 | 1:1 | Block_2 | 1.1654 | 9.079 | TRUE | NA | NA | NA | NA | 1.2377 | NA | NA | 1.7936
Q8N7X0 | Androglobin | ADGB | 1:1 | Block_2 | 2.8809 | 14.9217 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96PE1 | Adhesion G protein-coupled receptor A2 | ADGRA2 | 1:1 | Block_2 | 1.4513 | 6.5341 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IZF3 | Adhesion G protein-coupled receptor F4 | ADGRF4 | 1:1 | Block_2 | 3.8845 | 12.6443 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P08319 | All-trans-retinol dehydrogenase [NAD(+)] ADH4 | ADH4 | 1:1 | Block_2 | -3.05 | 1.7282 | FALSE | 21.198 | 5.6494 | -2.5185 | -1.2026 | 0.47547 | -2.5061 | -0.76104 | 1.4618
Q16186 | Proteasomal ubiquitin receptor ADRM1 | ADRM1 | 1:1 | Block_2 | 2.1463 | 8.7628 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NB90 | Ribosome biogenesis protein SPATA5 | AFG2A | 1:1 | Block_2 | 1.7895 | 10.4822 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.059
Q5VTM2_Q96P64 | Arf-GAP with GTPase, ANK repeat and PH domain-containing protein 4/9 | AGAP9_AGAP4 | 1:1 | Block_2 | 1.4331 | 13.2951 | TRUE | 12.5543 | 28.509 | NA | NA | NA | NA | NA | 1.5051
Q8NEM8 | Cytosolic carboxypeptidase 3 | AGBL3 | 1:1 | Block_2 | -1.663 | 10.5348 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q719I0 | Putative activator of 90 kDa heat shock protein ATPase homolog 2 | AHSA2P | 1:1 | Block_2 | -0.83085 | 10.5929 | TRUE | 13.2624 | 11.8457 | NA | NA | 0.26816 | NA | NA | 0.72929
P27144 | Adenylate kinase 4, mitochondrial | AK4 | 1:1 | Block_2 | 1.5859 | 10.0473 | FALSE | NA | NA | NA | NA | 1.9093 | NA | 1.7018 | 3.539
Q96M32 | Adenylate kinase 7 | AK7 | 1:1 | Block_2 | 1.3134 | 13.1058 | TRUE | NA | NA | NA | NA | 1.3692 | NA | NA | NA
Q12802 | A-kinase anchor protein 13 | AKAP13 | 1:1 | Block_2 | 0.57788 | 10.7214 | TRUE | NA | NA | NA | NA | 1.1353 | NA | 1.2196 | 2.836
Q86UN6 | A-kinase anchor protein 14 | AKAP14 | 1:1 | Block_2 | 1.9451 | 10.6467 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O43687 | A-kinase anchor protein 7 isoforms alpha and beta | AKAP7 | 1:1 | Block_2 | 0.73735 | 7.1127 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.81532
O43823 | A-kinase anchor protein 8 | AKAP8 | 1:1 | Block_2 | -0.42685 | 12.4679 | TRUE | 5.9403 | 9.8538 | NA | NA | 1.6932 | NA | 0.067081 | 1.9105
P14550 | Aldo-keto reductase family 1 member A1 | AKR1A1 | 1:1 | Block_2 | 2 | 10.1645 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P13196 | 5-aminolevulinate synthase, non-specific, mitochondrial | ALAS1 | 1:1 | Block_2 | 1.6895 | 12.0528 | FALSE | NA | NA | NA | NA | 1.8274 | NA | NA | NA
P47895 | Aldehyde dehydrogenase family 1 member A3 | ALDH1A3 | 1:1 | Block_2 | 1.5 | 7.6639 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O75891 | Cytosolic 10-formyltetrahydrofolate dehydrogenase | ALDH1L1 | 1:1 | Block_2 | -2.1592 | 12.1662 | FALSE | NA | NA | NA | NA | -1.4498 | NA | -1.9538 | 0.11662
Q9UM73 | ALK tyrosine kinase receptor | ALK | 1:1 | Block_2 | -0.90868 | 10.0736 | FALSE | 11.5181 | 15.386 | -0.47501 | 0.18268 | 0.65994 | -0.26238 | 0.41834 | 1.1888
Q86V81 | THO complex subunit 4 | ALYREF | 1:1 | Block_2 | 0.61093 | 8.3452 | TRUE | 24.989 | NA | NA | NA | 0.66605 | NA | 0.81765 | 5.9709
Q6DCA0 | AMMECR1-like protein | AMMECR1L | 1:1 | Block_2 | 1.3308 | 12.4298 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.721
Q9H1A4 | Anaphase-promoting complex subunit 1 | ANAPC1 | 1:1 | Block_2 | -0.19168 | 11.5201 | TRUE | 10.8706 | 9.1281 | NA | NA | NA | NA | NA | 0.30799
Q9UJX5 | Anaphase-promoting complex subunit 4 | ANAPC4 | 1:1 | Block_2 | 2.8821 | 9.4684 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q86XS5 | Angiopoietin-related protein 5 | ANGPTL5 | 1:1 | Block_2 | 0.38121 | 7.4041 | FALSE | 12.6256 | 14.5895 | NA | NA | 0.89136 | NA | NA | 1.1693
Q9P2G1 | Ankyrin repeat and IBR domain-containing protein 1 | ANKIB1 | 1:1 | Block_2 | 1.799 | 12.4449 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.9478
Q9NXR5 | Ankyrin repeat domain-containing protein 10 | ANKRD10 | 1:1 | Block_2 | -1.6531 | 7.3141 | TRUE | NA | NA | NA | NA | NA | NA | NA | -0.2721
A2A2Z9 | Ankyrin repeat domain-containing protein 18B | ANKRD18B | 1:1 | Block_2 | 0.10724 | 8.4316 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9GZV1 | Ankyrin repeat domain-containing protein 2 | ANKRD2 | 1:1 | Block_2 | 0.66813 | 5.9311 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.4508
Q8NF67 | Putative ankyrin repeat domain-containing protein 20A12 pseudogene | ANKRD20A12P | 1:1 | Block_2 | 2.369 | 13.7205 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q4UJ75 | Putative ankyrin repeat domain-containing protein 20A4 | ANKRD20A4P | 1:1 | Block_2 | -0.10194 | 12.4782 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.54413
Q8N2N9 | Ankyrin repeat domain-containing protein 36B | ANKRD36B | 1:1 | Block_2 | 1.6056 | 14.806 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9ULJ7 | Ankyrin repeat domain-containing protein 50 | ANKRD50 | 1:1 | Block_2 | 1.2729 | 11.6625 | TRUE | NA | NA | NA | NA | 1.6522 | NA | NA | 2.0766
Q8NB46 | Serine/threonine-protein phosphatase 6 regulatory ankyrin repeat subunit C | ANKRD52 | 1:1 | Block_2 | 1.8608 | 10.8542 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q92625 | Ankyrin repeat and SAM domain-containing protein 1A | ANKS1A | 1:1 | Block_2 | 0.3272 | 10.9662 | TRUE | 3.8084 | 8.1205 | NA | 0.9464 | 3.3044 | 0.86145 | 3.568 | 5.3532
Q6ZW76 | Ankyrin repeat and SAM domain-containing protein 3 | ANKS3 | 1:1 | Block_2 | 1.5626 | 12.7086 | TRUE | 16.8791 | 14.6823 | NA | NA | 1.567 | NA | NA | 3.1148
A6NFN9 | Protein ANKUB1 | ANKUB1 | 1:1 | Block_2 | 3.083 | 13.1995 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P23352 | Anosmin-1 | ANOS1 | 1:1 | Block_2 | 1.8149 | 10.7478 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P08133 | Annexin A6 | ANXA6 | 1:1 | Block_2 | 1.6631 | 10.6485 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q92870 | Amyloid beta precursor protein binding family B member 2 | APBB2 | 1:1 | Block_2 | 0.49414 | 10.5839 | TRUE | 2.6642 | NA | NA | NA | 0.86805 | NA | NA | 1.337
Q9BWW9 | Apolipoprotein L5 | APOL5 | 1:1 | Block_2 | 1.5 | 9.5714 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NP61 | ADP-ribosylation factor GTPase-activating protein 3 | ARFGAP3 | 1:1 | Block_2 | 1.2564 | 5.9914 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q13795 | ADP-ribosylation factor-related protein 1 | ARFRP1 | 1:1 | Block_2 | 1.185 | 11.7335 | FALSE | 13.3368 | 17.2974 | NA | NA | 1.6632 | NA | NA | 2.6058
A1A4S6 | Rho GTPase-activating protein 10 | ARHGAP10 | 1:1 | Block_2 | 0.7053 | 10.6482 | TRUE | NA | NA | NA | 1.422 | 4.841 | NA | 4.631 | 6.5101
Q6P4F7 | Rho GTPase-activating protein 11A | ARHGAP11A | 1:1 | Block_2 | 1.6235 | 11.7811 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q68EM7 | Rho GTPase-activating protein 17 | ARHGAP17 | 1:1 | Block_2 | 1.9339 | 5.8464 | FALSE | NA | NA | NA | NA | 3.0467 | NA | 2.375 | 4.4761
Q8N264 | Rho GTPase-activating protein 24 | ARHGAP24 | 1:1 | Block_2 | 0.13812 | 8.1082 | TRUE | NA | NA | NA | NA | 0.49143 | NA | NA | 1.9688
Q2M1Z3 | Rho GTPase-activating protein 31 | ARHGAP31 | 1:1 | Block_2 | 2.3793 | 14.0424 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.404
Q9NRY4 | Rho GTPase-activating protein 35 | ARHGAP35 | 1:1 | Block_2 | 2.0701 | 13.0757 | TRUE | NA | NA | NA | NA | 3.3668 | NA | 3.6877 | 5.7451
A6NI28 | Rho GTPase-activating protein 42 | ARHGAP42 | 1:1 | Block_2 | 7.1433 | 17.1094 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O94989 | Rho guanine nucleotide exchange factor 15 | ARHGEF15 | 1:1 | Block_2 | 0.13383 | 9.2131 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IW93 | Rho guanine nucleotide exchange factor 19 | ARHGEF19 | 1:1 | Block_2 | -0.32485 | 10.3653 | TRUE | NA | NA | NA | NA | 0.26367 | NA | NA | 0.61576
Q92974 | Rho guanine nucleotide exchange factor 2 | ARHGEF2 | 1:1 | Block_2 | 3.1585 | 9.343 | TRUE | NA | NA | NA | NA | 4.2997 | NA | NA | 4.5919
Q8N1W1 | Rho guanine nucleotide exchange factor 28 | ARHGEF28 | 1:1 | Block_2 | 0.95465 | 13.9632 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7306
Q9NXL2 | Rho guanine nucleotide exchange factor 38 | ARHGEF38 | 1:1 | Block_2 | 2.1342 | 9.1887 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q03989 | AT-rich interactive domain-containing protein 5A | ARID5A | 1:1 | Block_2 | 1.3101 | 11.4985 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5823
Q96B23 | Uncharacterized protein C18orf25 | ARK2N | 1:1 | Block_2 | 1.1414 | 12.2737 | TRUE | NA | NA | NA | NA | 1.5846 | NA | 1.7111 | 3.4997
P0DKL9 | ARL14 effector protein-like | ARL14EPL | 1:1 | Block_2 | 0.0016811 | 12.5903 | TRUE | 9.1644 | 14.1364 | NA | 0.051419 | 1.8987 | NA | 0.17975 | 1.6439
P56559 | ADP-ribosylation factor-like protein 4C | ARL4C | 1:1 | Block_2 | 0.51176 | 5.3457 | FALSE | NA | NA | NA | NA | 0.61445 | NA | NA | 0.92791
Q7Z6K5 | Arpin | ARPIN | 1:1 | Block_2 | 0.97367 | 11.7677 | FALSE | NA | NA | NA | NA | 1.0668 | NA | NA | 2.0853
Q93070 | Ecto-ADP-ribosyltransferase 4 | ART4 | 1:1 | Block_2 | 1.273 | 8.1026 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.9186
A6NK59 | Ankyrin repeat and SOCS box protein 14 | ASB14 | 1:1 | Block_2 | 1.0642 | 12.1247 | FALSE | NA | NA | NA | NA | 1.3326 | NA | NA | NA
Q8WXK1 | Ankyrin repeat and SOCS box protein 15 | ASB15 | 1:1 | Block_2 | 0.9374 | 12.1563 | FALSE | NA | NA | NA | NA | 1.4914 | NA | NA | 1.6733
Q9Y575 | Ankyrin repeat and SOCS box protein 3 | ASB3 | 1:1 | Block_2 | 2.399 | 8.739 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P50553 | Achaete-scute homolog 1 | ASCL1 | 1:1 | Block_2 | 1.6521 | 10.7696 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q99929 | Achaete-scute homolog 2 | ASCL2 | 1:1 | Block_2 | 1.5076 | 9.1819 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NR48 | Histone-lysine N-methyltransferase ASH1L | ASH1L | 1:1 | Block_2 | 2.3495 | 12.1653 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q76L83 | Putative Polycomb group protein ASXL2 | ASXL2 | 1:1 | Block_2 | 1.0781 | 11.0845 | TRUE | NA | NA | NA | NA | 1.4393 | NA | NA | 1.6138
P18846 | Cyclic AMP-dependent transcription factor ATF-1 | ATF1 | 1:1 | Block_2 | 0.33425 | 11.2485 | FALSE | NA | 9.3358 | NA | NA | 0.88641 | NA | 1.0908 | 2.857
P18848 | Cyclic AMP-dependent transcription factor ATF-4 | ATF4 | 1:1 | Block_2 | 1.9274 | 11.8004 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H0Y0 | Ubiquitin-like-conjugating enzyme ATG10 | ATG10 | 1:1 | Block_2 | 1.3093 | 8.5453 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BSB4 | Autophagy-related protein 101 | ATG101 | 1:1 | Block_2 | 1.6119 | 6.8093 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.6597
O75143 | Autophagy-related protein 13 | ATG13 | 1:1 | Block_2 | 2.2665 | 11.7826 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZNE5 | Beclin 1-associated autophagy-related key regulator | ATG14 | 1:1 | Block_2 | 2.5398 | 10.9439 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6DD88 | Atlastin-3 | ATL3 | 1:1 | Block_2 | 1.1397 | 11.4948 | TRUE | NA | NA | NA | NA | 1.2775 | NA | NA | 1.6701
Q16720 | Plasma membrane calcium-transporting ATPase 3 | ATP2B3 | 1:1 | Block_2 | 0.90883 | 12.1002 | FALSE | NA | NA | NA | NA | NA | NA | NA | 0.92443
P98194 | Calcium-transporting ATPase type 2C member 1 | ATP2C1 | 1:1 | Block_2 | 1.143 | 12.246 | TRUE | NA | NA | NA | NA | 1.7115 | NA | NA | 1.807
P61421 | V-type proton ATPase subunit d 1 | ATP6V0D1 | 1:1 | Block_2 | 0.72434 | 8.77 | TRUE | NA | NA | NA | 1.4322 | 3.2604 | 1.123 | 3.1704 | 4.8774
Q8N8Y2 | V-type proton ATPase subunit d 2 | ATP6V0D2 | 1:1 | Block_2 | 1.3378 | 5.7097 | TRUE | NA | NA | NA | NA | 1.3515 | NA | NA | NA
Q8NEY4 | V-type proton ATPase subunit C 2 | ATP6V1C2 | 1:1 | Block_2 | 0.9293 | 8.5899 | TRUE | NA | NA | NA | NA | 1.0528 | NA | NA | 1.0487
Q9ULK2 | Ataxin-7-like protein 1 | ATXN7L1 | 1:1 | Block_2 | 0.73228 | 14.1405 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.92217
Q6L9W6 | Beta-1,4-N-acetylgalactosaminyltransferase 3 | B4GALNT3 | 1:1 | Block_2 | 1.0212 | 8.8932 | FALSE | NA | 0.064789 | NA | NA | NA | NA | NA | 1.519
P56817 | Beta-secretase 1 | BACE1 | 1:1 | Block_2 | -5 | 7.8535 | FALSE | 12.6291 | 15.9069 | -0.50052 | 0.29295 | 0.70053 | -0.50375 | 0.35642 | 1.0706
Q9UIF9 | Bromodomain adjacent to zinc finger domain protein 2A | BAZ2A | 1:1 | Block_2 | 1.1203 | 11.7541 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8ND07 | Basal body-orientation factor 1 | BBOF1 | 1:1 | Block_2 | 0.7719 | 12.4104 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WY36 | HMG box transcription factor BBX | BBX | 1:1 | Block_2 | 0.39564 | 13.563 | FALSE | 14.1267 | 6.2501 | NA | NA | NA | NA | NA | 1.1923
O75815 | Breast cancer anti-estrogen resistance protein 3 | BCAR3 | 1:1 | Block_2 | -0.41901 | 12.4064 | TRUE | 31.705 | 17.6747 | NA | NA | 0.15124 | NA | NA | 1.0331
Q9C0K0 | B-cell lymphoma/leukemia 11B | BCL11B | 1:1 | Block_2 | 0.56638 | 11.0737 | TRUE | NA | NA | NA | NA | 0.78514 | NA | NA | 1.8504
Q9HD36 | Bcl-2-like protein 10 | BCL2L10 | 1:1 | Block_2 | 1.258 | 12.2987 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q92843 | Bcl-2-like protein 2 | BCL2L2 | 1:1 | Block_2 | -0.96636 | 7.3931 | FALSE | NA | NA | NA | NA | -0.63831 | NA | NA | -0.5
Q8WUZ0 | B-cell CLL/lymphoma 7 protein family member C | BCL7C | 1:1 | Block_2 | 1.8347 | 14.1044 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.4395
Q9NYF8 | Bcl-2-associated transcription factor 1 | BCLAF1 | 1:1 | Block_2 | 2.7342 | 5.9186 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NDZ0 | BEN domain-containing protein 2 | BEND2 | 1:1 | Block_2 | 2.5714 | 12.2022 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZU67 | BEN domain-containing protein 4 | BEND4 | 1:1 | Block_2 | 0.70027 | 10.0744 | FALSE | NA | 4.0102 | NA | NA | 0.98781 | NA | NA | 0.98196
Q7L4P6 | BEN domain-containing protein 5 | BEND5 | 1:1 | Block_2 | 3.4429 | 16.5134 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N7W2 | BEN domain-containing protein 7 | BEND7 | 1:1 | Block_2 | 0.64695 | 11.6891 | FALSE | 8.6062 | 12.9168 | NA | NA | 1.3863 | NA | NA | 1.0797
Q12934 | Filensin | BFSP1 | 1:1 | Block_2 | 1.5673 | 13.6532 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86UB2 | Basic immunoglobulin-like variable motif-containing protein | BIVM | 1:1 | Block_2 | 1.7569 | 13.3848 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TDH9 | Biogenesis of lysosome-related organelles complex 1 subunit 5 | BLOC1S5 | 1:1 | Block_2 | 1.5818 | 12.3455 | TRUE | NA | 24.4406 | NA | NA | 1.8085 | NA | 1.9396 | 3.3195
Q5H9B9 | Putative BMP-2-inducible kinase-like protein | BMP2KL | 1:1 | Block_2 | 1.6337 | 11.5481 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.0583
Q01954 | Zinc finger protein basonuclin-1 | BNC1 | 1:1 | Block_2 | 1.0789 | 10.2544 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.1447
P0C671 | Protein BNIP5 | BNIP5 | 1:1 | Block_2 | 2.5 | 10.4671 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q92994 | Transcription factor IIIB 90 kDa subunit | BRF1 | 1:1 | Block_2 | 1.0694 | 12.1883 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.3104
Q9BX63 | Fanconi anemia group J protein | BRIP1 | 1:1 | Block_2 | 1.4379 | 14.0766 | FALSE | NA | NA | NA | NA | 2.6834 | NA | NA | NA
Q6RI45 | Bromodomain and WD repeat-containing protein 3 | BRWD3 | 1:1 | Block_2 | 1.985 | 13.3209 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UPA5 | Protein bassoon | BSN | 1:1 | Block_2 | 1.8144 | 14.3023 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.9487
Q9BX70 | BTB/POZ domain-containing protein 2 | BTBD2 | 1:1 | Block_2 | 2.4665 | 10.5471 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y2F9 | BTB/POZ domain-containing protein 3 | BTBD3 | 1:1 | Block_2 | 2.5096 | 13.2496 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P203 | BTB/POZ domain-containing protein 7 | BTBD7 | 1:1 | Block_2 | 1.9676 | 11.7078 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5XKL5 | BTB/POZ domain-containing protein 8 | BTBD8 | 1:1 | Block_2 | 1.4427 | 11.8397 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96Q07 | BTB/POZ domain-containing protein 9 | BTBD9 | 1:1 | Block_2 | 2.5 | 14.3402 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.8073
Q6UXG8 | Butyrophilin-like protein 9 | BTNL9 | 1:1 | Block_2 | 1.4284 | 10.8377 | TRUE | NA | 9.9615 | NA | NA | 1.4656 | NA | NA | 2.4558
O60566 | Mitotic checkpoint serine/threonine-protein kinase BUB1 beta | BUB1B | 1:1 | Block_2 | 0.75875 | 14.1645 | TRUE | NA | NA | NA | NA | 1.3598 | NA | NA | 2.0702
Q9H7T3 | Uncharacterized protein C10orf95 | C10orf95 | 1:1 | Block_2 | 0.48582 | 9.0714 | TRUE | NA | NA | NA | NA | 0.4871 | NA | NA | 1.1238
Q96A22 | Uncharacterized protein C11orf52 | C11orf52 | 1:1 | Block_2 | 1.7325 | 14.6707 | TRUE | 14.0144 | 18.1949 | NA | NA | 2.1173 | NA | NA | 2.5497
Q86WS4 | Uncharacterized protein C12orf40 | C12orf40 | 1:1 | Block_2 | 2.0022 | 13.5831 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IXR9 | Uncharacterized protein C12orf56 | C12orf56 | 1:1 | Block_2 | 3.2786 | 13.5898 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WTQ4 | Uncharacterized protein C16orf78 | C16orf78 | 1:1 | Block_2 | 0.47835 | 11.4617 | TRUE | NA | NA | NA | NA | 2.0355 | NA | NA | 0.78559
Q7Z2V1 | Protein TNT | C16orf82 | 1:1 | Block_2 | -0.47157 | 13.0767 | TRUE | NA | NA | NA | NA | 1.208 | NA | NA | 0.12263
Q6ZW13 | Uncharacterized protein C16orf86 | C16orf86 | 1:1 | Block_2 | 0.77591 | 12.4479 | TRUE | NA | 16.956 | NA | NA | 2.5515 | NA | 2.4403 | 4.2818
A6NNT2 | Uncharacterized protein C16orf96 | C16orf96 | 1:1 | Block_2 | 1 | 9.7447 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WW18 | Uncharacterized protein C17orf50 | C17orf50 | 1:1 | Block_2 | 2.213 | 12.6002 | TRUE | NA | NA | NA | NA | 2.8141 | NA | NA | NA
Q6UX52 | Protein IL-40 | C17orf99 | 1:1 | Block_2 | 2.0372 | 9.4304 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.2499
Q8NDD1 | Uncharacterized protein C1orf131 | C1orf131 | 1:1 | Block_2 | 2.0478 | 16.9902 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.7378
Q8WYQ4 | Uncharacterized protein C22orf15 | C22orf15 | 1:1 | Block_2 | 1.6653 | 7.4647 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O95567 | Uncharacterized protein C22orf31 | C22orf31 | 1:1 | Block_2 | 1.2896 | 12.5353 | TRUE | NA | 26.9153 | NA | NA | 2.418 | NA | NA | 2.5031
Q6IC83 | Uncharacterized protein C22orf42 | C22orf42 | 1:1 | Block_2 | 1.8465 | 10.899 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q68DN1 | Uncharacterized protein C2orf16 | C2orf16 | 1:1 | Block_2 | 1.1001 | 11.2947 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q2NKX9 | UPF0561 protein C2orf68 | C2orf68 | 1:1 | Block_2 | 1.6769 | 12.2202 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.0435
Q8N5S3 | Uncharacterized protein C2orf73 | C2orf73 | 1:1 | Block_2 | -0.55639 | 11.2912 | TRUE | 9.3331 | 17.9768 | NA | NA | -0.2337 | NA | NA | NA
Q0P641 | Uncharacterized protein C2orf80 | C2orf80 | 1:1 | Block_2 | 0.67221 | 12.194 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.621
A6NN90 | Uncharacterized protein C2orf81 | C2orf81 | 1:1 | Block_2 | 1.6009 | 7.9385 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5JPI3 | Uncharacterized protein C3orf38 | C3orf38 | 1:1 | Block_2 | 1.8235 | 9.2418 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZUJ4 | Uncharacterized protein C3orf62 | C3orf62 | 1:1 | Block_2 | 2.2601 | 12.1188 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
A6NLC5 | UPF0524 protein C3orf70 | C3orf70 | 1:1 | Block_2 | 2.7986 | 7.9743 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q504U0 | Renal cancer differentiation gene 1 protein | C4orf46 | 1:1 | Block_2 | 1.1529 | 11.1499 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.2542
A7E2U8 | UPF0602 protein C4orf47 | C4orf47 | 1:1 | Block_2 | 1.2866 | 13.3862 | TRUE | NA | NA | NA | NA | 1.5859 | NA | NA | NA
D6RIA3 | Uncharacterized protein C4orf54 | C4orf54 | 1:1 | Block_2 | 1.5 | 11.2255 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q49AR2 | UPF0489 protein C5orf22 | C5orf22 | 1:1 | Block_2 | 2.8707 | 10.9465 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q7Z6I8 | UPF0461 protein C5orf24 | C5orf24 | 1:1 | Block_2 | 0.6651 | 10.8191 | TRUE | NA | NA | NA | NA | 1.4226 | NA | NA | 1.2291
Q96MH7 | Uncharacterized protein C5orf34 | C5orf34 | 1:1 | Block_2 | 0.76892 | 12.0435 | TRUE | NA | NA | NA | NA | 2.8824 | NA | NA | 1.7582
Q5T5N4 | Uncharacterized protein C6orf118 | C6orf118 | 1:1 | Block_2 | 1.1951 | 7.2604 | FALSE | NA | NA | NA | NA | 1.2992 | NA | NA | 1.2355
Q5TEZ5 | Uncharacterized protein C6orf163 | C6orf163 | 1:1 | Block_2 | 1.3192 | 14.8515 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BPX7 | UPF0415 protein C7orf25 | C7orf25 | 1:1 | Block_2 | 0.54945 | 10.7825 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.2204
Q9H7E9 | UPF0488 protein C8orf33 | C8orf33 | 1:1 | Block_2 | 1.1274 | 11.7292 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.1564
Q49A92 | Uncharacterized protein C8orf34 | C8orf34 | 1:1 | Block_2 | 0.46946 | 12.3136 | FALSE | 13.8869 | 9.3881 | NA | NA | 1.2806 | NA | NA | 2.2296
Q8NAV2 | Uncharacterized protein C8orf58 | C8orf58 | 1:1 | Block_2 | 0.74914 | 12.54 | TRUE | 10.2954 | 14.5362 | NA | NA | 2.5015 | NA | NA | 1.4859
Q5JTZ5 | Uncharacterized protein C9orf152 | C9orf152 | 1:1 | Block_2 | 1.5163 | 12.5003 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.9149
Q8TAL5 | Uncharacterized protein C9orf43 | C9orf43 | 1:1 | Block_2 | 0.64005 | 12.2929 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y2D0 | Carbonic anhydrase 5B, mitochondrial | CA5B | 1:1 | Block_2 | 1.5697 | 11.5243 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y6J0 | Calcineurin-binding protein cabin-1 | CABIN1 | 1:1 | Block_2 | 1.0896 | 12.8154 | TRUE | NA | NA | NA | NA | 1.4368 | NA | NA | 1.4024
Q86V35 | Calcium-binding protein 7 | CABP7 | 1:1 | Block_2 | 2.0322 | 13.1272 | TRUE | NA | NA | NA | NA | 2.2738 | NA | NA | 2.643
O75952 | Calcium-binding tyrosine phosphorylation-regulated protein | CABYR | 1:1 | Block_2 | 0.36899 | 13.8435 | TRUE | NA | NA | NA | NA | 0.98193 | NA | NA | 0.91698
Q8TD86 | Calmodulin-like protein 6 | CALML6 | 1:1 | Block_2 | 1.8428 | 11.2158 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6NHC0 | Calpain-8 | CAPN8 | 1:1 | Block_2 | 1.2877 | 9.1206 | FALSE | 13.1019 | 29.2601 | NA | NA | 1.485 | NA | NA | 2.5823
Q6IMN6 | Caprin-2 | CAPRIN2 | 1:1 | Block_2 | 2.0718 | 12.7856 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WXE0 | Caskin-2 | CASKIN2 | 1:1 | Block_2 | -0.39384 | 6.0066 | FALSE | NA | NA | NA | NA | 0.17486 | NA | NA | 1.3958
Q9UKL3 | CASP8-associated protein 2 | CASP8AP2 | 1:1 | Block_2 | 0.48788 | 15.3106 | TRUE | NA | 16.9977 | NA | NA | 1.5506 | NA | NA | 1.7401
Q86V15 | Zinc finger protein castor homolog 1 | CASZ1 | 1:1 | Block_2 | 2.7628 | 11.3427 | TRUE | 12.4107 | 25.6441 | NA | NA | NA | NA | NA | NA
Q8N7E2 | E3 ubiquitin-protein ligase CBLL2 | CBLL2 | 1:1 | Block_2 | -0.40303 | 8.3403 | TRUE | NA | NA | NA | NA | 0.0090185 | NA | NA | NA
O75828 | Carbonyl reductase [NADPH] 3 | CBR3 | 1:1 | Block_2 | 0.87448 | 11.3587 | TRUE | NA | NA | NA | NA | 1.0364 | NA | 0.91215 | 2.4966
O00257 | E3 SUMO-protein ligase CBX4 | CBX4 | 1:1 | Block_2 | 1.6741 | 12.3907 | TRUE | NA | NA | NA | NA | 2.1397 | NA | NA | 2.0617
Q9Y3M2 | Protein chibby homolog 1 | CBY1 | 1:1 | Block_2 | 2.0235 | 14.8763 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.1172
Q5T0F9 | Coiled-coil and C2 domain-containing protein 1B | CC2D1B | 1:1 | Block_2 | 0.35268 | 5.9291 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IWD4 | Coiled-coil domain-containing protein 117 | CCDC117 | 1:1 | Block_2 | 0.60161 | 10.4067 | FALSE | 10.3041 | 8.1265 | NA | NA | 0.87394 | NA | 0.71146 | 1.7462
Q86Z20 | Coiled-coil domain-containing protein 125 | CCDC125 | 1:1 | Block_2 | -0.29069 | 12.0461 | TRUE | NA | NA | NA | NA | 0.97234 | NA | NA | 0.93874
Q96BQ5 | Coiled-coil domain-containing protein 127 | CCDC127 | 1:1 | Block_2 | 1 | 13.0363 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IYE1 | Coiled-coil domain-containing protein 13 | CCDC13 | 1:1 | Block_2 | 2.5 | 14.2421 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.464
Q6PK04 | Coiled-coil domain-containing protein 137 | CCDC137 | 1:1 | Block_2 | 2.3613 | 13.1893 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96M89 | Coiled-coil domain-containing protein 138 | CCDC138 | 1:1 | Block_2 | -5.1874 | 6.5678 | TRUE | 24.7374 | 14.799 | NA | NA | -4.5987 | NA | NA | -3.2393
Q49A88 | Coiled-coil domain-containing protein 14 | CCDC14 | 1:1 | Block_2 | 3.2321 | 15.3172 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.5731
Q96MF4 | Coiled-coil domain-containing protein 140 | CCDC140 | 1:1 | Block_2 | 1.0353 | 11.758 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.1719
Q8IYA2 | Putative coiled-coil domain-containing protein 144C | CCDC144CP | 1:1 | Block_2 | 0.93107 | 11.3109 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NFR7 | Coiled-coil domain-containing protein 148 | CCDC148 | 1:1 | Block_2 | 0.93448 | 13.5358 | TRUE | 8.6251 | 3.9355 | NA | NA | 1.1088 | NA | NA | 1.0733
Q569K6 | Coiled-coil domain-containing protein 157 | CCDC157 | 1:1 | Block_2 | 1.2741 | 11.9221 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NDH2 | Leucine-rich repeat transmembrane protein CCDC168 | CCDC168 | 1:1 | Block_2 | 2.6417 | 13.1 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6NNP5 | Coiled-coil domain-containing protein 169 | CCDC169 | 1:1 | Block_2 | 1.4815 | 13.3118 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IYT3 | Coiled-coil domain-containing protein 170 | CCDC170 | 1:1 | Block_2 | 0.46313 | 13.7166 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.4827
Q8NCU1 | Uncharacterized protein CCDC197 | CCDC197 | 1:1 | Block_2 | 1.873 | 14.2934 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.0183
Q8N4L8 | Coiled-coil domain-containing protein 24 | CCDC24 | 1:1 | Block_2 | 2.7299 | 15.0594 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BUN5 | Coiled-coil domain-containing protein 28B | CCDC28B | 1:1 | Block_2 | 0.39654 | 9.5872 | TRUE | NA | NA | NA | NA | 0.83649 | NA | 0.49934 | 1.8674
Q4G0X9 | Coiled-coil domain-containing protein 40 | CCDC40 | 1:1 | Block_2 | 3.0011 | 9.3848 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96MW1 | Coiled-coil domain-containing protein 43 | CCDC43 | 1:1 | Block_2 | -0.046607 | 8.8418 | TRUE | 8.8036 | 8.2573 | NA | 1.1959 | 2.8522 | 0.58905 | 3.432 | 4.8629
Q8NA47 | Coiled-coil domain-containing protein 63 | CCDC63 | 1:1 | Block_2 | 0.45329 | 10.7324 | TRUE | 4.0634 | 8.7214 | NA | NA | 0.55273 | NA | NA | 0.73501
Q8IXS2 | Dynein regulatory complex subunit 2 | CCDC65 | 1:1 | Block_2 | 2.9585 | 10.3826 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A2RUB6 | Coiled-coil domain-containing protein 66 | CCDC66 | 1:1 | Block_2 | 1.458 | 13.7987 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.5167
Q96M83 | Coiled-coil domain-containing protein 7 | CCDC7 | 1:1 | Block_2 | 1.8507 | 14.0307 | TRUE | NA | NA | NA | NA | 3.8372 | NA | NA | 2.1138
Q6ZRK6 | Coiled-coil domain-containing protein 73 | CCDC73 | 1:1 | Block_2 | 1.7951 | 14.0417 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8974
A6NC98 | Coiled-coil domain-containing protein 88B | CCDC88B | 1:1 | Block_2 | -0.81099 | 7.9879 | TRUE | NA | NA | NA | NA | -0.56884 | NA | -0.51408 | 1.2524
Q8N998 | Coiled-coil domain-containing protein 89 | CCDC89 | 1:1 | Block_2 | 1.1852 | 14.2575 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.333
Q7Z6B0 | Coiled-coil domain-containing protein 91 | CCDC91 | 1:1 | Block_2 | 1.4287 | 12.0045 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.124
Q8TD31 | Coiled-coil alpha-helical rod protein 1 | CCHCR1 | 1:1 | Block_2 | 1.2251 | 15.1965 | TRUE | NA | NA | NA | NA | 1.4883 | NA | 1.3211 | 3.3816
P22362 | C-C motif chemokine 1 | CCL1 | 1:1 | Block_2 | 2.038 | 15.6644 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UK58 | Cyclin-L1 | CCNL1 | 1:1 | Block_2 | 0.69248 | 12.2992 | TRUE | NA | NA | NA | NA | 0.83322 | NA | NA | 1.0859
Q96S94 | Cyclin-L2 | CCNL2 | 1:1 | Block_2 | 1.7973 | 11.1382 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5T2Q4 | Cyclin-Y-like protein 2 | CCNYL2 | 1:1 | Block_2 | 0.66196 | 15.7556 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8912
O43303 | Centriolar coiled-coil protein of 110 kDa | CCP110 | 1:1 | Block_2 | 2.1326 | 12.6749 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.1564
Q6IQ19 | Centriole, cilia and spindle-associated protein | CCSAP | 1:1 | Block_2 | -0.2505 | 13.3722 | TRUE | 5.6439 | 10.9172 | NA | 0.054127 | 1.376 | -0.24712 | 0.61682 | 2.4678
Q6Q8B3 | Cell surface glycoprotein CD200 receptor 2 | CD200R1L | 1:1 | Block_2 | 1.9757 | 10.469 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O95400 | CD2 antigen cytoplasmic tail-binding protein 2 | CD2BP2 | 1:1 | Block_2 | 0.93139 | 6.7678 | TRUE | NA | NA | NA | NA | 1.1374 | NA | NA | 1.2916
A8K4G0 | CMRF35-like molecule 7 | CD300LB | 1:1 | Block_2 | -0.032972 | 10.3006 | TRUE | 21.4856 | 10.5012 | NA | 0.04561 | 0.47126 | NA | 0.64428 | 1.8872
P11912 | B-cell antigen receptor complex-associated protein alpha chain | CD79A | 1:1 | Block_2 | 2.5 | 8.9133 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q7L3B6 | Hsp90 co-chaperone Cdc37-like 1 | CDC37L1 | 1:1 | Block_2 | 2.0627 | 11.3056 | TRUE | NA | NA | NA | NA | NA | NA | 2.1449 | 3.5115
Q00587 | Cdc42 effector protein 1 | CDC42EP1 | 1:1 | Block_2 | 0.62332 | 10.6911 | TRUE | 19.878 | 11.3181 | NA | NA | 0.73132 | NA | NA | 1.0662
P55289 | Cadherin-12 | CDH12 | 1:1 | Block_2 | 0.85568 | 7.631 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
A6H8M9 | Cadherin-related family member 4 | CDHR4 | 1:1 | Block_2 | 0.65708 | 9.3948 | FALSE | NA | NA | NA | NA | 0.81367 | NA | NA | 1.9202
Q9NYV4 | Cyclin-dependent kinase 12 | CDK12 | 1:1 | Block_2 | -0.65524 | 9.0223 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O94921 | Cyclin-dependent kinase 14 | CDK14 | 1:1 | Block_2 | 2.0301 | 13.3838 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.6041
Q00537 | Cyclin-dependent kinase 17 | CDK17 | 1:1 | Block_2 | 0.085844 | 12.1428 | TRUE | NA | 30.2693 | NA | 1.8924 | 3.849 | 0.88658 | 3.7308 | 5.3556
Q9BWU1 | Cyclin-dependent kinase 19 | CDK19 | 1:1 | Block_2 | 2 | 11.9784 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.1825
O75956 | Cyclin-dependent kinase 2-associated protein 2 | CDK2AP2 | 1:1 | Block_2 | 2.6319 | 9.3759 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IVW4 | Cyclin-dependent kinase-like 3 | CDKL3 | 1:1 | Block_2 | 1.1907 | 11.7772 | TRUE | NA | NA | NA | NA | 1.5964 | NA | NA | 1.1914
O76039 | Cyclin-dependent kinase-like 5 | CDKL5 | 1:1 | Block_2 | 2.2854 | 11.8668 | TRUE | NA | NA | NA | NA | 2.3753 | NA | NA | 2.843
P49918 | Cyclin-dependent kinase inhibitor 1C | CDKN1C | 1:1 | Block_2 | 2.7355 | 8.6537 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q16878 | Cysteine dioxygenase type 1 | CDO1 | 1:1 | Block_2 | 2.7371 | 12.1083 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y6F8 | Testis-specific chromodomain protein Y 1 | CDY1 | 1:1 | Block_2 | 1.3596 | 13.2175 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q14002 | Carcinoembryonic antigen-related cell adhesion molecule 7 | CEACAM7 | 1:1 | Block_2 | 1.7215 | 13.0493 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8075
P17676 | CCAAT/enhancer-binding protein beta | CEBPB | 1:1 | Block_2 | 0.38 | 7.9608 | FALSE | 18.8682 | 9.1966 | NA | NA | 0.95073 | NA | 0.52072 | 2.1318
Q711Q0 | Cardiac-enriched FHL2-interacting protein | CEFIP | 1:1 | Block_2 | 0.93391 | 9.9292 | TRUE | 10.6045 | 7.7954 | NA | NA | NA | NA | NA | 1.0662
Q9BZC1 | CUGBP Elav-like family member 4 | CELF4 | 1:1 | Block_2 | 1.3613 | 8.5375 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.8245
Q86UT8 | Centrosomal AT-AC splicing factor | CENATAC | 1:1 | Block_2 | 2.1209 | 12.9864 | TRUE | NA | NA | NA | NA | 2.6186 | NA | 3.0394 | 4.6253
Q5TB80 | Centrosomal protein of 162 kDa | CEP162 | 1:1 | Block_2 | 0.79818 | 9.4917 | TRUE | 4.272 | 18.68 | NA | NA | NA | NA | NA | 1.3343
Q96MC4 | CEP295 N-terminal-like protein | CEP295NL | 1:1 | Block_2 | 0.34871 | 13.3089 | FALSE | NA | NA | NA | NA | 1.4254 | NA | NA | 1.868
Q8IYX8 | Centrosomal protein CEP57L1 | CEP57L1 | 1:1 | Block_2 | 1.8944 | 12.0654 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5SZL2 | Centrosomal protein of 85 kDa-like | CEP85L | 1:1 | Block_2 | 1.1116 | 12.7161 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.9163
Q96ST8 | Centrosomal protein of 89 kDa | CEP89 | 1:1 | Block_2 | 0.68519 | 13.2056 | TRUE | NA | NA | NA | NA | 0.7448 | NA | NA | 1.423
Q8IW35 | Centrosomal protein of 97 kDa | CEP97 | 1:1 | Block_2 | 2.3409 | 11.8172 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.6769
Q8TCT0 | Ceramide kinase | CERK | 1:1 | Block_2 | 2.1313 | 9.2689 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6P656 | Cilia- and flagella-associated protein 161 | CFAP161 | 1:1 | Block_2 | 0.68557 | 10.5979 | TRUE | NA | 4.7527 | NA | NA | 0.85821 | NA | NA | 0.84892
Q9Y6A4 | Cilia- and flagella-associated protein 20 | CFAP20 | 1:1 | Block_2 | 2.2762 | 13.9068 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.546
Q8TBY9 | Cilia- and flagella-associated protein 251 | CFAP251 | 1:1 | Block_2 | 0.98919 | 11.3009 | TRUE | NA | 4.5381 | NA | NA | NA | NA | NA | 1.1828
Q5T5A4 | Cilia- and flagella-associated protein 276 | CFAP276 | 1:1 | Block_2 | 1.0397 | 12.1625 | FALSE | NA | NA | NA | NA | 1.2352 | NA | NA | 1.7189
Q6V702 | Cilia- and flagella-associated protein 299 | CFAP299 | 1:1 | Block_2 | 1.0924 | 13.8737 | TRUE | NA | NA | NA | NA | 1.357 | NA | NA | 1.8428
Q8IYW2 | Cilia- and flagella-associated protein 46 | CFAP46 | 1:1 | Block_2 | 1.5248 | 13.3805 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.155
Q5T655 | Cilia- and flagella-associated protein 58 | CFAP58 | 1:1 | Block_2 | 0.97647 | 10.842 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NHU2 | Cilia- and flagella-associated protein 61 | CFAP61 | 1:1 | Block_2 | 1.4662 | 13.5856 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.3526
A5D8W1 | Cilia- and flagella-associated protein 69 | CFAP69 | 1:1 | Block_2 | 2.5474 | 12.4358 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9ULG3 | Uncharacterized protein CFAP92 | CFAP92 | 1:1 | Block_2 | -2.0683 | 9.5317 | TRUE | NA | NA | NA | NA | -1.1516 | NA | NA | -1.6606
D6REC4 | Cilia- and flagella-associated protein 99 | CFAP99 | 1:1 | Block_2 | -0.28689 | 12.9059 | TRUE | 18.5053 | 19.0373 | NA | NA | 0.33897 | NA | NA | 0.76112
Q9BUX1 | Glutathione-specific gamma-glutamylcyclotransferase 1 | CHAC1 | 1:1 | Block_2 | 1.2589 | 8.2953 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4152
Q5T1J5 | Putative coiled-coil-helix-coiled-coil-helix domain-containing protein CHCHD2P9, mitochondrial | CHCHD2P9 | 1:1 | Block_2 | 1.8869 | 10.5901 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.0227
Q14839 | Chromodomain-helicase-DNA-binding protein 4 | CHD4 | 1:1 | Block_2 | 0.2212 | 7.7754 | FALSE | 22.1447 | 13.8532 | NA | NA | NA | NA | NA | 0.86268
Q8TD26 | Chromodomain-helicase-DNA-binding protein 6 | CHD6 | 1:1 | Block_2 | 1.3844 | 13.9654 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5113
O96017 | Serine/threonine-protein kinase Chk2 | CHEK2 | 1:1 | Block_2 | 3.3291 | 11.1229 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y259 | Choline/ethanolamine kinase | CHKB | 1:1 | Block_2 | 1.2657 | 13.0573 | FALSE | NA | NA | NA | NA | 1.4995 | NA | NA | 1.2787
Q96CF2 | Charged multivesicular body protein 4c | CHMP4C | 1:1 | Block_2 | 0.83627 | 9.2584 | TRUE | NA | NA | NA | NA | 0.916 | NA | NA | 1.4052
P08172 | Muscarinic acetylcholine receptor M2 | CHRM2 | 1:1 | Block_2 | 1.5 | 10.4897 | TRUE | NA | NA | NA | NA | 1.9537 | NA | NA | 2.7491
P20309 | Muscarinic acetylcholine receptor M3 | CHRM3 | 1:1 | Block_2 | 0.4939 | 9.5923 | FALSE | 14.12 | 12.4639 | NA | NA | 1.666 | NA | 0.6602 | 1.7289
Q9Y4C5 | Carbohydrate sulfotransferase 2 | CHST2 | 1:1 | Block_2 | 0.6911 | 9.1942 | FALSE | NA | NA | NA | NA | 1.055 | NA | 0.7096 | 1.6297
Q8NCG5 | Carbohydrate sulfotransferase 4 | CHST4 | 1:1 | Block_2 | 0.53136 | 8.7813 | FALSE | NA | NA | NA | NA | 0.75726 | NA | NA | 1.4954
Q9H5X1 | Cytosolic iron-sulfur assembly component 2A | CIAO2A | 1:1 | Block_2 | 1.9075 | 10.8176 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.5505
Q8N365 | Circadian-associated transcriptional repressor | CIART | 1:1 | Block_2 | 3.5557 | 8.5384 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O75838 | Calcium and integrin-binding family member 2 | CIB2 | 1:1 | Block_2 | 1.6847 | 13.0512 | TRUE | NA | NA | NA | NA | 1.8726 | NA | NA | 2.2656
Q9UPZ9 | Serine/threonine-protein kinase ICK | CILK1 | 1:1 | Block_2 | 0.7497 | 14.1204 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4793
Q8IYA6 | Cytoskeleton-associated protein 2-like | CKAP2L | 1:1 | Block_2 | 0.22783 | 9.2321 | TRUE | 0.581 | 12.0983 | NA | NA | 2.7881 | NA | NA | 2.482
Q96B33 | Claudin-23 | CLDN23 | 1:1 | Block_2 | 1.4109 | 6.5997 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9ULY5 | C-type lectin domain family 4 member E | CLEC4E | 1:1 | Block_2 | 0.5973 | 13.9713 | TRUE | 3.6758 | 10.6772 | NA | NA | 1.8356 | NA | 0.95007 | 2.5451
Q8NHS4 | Clathrin heavy chain linker domain-containing protein 1 | CLHC1 | 1:1 | Block_2 | -0.80109 | 9.5608 | FALSE | 12.4399 | 10.481 | NA | NA | NA | NA | NA | -0.63927
O00299 | Chloride intracellular channel protein 1 | CLIC1 | 1:1 | Block_2 | 0.46789 | 9.6593 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.91303
O95833 | Chloride intracellular channel protein 3 | CLIC3 | 1:1 | Block_2 | 1.4053 | 8.4419 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.5076
P49761 | Dual specificity protein kinase CLK3 | CLK3 | 1:1 | Block_2 | 0.39166 | 9.1198 | TRUE | 2.4332 | 6.7408 | NA | NA | 0.66177 | NA | NA | 1.974
Q5K130 | Putative uncharacterized protein CLLU1-AS1 | CLLU1-AS1 | 1:1 | Block_2 | 0.65352 | 12.5877 | TRUE | NA | NA | NA | NA | 1.287 | NA | NA | 1.4778
O76031 | ATP-dependent Clp protease ATP-binding subunit clpX-like, mitochondrial | CLPX | 1:1 | Block_2 | 1 | 12.482 | TRUE | 2.7497 | 10.0614 | NA | NA | NA | NA | NA | 1.1811
P09496 | Clathrin light chain A | CLTA | 1:1 | Block_2 | 1 | 8.1 | FALSE | NA | NA | NA | 2.1087 | 4.9915 | 2.6246 | 4.3982 | 5.9397
P53675 | Clathrin heavy chain 2 | CLTCL1 | 1:1 | Block_2 | 1.7326 | 11.0922 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96AJ1 | Clusterin-associated protein 1 | CLUAP1 | 1:1 | Block_2 | 1.4149 | 10.0971 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.8179
Q8NA66 | Cyclic nucleotide-binding domain-containing protein 1 | CNBD1 | 1:1 | Block_2 | 0.94362 | 12.8929 | FALSE | 13.6809 | 7.3931 | NA | NA | 1.4958 | NA | NA | 1.4941
Q16281 | Cyclic nucleotide-gated cation channel alpha-3 | CNGA3 | 1:1 | Block_2 | 0.67709 | 12.8542 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.75881
Q8WXI2 | Connector enhancer of kinase suppressor of ras 2 | CNKSR2 | 1:1 | Block_2 | 2.3179 | 13.5213 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.3744
Q9H9A5 | CCR4-NOT transcription complex subunit 10 | CNOT10 | 1:1 | Block_2 | 2.0223 | 12.9203 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O95628 | CCR4-NOT transcription complex subunit 4 | CNOT4 | 1:1 | Block_2 | 2.2933 | 7.7733 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BT09 | Protein canopy homolog 3 | CNPY3 | 1:1 | Block_2 | 3.5408 | 8.1941 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O75128 | Protein cordon-bleu | COBL | 1:1 | Block_2 | 2.4308 | 13.6924 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q53SF7 | Cordon-bleu protein-like 1 | COBLL1 | 1:1 | Block_2 | 1.4751 | 10.7132 | TRUE | 11.2775 | 26.5461 | NA | NA | 1.8254 | NA | NA | 2.0154
P83436 | Conserved oligomeric Golgi complex subunit 7 | COG7 | 1:1 | Block_2 | 0.94007 | 6.2565 | FALSE | NA | NA | NA | NA | 1.2973 | NA | NA | 1.6925
Q6ZST2 | Putative uncharacterized protein COL25A1-DT | COL25A1-DT | 1:1 | Block_2 | 0.66813 | 10.8618 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.96426
Q96A83 | Collagen alpha-1(XXVI) chain | COL26A1 | 1:1 | Block_2 | -0.16336 | 9.2269 | TRUE | 13.2984 | 19.1381 | NA | 0.17225 | 2.6004 | NA | 0.23139 | 1.7909
A6NMZ7 | Collagen alpha-6(VI) chain | COL6A6 | 1:1 | Block_2 | 0.32755 | 8.3261 | FALSE | NA | NA | NA | NA | 0.5041 | NA | NA | 0.90995
P61923 | Coatomer subunit zeta-1 | COPZ1 | 1:1 | Block_2 | 1.3526 | 9.7667 | FALSE | NA | NA | NA | NA | 4.8051 | 2.0779 | 4.5305 | 6.5752
Q9Y2Z9 | Ubiquinone biosynthesis monooxygenase COQ6, mitochondrial | COQ6 | 1:1 | Block_2 | 0.38721 | 11.192 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.2792
Q8WVH0 | Complexin-3 | CPLX3 | 1:1 | Block_2 | 0.5928 | 9.475 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6NV74 | CRACD-like protein | CRACDL | 1:1 | Block_2 | 1.5124 | 13.0168 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.894
Q8N4Y2 | EF-hand calcium-binding domain-containing protein 4A | CRACR2B | 1:1 | Block_2 | 0.64165 | 15.6096 | TRUE | NA | NA | NA | NA | 0.72224 | NA | NA | 1.1752
Q96BA8 | Cyclic AMP-responsive element-binding protein 3-like protein 1 | CREB3L1 | 1:1 | Block_2 | 1.1621 | 5.9154 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.3474
P34998 | Corticotropin-releasing factor receptor 1 | CRHR1 | 1:1 | Block_2 | 4.826 | 9.4054 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9HC73 | Cytokine receptor-like factor 2 | CRLF2 | 1:1 | Block_2 | 1.5954 | 8.1776 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.3089
Q6UUV9 | CREB-regulated transcription coactivator 1 | CRTC1 | 1:1 | Block_2 | 1.8714 | 11.2226 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.3313
Q68DQ2 | Very large A-kinase anchor protein | CRYBG3 | 1:1 | Block_2 | 2.1294 | 12.4904 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.1751
P47710 | Alpha-S1-casein | CSN1S1 | 1:1 | Block_2 | 1.0961 | 8.5077 | FALSE | NA | 1.816 | NA | NA | NA | NA | NA | 1.4783
Q1MSJ5 | Centrosome and spindle pole-associated protein 1 | CSPP1 | 1:1 | Block_2 | 1 | 12.769 | TRUE | NA | NA | NA | NA | 1.3801 | NA | 1.0977 | 3.0878
Q5H943 | Kita-kyushu lung cancer antigen 1 | CT83 | 1:1 | Block_2 | 2.4931 | 11.7107 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IX95 | Putative cTAGE family member 3 | CTAGE3P | 1:1 | Block_2 | 0.78761 | 11.0735 | FALSE | NA | NA | NA | NA | NA | NA | NA | 0.94934
Q05D32 | CTD small phosphatase-like protein 2 | CTDSPL2 | 1:1 | Block_2 | 2.2489 | 10.9374 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P26232 | Catenin alpha-2 | CTNNA2 | 1:1 | Block_2 | 2.6761 | 12.9029 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WYA6 | Beta-catenin-like protein 1 | CTNNBL1 | 1:1 | Block_2 | 0.95632 | 12.3699 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.6467
Q9NWM3 | CUE domain-containing protein 1 | CUEDC1 | 1:1 | Block_2 | 1.5741 | 12.2242 | FALSE | NA | NA | NA | NA | 1.7365 | NA | 1.7194 | 3.1677
Q13620 | Cullin-4B | CUL4B | 1:1 | Block_2 | 1.1709 | 10.9445 | TRUE | 13.8115 | 8.2333 | NA | NA | 1.57 | NA | NA | 1.4312
O14529 | Homeobox protein cut-like 2 | CUX2 | 1:1 | Block_2 | 0.27669 | 14.5705 | TRUE | NA | NA | NA | NA | 1.5086 | NA | NA | 1.1513
Q9HCG8 | Pre-mRNA-splicing factor CWC22 homolog | CWC22 | 1:1 | Block_2 | 0.15742 | 11.652 | TRUE | 12.0818 | 13.2913 | NA | 0.2537 | 1.0559 | NA | 0.32818 | 2.251
A8MYA2 | Uncharacterized protein CXorf49 | CXorf49 | 1:1 | Block_2 | 2.9525 | 12.5275 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NQC7 | Ubiquitin carboxyl-terminal hydrolase CYLD | CYLD | 1:1 | Block_2 | 1.8304 | 11.4592 | TRUE | NA | NA | NA | NA | 2.2865 | NA | 2.4926 | 3.9755
P05177 | Cytochrome P450 1A2 | CYP1A2 | 1:1 | Block_2 | 1.5489 | 11.2227 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6V0L0 | Cytochrome P450 26C1 | CYP26C1 | 1:1 | Block_2 | 1.1636 | 14.0542 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P13584 | Cytochrome P450 4B1 | CYP4B1 | 1:1 | Block_2 | 1.3881 | 10.6915 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q99418 | Cytohesin-2 | CYTH2 | 1:1 | Block_2 | 0.94917 | 5.0567 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NYF0 | Dapper homolog 1 | DACT1 | 1:1 | Block_2 | 1.8656 | 9.5663 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NQZ3 | Deleted in azoospermia protein 1 | DAZ1 | 1:1 | Block_2 | 0.98932 | 11.6818 | FALSE | NA | NA | NA | NA | 1.4545 | NA | NA | NA
Q10586 | D site-binding protein | DBP | 1:1 | Block_2 | -1.3339 | 12.7137 | TRUE | 13.9477 | 20.2756 | NA | NA | -0.59638 | NA | NA | 0.44699
Q6ZNG2 | Homeobox protein DBX2 | DBX2 | 1:1 | Block_2 | 1.341 | 13.3201 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4461
Q8N568 | Serine/threonine-protein kinase DCLK2 | DCLK2 | 1:1 | Block_2 | -0.17631 | 9.4645 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.43608
Q9NPI6 | mRNA-decapping enzyme 1A | DCP1A | 1:1 | Block_2 | 0.46241 | 10.437 | TRUE | 2.9335 | 2.0416 | NA | 0.58381 | 1.3953 | NA | 1.4636 | 3.1763
Q8IU60 | m7GpppN-mRNA hydrolase | DCP2 | 1:1 | Block_2 | 1.9593 | 12.0965 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.1111
Q8IXT1 | DNA damage-induced apoptosis suppressor protein | DDIAS | 1:1 | Block_2 | 1.8977 | 10.4401 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q13206 | Probable ATP-dependent RNA helicase DDX10 | DDX10 | 1:1 | Block_2 | -1.3105 | 11.8933 | TRUE | NA | NA | NA | NA | NA | NA | NA | -0.29565
Q9UHI6 | Probable ATP-dependent RNA helicase DDX20 | DDX20 | 1:1 | Block_2 | 1.0283 | 9.9015 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.5127
Q9NR30 | Nucleolar RNA helicase 2 | DDX21 | 1:1 | Block_2 | 1.5256 | 11.7006 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.3672
Q9NXZ2 | Probable ATP-dependent RNA helicase DDX43 | DDX43 | 1:1 | Block_2 | 2.2507 | 11.6733 | FALSE | NA | NA | NA | NA | 2.5307 | NA | NA | NA
Q8NHQ9 | ATP-dependent RNA helicase DDX55 | DDX55 | 1:1 | Block_2 | 0.38062 | 12.651 | TRUE | NA | 3.6612 | NA | NA | 0.72957 | NA | 0.4705 | 2.589
Q5T1V6 | Probable ATP-dependent RNA helicase DDX59 | DDX59 | 1:1 | Block_2 | 1.4523 | 11.3375 | TRUE | NA | NA | NA | NA | 1.5046 | NA | NA | 1.6517
Q8IY21 | Probable ATP-dependent RNA helicase DDX60 | DDX60 | 1:1 | Block_2 | 1.8131 | 13.7773 | TRUE | NA | NA | NA | NA | 2.6145 | NA | 2.1773 | 4.3029
Q5H9U9 | Probable ATP-dependent RNA helicase DDX60-like | DDX60L | 1:1 | Block_2 | 0.5413 | 12.0204 | TRUE | NA | NA | NA | NA | 1.9049 | NA | NA | 1.5775
Q5J5C9 | Beta-defensin 121 | DEFB121 | 1:1 | Block_2 | 0.44387 | 9.2447 | FALSE | NA | NA | NA | NA | 1.2688 | NA | NA | 0.79517
Q68D51 | DENN domain-containing protein 2C | DENND2C | 1:1 | Block_2 | 0.93636 | 10.354 | TRUE | NA | NA | NA | NA | 1.5257 | NA | 1.362 | 3.5393
Q16760 | Diacylglycerol kinase delta | DGKD | 1:1 | Block_2 | 1.1895 | 9.75 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.1986
Q86XP1 | Diacylglycerol kinase eta | DGKH | 1:1 | Block_2 | 0.30355 | 15.1098 | FALSE | 10.8308 | 15.3361 | NA | NA | 0.99394 | NA | 0.75916 | 2.5607
Q9H2U1 | ATP-dependent DNA/RNA helicase DHX36 | DHX36 | 1:1 | Block_2 | 2.1464 | 13.756 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.074
Q8IY37 | Probable ATP-dependent RNA helicase DHX37 | DHX37 | 1:1 | Block_2 | 1.66 | 11.8016 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q14562 | ATP-dependent RNA helicase DHX8 | DHX8 | 1:1 | Block_2 | 0.85296 | 9.9626 | TRUE | NA | NA | NA | NA | 0.99081 | NA | NA | 1.0189
Q9NSV4 | Protein diaphanous homolog 3 | DIAPH3 | 1:1 | Block_2 | 2.2115 | 11.8736 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.2502
Q9P265 | Disco-interacting protein 2 homolog B | DIP2B | 1:1 | Block_2 | 2.2733 | 7.8718 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.2882
Q8TF46 | DIS3-like exonuclease 1 | DIS3L | 1:1 | Block_2 | 0.90587 | 11.8144 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.1083
P36957 | Dihydrolipoyllysine-residue succinyltransferase component of 2-oxoglutarate dehydrogenase complex, mitochondrial | DLST | 1:1 | Block_2 | 0.58635 | 4.071 | FALSE | NA | 2.5481 | NA | NA | 1.1951 | NA | 0.73681 | 1.5756
Q9NPF5 | DNA methyltransferase 1-associated protein 1 | DMAP1 | 1:1 | Block_2 | 2.055 | 10.9462 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q14565 | Meiotic recombination protein DMC1/LIM15 homolog | DMC1 | 1:1 | Block_2 | 1.8875 | 12.002 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.908
Q8IYS4 | Dynein axonemal assembly factor 8 | DNAAF8 | 1:1 | Block_2 | 3.1861 | 14.0036 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P2D7 | Dynein axonemal heavy chain 1 | DNAH1 | 1:1 | Block_2 | 1.7923 | 13.9212 | TRUE | NA | NA | NA | NA | 1.9331 | NA | NA | NA
Q8IVF4 | Dynein axonemal heavy chain 10 | DNAH10 | 1:1 | Block_2 | 1.6798 | 12.6285 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UFH2 | Dynein axonemal heavy chain 17 | DNAH17 | 1:1 | Block_2 | 0.072548 | 11.7055 | TRUE | NA | NA | NA | NA | 1.1626 | NA | NA | 0.59154
Q9C0G6 | Dynein axonemal heavy chain 6 | DNAH6 | 1:1 | Block_2 | -5.2154 | 6.6494 | FALSE | 16.9823 | 14.6967 | NA | -4.2768 | -1.4957 | NA | -4.6267 | -1.2537
Q5VTH9 | Dynein axonemal intermediate chain 4 | DNAI4 | 1:1 | Block_2 | 3.762 | 13.9478 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6TDU7 | Dynein axonemal intermediate chain 7 | DNAI7 | 1:1 | Block_2 | 1.5319 | 12.7312 | TRUE | NA | NA | NA | NA | 1.697 | NA | NA | NA
P59910 | DnaJ homolog subfamily B member 13 | DNAJB13 | 1:1 | Block_2 | 2.2494 | 12.9797 | FALSE | 18.0194 | 43.3061 | NA | NA | NA | NA | NA | NA
Q8IXB1 | DnaJ homolog subfamily C member 10 | DNAJC10 | 1:1 | Block_2 | 1.9518 | 11.5544 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UKB3 | DnaJ homolog subfamily C member 12 | DNAJC12 | 1:1 | Block_2 | 1.5455 | 12.3557 | TRUE | 19.0255 | 16.2306 | NA | NA | 2.3642 | NA | 1.9032 | 3.5548
Q96DA6 | Mitochondrial import inner membrane translocase subunit TIM14 | DNAJC19 | 1:1 | Block_2 | 1.3771 | 8.782 | TRUE | NA | NA | NA | NA | 1.5661 | NA | NA | 2.0127
Q9NX36 | DnaJ homolog subfamily C member 28 | DNAJC28 | 1:1 | Block_2 | 1.7722 | 15.1674 | TRUE | 4.7584 | 30.5856 | NA | NA | 2.3545 | NA | NA | 2.9703
Q4LDG9 | Dynein axonemal light chain 1 | DNAL1 | 1:1 | Block_2 | 0.3591 | 10.8066 | TRUE | 12.8361 | 11.0355 | NA | NA | 0.49546 | NA | NA | 1.7321
Q96BY6 | Dedicator of cytokinesis protein 10 | DOCK10 | 1:1 | Block_2 | 2.5816 | 15.1188 | TRUE | 31.8462 | NA | NA | NA | NA | NA | NA | NA
Q5T8R8 | Uncharacterized protein DOCK8-AS1 | DOCK8-AS1 | 1:1 | Block_2 | 1.1909 | 11.3321 | TRUE | 7.94 | 9.1077 | NA | NA | NA | NA | NA | 2.0483
Q92782 | Zinc finger protein neuro-d4 | DPF1 | 1:1 | Block_2 | 3.9908 | 12.0945 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q92784 | Zinc finger protein DPF3 | DPF3 | 1:1 | Block_2 | 2 | 11.3958 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H4G8 | Putative DPH3 homolog B | DPH3P1 | 1:1 | Block_2 | 1.6878 | 11.4021 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8058
P14416 | Dopamine D2 receptor | DRD2 | 1:1 | Block_2 | 2.5 | 10.2833 | TRUE | 14.0297 | 7.0133 | NA | NA | NA | NA | NA | 2.9648
Q9NRR4 | Ribonuclease 3 | DROSHA | 1:1 | Block_2 | 1.7881 | 12.2579 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.2163
Q9BVC3 | Sister chromatid cohesion protein DCC1 | DSCC1 | 1:1 | Block_2 | -0.24795 | 10.9406 | TRUE | NA | NA | NA | NA | 0.07923 | NA | NA | 0.16729
Q9NZW4 | Dentin sialophosphoprotein | DSPP | 1:1 | Block_2 | 2.2154 | 9.3361 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.4275
Q9NZJ0 | Denticleless protein homolog | DTL | 1:1 | Block_2 | 2.4859 | 12.4021 | FALSE | 3.1 | 25.2024 | NA | NA | 2.5075 | NA | NA | 3.208
Q9NRD8 | Dual oxidase 2 | DUOX2 | 1:1 | Block_2 | 1.5 | 7.5914 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H1R2 | Dual specificity protein phosphatase 15 | DUSP15 | 1:1 | Block_2 | 1.0113 | 8.4751 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.0647
Q8WTR2 | Dual specificity protein phosphatase 19 | DUSP19 | 1:1 | Block_2 | 0.96653 | 8.7913 | FALSE | NA | NA | NA | NA | 1.3802 | NA | 1.1311 | 2.5517
Q13409 | Cytoplasmic dynein 1 intermediate chain 2 | DYNC1I2 | 1:1 | Block_2 | 1.6428 | 13.8992 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.9079
Q8NCM8 | Cytoplasmic dynein 2 heavy chain 1 | DYNC2H1 | 1:1 | Block_2 | 1.9488 | 9.0808 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.9633
Q8WVS4 | Cytoplasmic dynein 2 intermediate chain 1 | DYNC2I1 | 1:1 | Block_2 | 1.2501 | 12.0506 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TCX1 | Cytoplasmic dynein 2 light intermediate chain 1 | DYNC2LI1 | 1:1 | Block_2 | 1.7385 | 12.6768 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IYY4 | Cilium assembly protein DZIP1L | DZIP1L | 1:1 | Block_2 | 2.2493 | 16.013 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.6855
Q86Y13 | E3 ubiquitin-protein ligase DZIP3 | DZIP3 | 1:1 | Block_2 | 1.3759 | 13.2636 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96CJ1 | ELL-associated factor 2 | EAF2 | 1:1 | Block_2 | 1.2705 | 12.898 | TRUE | 5.2392 | NA | NA | NA | NA | NA | NA | 1.7841
Q56P03 | E2F-associated phosphoprotein | EAPP | 1:1 | Block_2 | 1.5539 | 12.839 | TRUE | NA | NA | NA | NA | 1.5571 | NA | NA | 2.2185
Q9BQ95 | Evolutionarily conserved signaling intermediate in Toll pathway, mitochondrial | ECSIT | 1:1 | Block_2 | -0.15242 | 11.0616 | TRUE | 10.7523 | 14.5906 | NA | 0.8117 | 3.1053 | 0.029555 | 1.6428 | 3.7305
Q9H8V3 | Protein ECT2 | ECT2 | 1:1 | Block_2 | 1.6842 | 12.7647 | TRUE | 18.5491 | 11.6484 | NA | NA | NA | NA | NA | 1.7017
P24530 | Endothelin receptor type B | EDNRB | 1:1 | Block_2 | 1.25 | 10.4735 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.0497
Q5JPI9 | EEF1A lysine methyltransferase 2 | EEF1AKMT2 | 1:1 | Block_2 | 2.9479 | 11.2946 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5T9C2 | Protein FAM102A | EEIG1 | 1:1 | Block_2 | 0.001 | 11.5649 | TRUE | 14.0886 | 19.502 | NA | 0.0025786 | 1.1008 | NA | 0.69959 | 2.5322
Q5T8I3 | Protein FAM102B | EEIG2 | 1:1 | Block_2 | 0.94372 | 9.5704 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.556
Q7L9B9 | Endonuclease/exonuclease/phosphatase family domain-containing protein 1 | EEPD1 | 1:1 | Block_2 | 0.56943 | 8.5158 | FALSE | NA | NA | NA | NA | 0.5876 | NA | NA | 2.2127
A6NFE3 | EF-hand calcium-binding domain-containing protein 10 | EFCAB10 | 1:1 | Block_2 | 0.096855 | 10.8232 | FALSE | 16.2662 | 21.4656 | NA | NA | 1.085 | NA | NA | 1.2714
Q6NXP0 | EF-hand calcium-binding domain-containing protein 12 | EFCAB12 | 1:1 | Block_2 | 1.7082 | 14.0442 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A8MWE9 | EF-hand calcium-binding domain-containing protein 8 | EFCAB8 | 1:1 | Block_2 | 3.1259 | 13.6893 | TRUE | NA | NA | NA | NA | NA | NA | NA | 4.1727
Q9HA90 | EF-hand and coiled-coil domain-containing protein 1 | EFCC1 | 1:1 | Block_2 | 2.9578 | 12.7084 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.3001
Q9NZN4 | EH domain-containing protein 2 | EHD2 | 1:1 | Block_2 | 2.2889 | 12.292 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P41214 | Eukaryotic translation initiation factor 2D | EIF2D | 1:1 | Block_2 | 0.060385 | 13.074 | TRUE | NA | NA | NA | 1.2969 | 3.6291 | 1.4965 | 3.8392 | 5.1486
O75821 | Eukaryotic translation initiation factor 3 subunit G | EIF3G | 1:1 | Block_2 | 2.6613 | 13.5281 | TRUE | NA | NA | NA | NA | NA | NA | NA | 4.1737
Q9NRA8 | Eukaryotic translation initiation factor 4E transporter | EIF4ENIF1 | 1:1 | Block_2 | 2.2958 | 10.0864 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9GZV4 | Eukaryotic translation initiation factor 5A-2 | EIF5A2 | 1:1 | Block_2 | 1.0416 | 9.0679 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5219
Q53HC9 | EARP and GARP complex-interacting protein 1 | EIPR1 | 1:1 | Block_2 | 1.6933 | 11.7952 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.2878
Q14576 | ELAV-like protein 3 | ELAVL3 | 1:1 | Block_2 | 1.2006 | 9.2673 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.2685
Q6PCB8 | Embigin | EMB | 1:1 | Block_2 | -2.263 | 5.4556 | TRUE | NA | NA | NA | NA | NA | NA | NA | -1.6995
Q92979 | Ribosomal RNA small subunit methyltransferase NEP1 | EMG1 | 1:1 | Block_2 | 1.4805 | 2.7427 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.6423
Q6ZMW3 | Echinoderm microtubule-associated protein-like 6 | EML6 | 1:1 | Block_2 | 1.812 | 11.6823 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q05925 | Homeobox protein engrailed-1 | EN1 | 1:1 | Block_2 | -0.05389 | 11.2132 | FALSE | NA | NA | NA | NA | NA | NA | NA | 0.644
Q8IYW4 | ENTH domain-containing protein 1 | ENTHD1 | 1:1 | Block_2 | 5.464 | 15.5088 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O75355 | Ectonucleoside triphosphate diphosphohydrolase 3 | ENTPD3 | 1:1 | Block_2 | 0.60069 | 11.5975 | TRUE | NA | NA | NA | NA | 0.85029 | NA | NA | 1.4647
Q9Y227 | Ectonucleoside triphosphate diphosphohydrolase 4 | ENTPD4 | 1:1 | Block_2 | 1.9337 | 12.3574 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NPA8 | Transcription and mRNA export factor ENY2 | ENY2 | 1:1 | Block_2 | -0.026286 | 11.7384 | FALSE | 16.8542 | 9.9248 | NA | 0.81916 | 1.4657 | 0.95503 | 2.0875 | 4.0277
Q9HCS5 | Band 4.1-like protein 4A | EPB41L4A | 1:1 | Block_2 | -1.537 | 7.0511 | FALSE | NA | NA | NA | NA | NA | NA | NA | -0.79621
Q52LR7 | Enhancer of polycomb homolog 2 | EPC2 | 1:1 | Block_2 | 3.2869 | 13.6099 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H201 | Epsin-3 | EPN3 | 1:1 | Block_2 | 0.52762 | 9.6781 | FALSE | 11.5992 | 12.94 | NA | NA | 2.0688 | NA | NA | 1.6389
Q92889 | DNA repair endonuclease XPF | ERCC4 | 1:1 | Block_2 | 2 | 10.2371 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.0196
Q86X53 | Glutamate-rich protein 1 | ERICH1 | 1:1 | Block_2 | 0.53626 | 11.4548 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.1812
Q5RHP9 | Glutamate-rich protein 3 | ERICH3 | 1:1 | Block_2 | 0.33016 | 11.4768 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.54053
A6NGS2 | Glutamate-rich protein 4 | ERICH4 | 1:1 | Block_2 | -0.064256 | 10.5177 | TRUE | NA | NA | NA | NA | 0.27449 | NA | NA | 0.43808
Q7L0X2 | Glutamate-rich protein 6 | ERICH6 | 1:1 | Block_2 | 1.9636 | 11.4905 | FALSE | NA | NA | NA | NA | 2.0993 | NA | NA | NA
Q9UJM3 | ERBB receptor feedback inhibitor 1 | ERRFI1 | 1:1 | Block_2 | 4.6151 | 14.3269 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O71037 | Endogenous retrovirus group K member 19 Env polyprotein | ERVK-19 | 1:1 | Block_2 | 0.97431 | 9.7002 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.046
Q9NVF9 | Ethanolamine kinase 2 | ETNK2 | 1:1 | Block_2 | 2.1456 | 11.2899 | TRUE | NA | NA | NA | NA | 2.1561 | NA | NA | 2.4364
Q9H790 | Exonuclease V | EXO5 | 1:1 | Block_2 | -0.20062 | 13.961 | TRUE | NA | NA | NA | NA | 0.91928 | NA | NA | 0.51913
Q9UPT5 | Exocyst complex component 7 | EXOC7 | 1:1 | Block_2 | 0.74345 | 10.7868 | FALSE | NA | 11.7517 | NA | 0.90953 | 2.4401 | 1.2351 | 2.603 | 4.722
Q9NPD3 | Exosome complex component RRP41 | EXOSC4 | 1:1 | Block_2 | 0.82976 | 5.8047 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.1232
O43909 | Exostosin-like 3 | EXTL3 | 1:1 | Block_2 | 1.7061 | 10.2591 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.8789
Q6NZ36 | Fanconi anemia core complex-associated protein 20 | FAAP20 | 1:1 | Block_2 | 1.1533 | 8.5087 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.1949
Q5XKR9 | Protein FAM104B | FAM104B | 1:1 | Block_2 | 1.2264 | 12.709 | TRUE | 3.0594 | 1.5596 | NA | NA | NA | NA | NA | NA
Q86XD5 | Protein FAM131B | FAM131B | 1:1 | Block_2 | 2.8923 | 10.6829 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96AQ9 | Protein FAM131C | FAM131C | 1:1 | Block_2 | 0.5354 | 12.1142 | FALSE | 0.91408 | 10.041 | NA | NA | 0.89975 | NA | NA | 1.0218
Q49AJ0 | Protein FAM135B | FAM135B | 1:1 | Block_2 | 0.059293 | 13.5084 | TRUE | NA | NA | NA | NA | 0.42882 | NA | 0.10025 | 1.2484
Q9NYF5 | Protein FAM13B | FAM13B | 1:1 | Block_2 | 2.151 | 12.2768 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.5186
Q3B820 | Protein FAM161A | FAM161A | 1:1 | Block_2 | 0.31796 | 14.4714 | TRUE | NA | NA | NA | NA | 0.85013 | NA | NA | 1.2565
Q96MY7 | Protein FAM161B | FAM161B | 1:1 | Block_2 | 0.69874 | 14.022 | TRUE | 27.0945 | 25.9625 | NA | NA | 1.5613 | NA | NA | 1.6828
Q8NB25 | Protein FAM184A | FAM184A | 1:1 | Block_2 | 0.81777 | 11.8885 | TRUE | 19.4449 | 23.2285 | NA | NA | 1.959 | NA | NA | 1.7504
A6NE01 | Protein FAM186A | FAM186A | 1:1 | Block_2 | 0.27192 | 12.7485 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P78312 | Protein FAM193A | FAM193A | 1:1 | Block_2 | 1.3926 | 13.7166 | TRUE | 7.8499 | 8.6147 | NA | NA | 2.4068 | NA | 1.5158 | 3.6017
Q9H8W3 | Protein FAM204A | FAM204A | 1:1 | Block_2 | 1.6388 | 13.2803 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.2466
Q96M60 | Protein FAM227B | FAM227B | 1:1 | Block_2 | -0.23798 | 11.4537 | TRUE | NA | NA | NA | NA | 0.23157 | NA | NA | 0.097424
Q8NA70 | Protein FAM47B | FAM47B | 1:1 | Block_2 | 1.3443 | 14.992 | TRUE | NA | 5.158 | NA | NA | 1.9918 | NA | NA | 1.6011
Q5TYM5 | Protein FAM72A | FAM72A | 1:1 | Block_2 | 3 | 5.5644 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5T0W9 | Protein FAM83B | FAM83B | 1:1 | Block_2 | 1.7605 | 11.5134 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8413
A6NDY2 | Putative protein FAM90A10 | FAM90A10 | 1:1 | Block_2 | -3.1374 | 7.0868 | FALSE | 11.6497 | 3.597 | NA | NA | -0.97747 | NA | NA | -1.0788
Q658Y4 | Protein FAM91A1 | FAM91A1 | 1:1 | Block_2 | 0.79947 | 12.5563 | TRUE | 8.1652 | 19.5579 | NA | NA | NA | NA | 0.89079 | 2.5214
Q8TC84 | Fibronectin type 3 and ankyrin repeat domains protein 1 | FANK1 | 1:1 | Block_2 | 0.083476 | 8.6442 | FALSE | 9.5725 | 8.1219 | NA | NA | 0.86226 | NA | NA | 1.15
Q6V0I7 | Protocadherin Fat 4 | FAT4 | 1:1 | Block_2 | 1 | 5.0879 | TRUE | 17.1404 | 10.1863 | NA | NA | 1.3294 | NA | NA | 1.6688
Q9HCM7 | Fibrosin-1-like protein | FBRSL1 | 1:1 | Block_2 | 1.1253 | 12.8879 | TRUE | NA | NA | NA | NA | 1.4806 | NA | NA | 1.7601
Q9UKA1 | F-box/LRR-repeat protein 5 | FBXL5 | 1:1 | Block_2 | 1.2325 | 9.9813 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IX29 | F-box only protein 16 | FBXO16 | 1:1 | Block_2 | 0.31104 | 11.1179 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.60968
Q8NEZ5 | F-box only protein 22 | FBXO22 | 1:1 | Block_2 | 1.0723 | 10.0099 | FALSE | NA | 8.3362 | NA | NA | 1.1343 | NA | NA | 1.7869
Q6PJ61 | F-box only protein 46 | FBXO46 | 1:1 | Block_2 | 1.7121 | 13.4384 | FALSE | 14.7793 | 12.7964 | NA | NA | NA | NA | NA | 2.0826
Q86WN1 | F-BAR and double SH3 domains protein 1 | FCHSD1 | 1:1 | Block_2 | 2.0137 | 9.4095 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UHY8 | Fasciculation and elongation protein zeta-2 | FEZ2 | 1:1 | Block_2 | 2.8033 | 14.8887 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.7952
Q7Z6J4 | FYVE, RhoGEF and PH domain-containing protein 2 | FGD2 | 1:1 | Block_2 | 1.5414 | 10.1276 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZV73 | FYVE, RhoGEF and PH domain-containing protein 6 | FGD6 | 1:1 | Block_2 | 0.25071 | 11.488 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N612 | FHF complex subunit HOOK interacting protein 1B | FHIP1B | 1:1 | Block_2 | 2.5563 | 11.9034 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5HY92 | Fidgetin | FIGN | 1:1 | Block_2 | 0.241 | 10.3971 | TRUE | NA | NA | NA | NA | 1.017 | NA | 0.31342 | 1.7846
Q6PIW4 | Fidgetin-like protein 1 | FIGNL1 | 1:1 | Block_2 | 0.5397 | 10.1634 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NFF5 | FAD synthase | FLAD1 | 1:1 | Block_2 | 1.1167 | 7.2433 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.743
Q8IVF7 | Formin-like protein 3 | FMNL3 | 1:1 | Block_2 | 2.4497 | 13.396 | FALSE | NA | NA | NA | NA | 2.945 | NA | NA | 4.1639
P53539 | Protein FosB | FOSB | 1:1 | Block_2 | 1.721 | 11.2409 | FALSE | NA | 6.1371 | NA | NA | NA | NA | NA | 2.006
Q01167 | Forkhead box protein K2 | FOXK2 | 1:1 | Block_2 | -0.64979 | 12.5904 | TRUE | 11.4942 | 11.1668 | NA | -0.50023 | 0.10355 | NA | 0.33525 | 1.8223
Q5SZK8 | FRAS1-related extracellular matrix protein 2 | FREM2 | 1:1 | Block_2 | 1.2274 | 7.8569 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.248
Q8N878 | FERM domain-containing protein 1 | FRMD1 | 1:1 | Block_2 | 2.5532 | 12.6006 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A1L4K1 | Fibronectin type III and SPRY domain-containing protein 2 | FSD2 | 1:1 | Block_2 | 1.7029 | 11.517 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8354
Q8NA03 | Fibrous sheath-interacting protein 1 | FSIP1 | 1:1 | Block_2 | 1.5529 | 12.694 | TRUE | NA | NA | NA | NA | 1.7661 | NA | NA | 3.237
P02792 | Ferritin light chain | FTL | 1:1 | Block_2 | 1.0754 | 8.7579 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.8777
Q5VWT5 | FYN-binding protein 2 | FYB2 | 1:1 | Block_2 | 0.66379 | 12.1559 | TRUE | 12.4925 | 22.2634 | NA | 1.124 | 2.6022 | NA | 1.8953 | 3.7622
Q13480 | GRB2-associated-binding protein 1 | GAB1 | 1:1 | Block_2 | 1.6847 | 11.0413 | TRUE | NA | NA | NA | 1.7646 | 3.1333 | 1.7621 | 3.375 | 5.0746
O00591 | Gamma-aminobutyric acid receptor subunit pi | GABRP | 1:1 | Block_2 | 1.5503 | 9.6179 | TRUE | NA | NA | NA | NA | 3.3661 | NA | NA | 1.911
O95257 | Growth arrest and DNA damage-inducible protein GADD45 gamma | GADD45G | 1:1 | Block_2 | -1.5059 | 7.5886 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O76087 | G antigen 7 | GAGE7 | 1:1 | Block_2 | 1.2139 | 9.2062 | TRUE | NA | NA | NA | NA | 1.4862 | NA | NA | 1.6065
Q01415 | N-acetylgalactosamine kinase | GALK2 | 1:1 | Block_2 | 2.4577 | 8.4259 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.6636
P17677 | Neuromodulin | GAP43 | 1:1 | Block_2 | 1.1707 | 10.8303 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4622
Q75VX8 | GRB2-associated and regulator of MAPK protein 2 | GAREM2 | 1:1 | Block_2 | 1.7526 | 15.6864 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.842
Q96KD3 | Golgi-associated RAB2 interactor protein 1B | GARIN1B | 1:1 | Block_2 | -0.65915 | 9.3594 | TRUE | 4.2487 | NA | NA | NA | -0.31838 | NA | NA | -0.19309
Q8N9W8 | Golgi-associated RAB2 interactor protein 2 | GARIN2 | 1:1 | Block_2 | 0.79316 | 12.6947 | TRUE | NA | NA | NA | NA | 1.7413 | NA | NA | 1.768
Q8N5Q1 | Golgi-associated RAB2 interactor protein 5B | GARIN5B | 1:1 | Block_2 | -0.48691 | 12.6847 | TRUE | 7.5135 | 14.8934 | NA | NA | -0.19599 | NA | NA | 1.2315
O15063 | Granule associated Rac and RHOG effector protein 1 | GARRE1 | 1:1 | Block_2 | 0.70586 | 12.083 | FALSE | NA | NA | NA | NA | 1.5583 | NA | NA | 1.7434
P54826 | Growth arrest-specific protein 1 | GAS1 | 1:1 | Block_2 | 0.78309 | 8.5581 | TRUE | 10.8268 | 20.2648 | NA | NA | 1.9451 | NA | NA | 1.9299
Q86XJ1 | GAS2-like protein 3 | GAS2L3 | 1:1 | Block_2 | 1.1963 | 11.5631 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.3671
O60861 | Growth arrest-specific protein 7 | GAS7 | 1:1 | Block_2 | -0.90912 | 7.8918 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.9303
O43716 | Glutamyl-tRNA | GATC | 1:1 | Block_2 | 2.3574 | 9.4018 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H227 | Cytosolic beta-glucosidase | GBA3 | 1:1 | Block_2 | 0.25612 | 11.3713 | TRUE | NA | NA | NA | NA | 0.8778 | NA | NA | 1.1808
Q9H0R5 | Guanylate-binding protein 3 | GBP3 | 1:1 | Block_2 | 1.4097 | 11.9119 | TRUE | 22.6778 | 3.5163 | NA | NA | 2.0823 | NA | 2.2614 | 4.1936
O75600 | 2-amino-3-ketobutyrate coenzyme A ligase, mitochondrial | GCAT | 1:1 | Block_2 | -1.2338 | 3.9418 | FALSE | NA | NA | NA | NA | -1.0219 | NA | NA | -0.86667
P35557 | Hexokinase-4 | GCK | 1:1 | Block_2 | -0.96327 | 6.3335 | FALSE | 11.3047 | 26.7293 | NA | -0.82127 | 0.45335 | NA | -0.40882 | 1.2933
Q9P109 | Beta-1,3-galactosyl-O-glycosyl-glycoprotein beta-1,6-N-acetylglucosaminyltransferase 4 | GCNT4 | 1:1 | Block_2 | 1.897 | 10.3951 | FALSE | NA | NA | NA | NA | 2.378 | NA | NA | NA
Q8N6F7 | Germinal center-associated signaling and motility protein | GCSAM | 1:1 | Block_2 | 1.8555 | 13.054 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.0545
Q9NXN4 | Ganglioside-induced differentiation-associated protein 2 | GDAP2 | 1:1 | Block_2 | 1.9945 | 10.9375 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q969S9 | Ribosome-releasing factor 2, mitochondrial | GFM2 | 1:1 | Block_2 | 0.96968 | 11.4389 | FALSE | 18.0975 | 13.272 | NA | NA | NA | NA | NA | 2.0952
O94808 | Glutamine--fructose-6-phosphate aminotransferase [isomerizing] 2 | GFPT2 | 1:1 | Block_2 | 1.5192 | 13.7074 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O75420 | GRB10-interacting GYF protein 1 | GIGYF1 | 1:1 | Block_2 | 1.0474 | 11.6581 | TRUE | 7.7372 | 8.03 | NA | NA | 1.7501 | NA | NA | 2.3409
Q9UG22 | GTPase IMAP family member 2 | GIMAP2 | 1:1 | Block_2 | 2.0782 | 9.6129 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6P9H5 | GTPase IMAP family member 6 | GIMAP6 | 1:1 | Block_2 | 2.624 | 10.9803 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BRT9 | DNA replication complex GINS protein SLD5 | GINS4 | 1:1 | Block_2 | 2.5 | 11.6871 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NVN8 | Guanine nucleotide-binding protein-like 3-like protein | GNL3L | 1:1 | Block_2 | 2.6105 | 12.8611 | FALSE | NA | NA | NA | NA | 3.5195 | NA | 2.8357 | 4.769
Q08379 | Golgin subfamily A member 2 | GOLGA2 | 1:1 | Block_2 | 2 | 11.7507 | FALSE | NA | NA | NA | NA | 2.125 | NA | NA | 3.7363
Q9HBQ8 | Putative golgin subfamily A member 2B | GOLGA2P5 | 1:1 | Block_2 | 2.5365 | 13.616 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q13439 | Golgin subfamily A member 4 | GOLGA4 | 1:1 | Block_2 | 0.50977 | 9.8097 | TRUE | 10.4814 | 6.0792 | NA | NA | 0.61341 | NA | NA | 1.7329
A8MQT2 | Golgin subfamily A member 8B | GOLGA8B | 1:1 | Block_2 | 1.5 | 13.2194 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O95872 | G patch domain and ankyrin repeat-containing protein 1 | GPANK1 | 1:1 | Block_2 | 2 | 10.3929 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N954 | G patch domain-containing protein 11 | GPATCH11 | 1:1 | Block_2 | 1 | 11.56 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4596
Q5T3I0 | G patch domain-containing protein 4 | GPATCH4 | 1:1 | Block_2 | 2.5 | 12.4927 | TRUE | 18.7656 | 56.025 | NA | NA | 3.3969 | NA | NA | 5.0035
Q9UKJ3 | G patch domain-containing protein 8 | GPATCH8 | 1:1 | Block_2 | -1.2492 | 8.3319 | FALSE | NA | NA | NA | NA | 1.2619 | NA | -1.0679 | 1.2134
Q8N158 | Glypican-2 | GPC2 | 1:1 | Block_2 | 0.81007 | 10.0181 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.96982
Q5JY77 | G-protein coupled receptor-associated sorting protein 1 | GPRASP1 | 1:1 | Block_2 | -0.03819 | 13.3138 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5395
Q96D09 | G-protein coupled receptor-associated sorting protein 2 | GPRASP2 | 1:1 | Block_2 | 1.5028 | 10.8921 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.0429
Q6PI77 | Protein BHLHb9 | GPRASP3 | 1:1 | Block_2 | 1.62 | 11.1272 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O60269 | G protein-regulated inducer of neurite outgrowth 2 | GPRIN2 | 1:1 | Block_2 | 2.5089 | 13.4204 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZVF9 | G protein-regulated inducer of neurite outgrowth 3 | GPRIN3 | 1:1 | Block_2 | 0.69842 | 10.5988 | TRUE | 9.1261 | 2.4769 | NA | NA | 1.6511 | NA | 1.2744 | 2.6621
Q96CP6 | Protein Aster-A | GRAMD1A | 1:1 | Block_2 | 1.3153 | 10.5889 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.6412
Q16478 | Glutamate receptor ionotropic, kainate 5 | GRIK5 | 1:1 | Block_2 | 0.68944 | 11.6885 | TRUE | NA | NA | NA | NA | 1.8764 | NA | NA | 1.7559
P35626 | Beta-adrenergic receptor kinase 2 | GRK3 | 1:1 | Block_2 | 0.63174 | 10.8038 | TRUE | NA | NA | NA | NA | 1.2088 | NA | 0.67868 | 2.0643
Q96QA5 | Gasdermin-A | GSDMA | 1:1 | Block_2 | 2.0713 | 9.5645 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NRJ2 | Putative uncharacterized protein GSN-AS1 | GSN-AS1 | 1:1 | Block_2 | -0.57953 | 11.5027 | TRUE | 0.05952 | 4.6921 | NA | NA | -0.33031 | NA | NA | 0.70989
P15170 | Eukaryotic peptide chain release factor GTP-binding subunit ERF3A | GSPT1 | 1:1 | Block_2 | 1.8228 | 7.9976 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.4183
Q86UP8 | General transcription factor II-I repeat domain-containing protein 2A | GTF2IRD2 | 1:1 | Block_2 | 1.9503 | 12.1479 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WUA4 | General transcription factor 3C polypeptide 2 | GTF3C2 | 1:1 | Block_2 | 0.93846 | 10.4308 | FALSE | 3.8587 | 8.9726 | NA | NA | NA | NA | NA | 1.7037
P43080 | Guanylyl cyclase-activating protein 1 | GUCA1A | 1:1 | Block_2 | 2.4109 | 9.1875 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O95843 | Guanylyl cyclase-activating protein 3 | GUCA1C | 1:1 | Block_2 | 0.069926 | 12.906 | TRUE | NA | NA | NA | NA | 0.68139 | NA | NA | 0.54752
Q8N442 | Translation factor GUF1, mitochondrial | GUF1 | 1:1 | Block_2 | 0.94688 | 12.3793 | FALSE | NA | NA | NA | NA | 2.1515 | NA | 0.94715 | 2.4512
Q3ZCU0 | Protein GVQW3 | GVQW3 | 1:1 | Block_2 | 1.9033 | 10.9843 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P22492 | Histone H1t | H1-6 | 1:1 | Block_2 | -0.86833 | 7.6609 | FALSE | NA | NA | NA | NA | 0.29636 | NA | NA | 0.87103
O96004 | Heart- and neural crest derivatives-expressed protein 1 | HAND1 | 1:1 | Block_2 | 0.76017 | 10.6102 | FALSE | NA | NA | NA | NA | 0.95307 | NA | NA | 1.3592
Q7Z4H7 | HAUS augmin-like complex subunit 6 | HAUS6 | 1:1 | Block_2 | 1.4916 | 11.8168 | TRUE | 16.104 | 18.922 | NA | NA | 1.9518 | NA | NA | 2.822
P69892 | Hemoglobin subunit gamma-2 | HBG2 | 1:1 | Block_2 | 2.7585 | 10.8263 | TRUE | NA | NA | NA | NA | 4.227 | NA | NA | 5.3501
Q9UBN7 | Histone deacetylase 6 | HDAC6 | 1:1 | Block_2 | 0.001 | 10.5097 | TRUE | 45.6817 | 21.6864 | NA | 1.5357 | 4.0754 | 1.3613 | 3.8509 | 7.2726
Q8N4P3 | Guanosine-3',5'-bis(diphosphate) 3'-pyrophosphohydrolase MESH1 | HDDC3 | 1:1 | Block_2 | -0.99095 | 6.3341 | FALSE | NA | NA | NA | NA | -0.62164 | NA | NA | -0.56119
Q7Z353 | Highly divergent homeobox | HDX | 1:1 | Block_2 | 1.4909 | 8.8729 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86WZ0 | HEAT repeat-containing protein 4 | HEATR4 | 1:1 | Block_2 | 2.4268 | 14.1142 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UBI9 | Headcase protein homolog | HECA | 1:1 | Block_2 | -2.224 | 11.113 | FALSE | 12.3859 | 8.3715 | NA | NA | -1.155 | NA | NA | -0.80689
Q5U5R9 | Probable E3 ubiquitin-protein ligase HECTD2 | HECTD2 | 1:1 | Block_2 | 0.43843 | 11.4553 | FALSE | NA | NA | NA | NA | 0.53888 | NA | NA | 1.2169
Q9BYK8 | Helicase with zinc finger domain 2 | HELZ2 | 1:1 | Block_2 | 1.1239 | 12.3359 | TRUE | 7.512 | 17.4342 | NA | NA | 1.636 | NA | NA | 1.9308
O95714 | E3 ubiquitin-protein ligase HERC2 | HERC2 | 1:1 | Block_2 | 1.0945 | 4.7807 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P63121 | Endogenous retrovirus group K member 113 Pro protein | HERVK-113_VP113 | 1:1 | Block_2 | 0.76812 | 11.515 | TRUE | 3.7525 | 12.6529 | NA | NA | NA | NA | NA | 0.80295
Q9BYE0 | Transcription factor HES-7 | HES7 | 1:1 | Block_2 | 1.3867 | 9.5759 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UBP5 | Hairy/enhancer-of-split related with YRPW motif protein 2 | HEY2 | 1:1 | Block_2 | 1.6961 | 7.5961 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IV36 | Protein HID1 | HID1 | 1:1 | Block_2 | 1.5476 | 12.0046 | TRUE | NA | NA | NA | NA | 1.9306 | NA | NA | NA
Q5T1R4 | Transcription factor HIVEP3 | HIVEP3 | 1:1 | Block_2 | -0.45184 | 13.0971 | TRUE | NA | NA | NA | NA | 0.79726 | NA | NA | 0.43234
Q16534 | Hepatic leukemia factor | HLF | 1:1 | Block_2 | 0.24894 | 10.6238 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q86SG4 | Putative Dresden prostate carcinoma protein 2 | HMGN2P46 | 1:1 | Block_2 | 0.30418 | 12.9412 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.71583
P61978 | Heterogeneous nuclear ribonucleoprotein K | HNRNPK | 1:1 | Block_2 | 2.3428 | 13.5707 | TRUE | NA | NA | NA | NA | NA | NA | 2.6031 | 5.505
Q6PI97 | Cilia- and flagella-associated protein HOATZ | HOATZ | 1:1 | Block_2 | 0.93872 | 12.9652 | TRUE | NA | NA | NA | NA | 1.6817 | NA | NA | 1.1164
O43364 | Homeobox protein Hox-A2 | HOXA2 | 1:1 | Block_2 | 1.2333 | 10.0511 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P14651 | Homeobox protein Hox-B3 | HOXB3 | 1:1 | Block_2 | 1.5295 | 10.643 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P17483 | Homeobox protein Hox-B4 | HOXB4 | 1:1 | Block_2 | 1.1273 | 8.4471 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P13378 | Homeobox protein Hox-D8 | HOXD8 | 1:1 | Block_2 | 1.1108 | 13.7394 | TRUE | 10.9224 | 8.2132 | NA | NA | 1.2887 | NA | NA | 1.5173
P28356 | Homeobox protein Hox-D9 | HOXD9 | 1:1 | Block_2 | 0.52 | 13.1833 | TRUE | 24.9396 | 23.3641 | NA | NA | 1.8475 | NA | NA | 2.0661
P35367 | Histamine H1 receptor | HRH1 | 1:1 | Block_2 | 0.55026 | 10.9709 | TRUE | NA | NA | NA | NA | 1.2406 | NA | NA | 0.93038
O75031 | Heat shock factor 2-binding protein | HSF2BP | 1:1 | Block_2 | 1.0942 | 12.4739 | TRUE | NA | 7.6969 | NA | NA | NA | NA | NA | 1.2127
Q96JZ2 | Hematopoietic SH2 domain-containing protein | HSH2D | 1:1 | Block_2 | 1.5 | 12.4001 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.5721
Q58FF6 | Putative heat shock protein HSP 90-beta 4 | HSP90AB4P | 1:1 | Block_2 | 2.8179 | 11.1523 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96EW2 | HSPB1-associated protein 1 | HSPBAP1 | 1:1 | Block_2 | 1.1372 | 10.0204 | TRUE | NA | 14.3635 | NA | NA | 1.4224 | NA | NA | 2.0312
P28223 | 5-hydroxytryptamine receptor 2A | HTR2A | 1:1 | Block_2 | 1.0297 | 6.6554 | TRUE | 17.3301 | 23.3557 | NA | NA | 1.1779 | NA | NA | 1.104
P57058 | Hormonally up-regulated neu tumor-associated kinase | HUNK | 1:1 | Block_2 | 0.68833 | 11.5776 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.2611
Q4G0P3 | Hydrocephalus-inducing protein homolog | HYDIN | 1:1 | Block_2 | 0.3917 | 12.3051 | TRUE | NA | NA | NA | NA | 0.46869 | NA | NA | 2.0526
Q8NDH6 | Islet cell autoantigen 1-like protein | ICA1L | 1:1 | Block_2 | 0.60314 | 9.6099 | TRUE | 4.1293 | 13.8876 | NA | NA | 1.1984 | NA | 0.62094 | 1.5457
Q9BYX4 | Interferon-induced helicase C domain-containing protein 1 | IFIH1 | 1:1 | Block_2 | 4.0375 | 16.3312 | FALSE | NA | NA | NA | NA | NA | NA | NA | 5.2983
P01562 | Interferon alpha-1/13 | IFNA1_IFNA13 | 1:1 | Block_2 | 1.3684 | 13.5444 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.6103
Q12894 | Interferon-related developmental regulator 2 | IFRD2 | 1:1 | Block_2 | -1.5881 | 10.8745 | TRUE | 14.0604 | NA | NA | NA | -0.1611 | NA | NA | -0.13695
Q9UG01 | Intraflagellar transport protein 172 homolog | IFT172 | 1:1 | Block_2 | 2.1384 | 11.9794 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y547 | Intraflagellar transport protein 25 homolog | IFT25 | 1:1 | Block_2 | -0.85318 | 11.8019 | TRUE | 16.9035 | NA | NA | NA | -0.64189 | NA | -0.49213 | 1.0119
Q96FT9 | Intraflagellar transport protein 43 homolog | IFT43 | 1:1 | Block_2 | 2.6461 | 13.5158 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96LB3 | Intraflagellar transport protein 74 homolog | IFT74 | 1:1 | Block_2 | 2.1474 | 10.218 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WYA0 | Intraflagellar transport protein 81 homolog | IFT81 | 1:1 | Block_2 | 2 | 13.9276 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.4252
Q13099 | Intraflagellar transport protein 88 homolog | IFT88 | 1:1 | Block_2 | 1.9796 | 12.7274 | TRUE | 9.3623 | 12.8233 | NA | NA | 2.2451 | NA | NA | 2.8526
Q86VG3 | Intraflagellar transport-associated protein | IFTAP | 1:1 | Block_2 | 1.8154 | 10.8128 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6U949 | Putative insulin-like growth factor 2 antisense gene protein | IGF2-AS | 1:1 | Block_2 | 1.037 | 12.5589 | TRUE | NA | NA | NA | NA | 1.9832 | NA | NA | 2.2816
Q8N9C0 | Immunoglobulin superfamily member 22 | IGSF22 | 1:1 | Block_2 | -1.645 | 11.7302 | TRUE | 14.4265 | 23.549 | NA | -1.0134 | 0.69837 | NA | -0.7195 | 0.67346
Q8IYA8 | Interactor of HORMAD1 protein 1 | IHO1 | 1:1 | Block_2 | 2.2194 | 12.651 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q01113 | Interleukin-9 receptor | IL9R | 1:1 | Block_2 | 1.0819 | 10.0076 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.2041
Q12905 | Interleukin enhancer-binding factor 2 | ILF2 | 1:1 | Block_2 | 2.5 | 8.5621 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H160 | Inhibitor of growth protein 2 | ING2 | 1:1 | Block_2 | 2 | 9.2975 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96EL1 | PAK4-inhibitor INKA1 | INKA1 | 1:1 | Block_2 | 3.1326 | 11.8428 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9ULG1 | Chromatin-remodeling ATPase INO80 | INO80 | 1:1 | Block_2 | 0.86895 | 15.3663 | TRUE | NA | NA | NA | NA | 1.8871 | NA | NA | 2.3011
Q96T92 | Insulinoma-associated protein 2 | INSM2 | 1:1 | Block_2 | 2 | 16.8838 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZSG2 | Inhibitory synaptic factor 2A | INSYN2A | 1:1 | Block_2 | 0.78294 | 12.5649 | TRUE | NA | NA | NA | NA | 1.3029 | NA | NA | 0.8804
Q9NVM9 | Integrator complex subunit 13 | INTS13 | 1:1 | Block_2 | 2.5875 | 13.165 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q75QN2 | Integrator complex subunit 8 | INTS8 | 1:1 | Block_2 | 1.5 | 8.7178 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q92551 | Inositol hexakisphosphate kinase 1 | IP6K1 | 1:1 | Block_2 | 2.7181 | 13.9757 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.7852
Q8NFU5 | Inositol polyphosphate multikinase | IPMK | 1:1 | Block_2 | 0.85981 | 10.865 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96DY2 | Dynein regulatory complex protein 10 | IQCD | 1:1 | Block_2 | 1.9304 | 10.156 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6IPM2 | IQ domain-containing protein E | IQCE | 1:1 | Block_2 | 2.3766 | 14.694 | TRUE | NA | NA | NA | NA | NA | NA | NA | 4.3644
Q9H095 | Dynein regulatory complex protein 9 | IQCG | 1:1 | Block_2 | 0.88355 | 13.9731 | TRUE | 5.0886 | 7.6319 | NA | NA | 2.1819 | NA | NA | 1.6912
B3KU38 | IQCJ-SCHIP1 readthrough transcript protein | IQCJ-SCHIP1 | 1:1 | Block_2 | -0.94401 | 11.5282 | TRUE | NA | NA | NA | NA | -0.74305 | NA | NA | -0.45663
Q9H0B3 | IQ domain-containing protein N | IQCN | 1:1 | Block_2 | -1.1081 | 9.9512 | TRUE | NA | NA | NA | NA | 0.45213 | NA | NA | 0.39649
Q6DN90 | IQ motif and SEC7 domain-containing protein 1 | IQSEC1 | 1:1 | Block_2 | 1.1998 | 8.5029 | FALSE | 18.8881 | 26.598 | NA | NA | NA | NA | NA | 2.4438
Q5JU85 | IQ motif and SEC7 domain-containing protein 2 | IQSEC2 | 1:1 | Block_2 | 0.94302 | 10.6214 | FALSE | 0.87253 | NA | NA | NA | NA | NA | NA | 1.2201
Q7Z5L9 | Interferon regulatory factor 2-binding protein 2 | IRF2BP2 | 1:1 | Block_2 | -2.164 | 7.8923 | FALSE | 13.6853 | 15.8512 | NA | NA | -0.9618 | NA | -1.7842 | -0.4795
Q9H1B7 | Probable E3 ubiquitin-protein ligase IRF2BPL | IRF2BPL | 1:1 | Block_2 | 2.2446 | 11.721 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.5596
Q8WZA9 | Immunity-related GTPase family Q protein | IRGQ | 1:1 | Block_2 | 0.7331 | 10.7215 | FALSE | 10.879 | 9.467 | NA | NA | 1.1941 | NA | 1.0765 | 2.2956
Q86U28 | Iron-sulfur cluster assembly 2 homolog, mitochondrial | ISCA2 | 1:1 | Block_2 | 2.1114 | 8.279 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P61371 | Insulin gene enhancer protein ISL-1 | ISL1 | 1:1 | Block_2 | 1.2073 | 12.9555 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6GPH6 | Inositol 1,4,5-trisphosphate receptor-interacting protein-like 1 | ITPRIPL1 | 1:1 | Block_2 | 0.38126 | 9.9868 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.50027
Q8IYV9 | Izumo sperm-egg fusion protein 1 | IZUMO1 | 1:1 | Block_2 | 3.9122 | 13.7629 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q92613 | Protein Jade-3 | JADE3 | 1:1 | Block_2 | 1.4434 | 11.0075 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96AA8 | Janus kinase and microtubule-interacting protein 2 | JAKMIP2 | 1:1 | Block_2 | 1.0033 | 14.9366 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9HDC5 | Junctophilin-1 | JPH1 | 1:1 | Block_2 | 0.38996 | 11.598 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.74927
Q14678 | KN motif and ankyrin repeat domain-containing protein 1 | KANK1 | 1:1 | Block_2 | 3.3493 | 11.5952 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q7Z3B3 | KAT8 regulatory NSL complex subunit 1 | KANSL1 | 1:1 | Block_2 | 1.1317 | 11.1996 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H8E8 | Cysteine-rich protein 2-binding protein | KAT14 | 1:1 | Block_2 | 1.5975 | 11.5911 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BW62 | Katanin p60 ATPase-containing subunit A-like 1 | KATNAL1 | 1:1 | Block_2 | 2 | 13.4466 | TRUE | NA | NA | NA | NA | 4.4787 | 2.3126 | 4.477 | 6.6591
Q8WVZ9 | Kelch repeat and BTB domain-containing protein 7 | KBTBD7 | 1:1 | Block_2 | 2.9631 | 11.9382 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P22001 | Potassium voltage-gated channel subfamily A member 3 | KCNA3 | 1:1 | Block_2 | 2.3105 | 9.1176 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NZV8 | Potassium voltage-gated channel subfamily D member 2 | KCND2 | 1:1 | Block_2 | 3.5482 | 9.0275 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P15382 | Putative potassium voltage-gated channel subfamily E member 1B | KCNE1 | 1:1 | Block_2 | 1.5289 | 8.9124 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N5I3 | Potassium channel regulatory protein | KCNRG | 1:1 | Block_2 | 1.5828 | 10.3855 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q68DU8 | BTB/POZ domain-containing protein KCTD16 | KCTD16 | 1:1 | Block_2 | 0.31926 | 11.1634 | FALSE | 11.8814 | 7.3737 | NA | NA | 0.73146 | NA | NA | 1.1706
Q17RG1 | BTB/POZ domain-containing protein KCTD19 | KCTD19 | 1:1 | Block_2 | 0.57494 | 11.3897 | FALSE | 10.9302 | 23.2165 | NA | NA | 0.63836 | NA | NA | 1.2717
Q14681 | BTB/POZ domain-containing protein KCTD2 | KCTD2 | 1:1 | Block_2 | 2.7559 | 13.7864 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.9979
Q6ZWB6 | BTB/POZ domain-containing protein KCTD8 | KCTD8 | 1:1 | Block_2 | 1.5082 | 15.3608 | TRUE | 11.0103 | 21.4426 | NA | NA | NA | NA | NA | NA
Q8NAX2 | Keratinocyte differentiation factor 1 | KDF1 | 1:1 | Block_2 | 0.57085 | 13.1036 | TRUE | NA | 0.16501 | NA | NA | 0.83571 | NA | NA | 0.98403
Q9BY66 | Lysine-specific demethylase 5D | KDM5D | 1:1 | Block_2 | 0.64794 | 12.8462 | TRUE | NA | NA | NA | NA | 0.84708 | NA | NA | 1.3424
O75525 | KH domain-containing, RNA-binding, signal transduction-associated protein 3 | KHDRBS3 | 1:1 | Block_2 | 1.4416 | 7.1695 | FALSE | 9.7314 | 9.4816 | NA | NA | NA | NA | NA | 2.1385
Q6ZU52 | Uncharacterized protein KIAA0408 | KIAA0408 | 1:1 | Block_2 | 4.8599 | 14.9635 | TRUE | NA | NA | NA | NA | NA | NA | NA | 5.0254
Q9BVV6 | Protein TALPID3 | KIAA0586 | 1:1 | Block_2 | 0.9581 | 11.2387 | TRUE | 4.8906 | 1.9573 | NA | NA | 1.4941 | NA | NA | 1.4836
Q9ULL0 | Acrosomal protein KIAA1210 | KIAA1210 | 1:1 | Block_2 | -0.18387 | 13.2076 | FALSE | NA | NA | NA | NA | 0.21768 | NA | NA | 0.39227
Q5JYT7 | Uncharacterized protein KIAA1755 | KIAA1755 | 1:1 | Block_2 | 1.5501 | 13.8834 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.1103
Q9NS87 | Kinesin-like protein KIF15 | KIF15 | 1:1 | Block_2 | 3.0685 | 12.3337 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NI77 | Kinesin-like protein KIF18A | KIF18A | 1:1 | Block_2 | -2.4029 | 12.5576 | TRUE | NA | NA | NA | NA | -2.1651 | NA | NA | -1.7259
Q2TAC6 | Kinesin-like protein KIF19 | KIF19 | 1:1 | Block_2 | 0.97939 | 13.9013 | TRUE | 8.1054 | 35.1857 | NA | NA | 1.0646 | NA | NA | 1.6112
Q12756 | Kinesin-like protein KIF1A | KIF1A | 1:1 | Block_2 | 2.192 | 11.0847 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.2037
Q5T7B8 | Kinesin-like protein KIF24 | KIF24 | 1:1 | Block_2 | 2.9169 | 13.6384 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.0894
Q9ULI4 | Kinesin-like protein KIF26A | KIF26A | 1:1 | Block_2 | 1.1976 | 15.045 | TRUE | 12.8691 | 5.1702 | NA | NA | NA | NA | NA | 2.0363
Q99661 | Kinesin-like protein KIF2C | KIF2C | 1:1 | Block_2 | 2.5 | 11.8206 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O15066 | Kinesin-like protein KIF3B | KIF3B | 1:1 | Block_2 | 2 | 13.37 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.2069
Q2VIQ3 | Chromosome-associated kinesin KIF4B | KIF4B | 1:1 | Block_2 | 2 | 11.8343 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O60282 | Kinesin heavy chain isoform 5C | KIF5C | 1:1 | Block_2 | 2 | 10.5965 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q2M1P5 | Kinesin-like protein KIF7 | KIF7 | 1:1 | Block_2 | 2 | 13.8021 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9HAQ2 | Kinesin-like protein KIF9 | KIF9 | 1:1 | Block_2 | 0.66181 | 12.441 | TRUE | 13.469 | 21.9659 | NA | NA | 1.6282 | NA | NA | 2.2197
Q8N743 | Killer cell immunoglobulin-like receptor 3DL3 | KIR3DL3 | 1:1 | Block_2 | 2.5 | 8.1651 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y4X4 | Krueppel-like factor 12 | KLF12 | 1:1 | Block_2 | 0.79424 | 10.0925 | TRUE | NA | NA | NA | NA | 0.96687 | NA | NA | 0.96919
Q8TBB5 | Kelch domain-containing protein 4 | KLHDC4 | 1:1 | Block_2 | 0.023953 | 12.6015 | TRUE | 11.9525 | 11.7403 | NA | NA | 1.3764 | NA | 1.0234 | 2.9587
Q96G42 | Kelch domain-containing protein 7B | KLHDC7B | 1:1 | Block_2 | 1.2234 | 11.629 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.3026
Q2TBA0 | Kelch-like protein 40 | KLHL40 | 1:1 | Block_2 | 0.62691 | 9.2382 | FALSE | 8.6142 | 13.6715 | NA | NA | 2.6048 | NA | NA | 1.8859
Q107X0 | Putative protein KRIP1 | KLKP1 | 1:1 | Block_2 | 1.478 | 12.8146 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IZD2 | Inactive histone-lysine N-methyltransferase 2E | KMT2E | 1:1 | Block_2 | 1.3825 | 9.1818 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.6528
Q9NQR1 | N-lysine methyltransferase KMT5A | KMT5A | 1:1 | Block_2 | 0.82528 | 12.4491 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q4FZB7 | Histone-lysine N-methyltransferase KMT5B | KMT5B | 1:1 | Block_2 | 0.93424 | 11.5173 | TRUE | NA | NA | NA | NA | 1.0998 | NA | NA | 1.1838
Q13601 | KRR1 small subunit processome component homolog | KRR1 | 1:1 | Block_2 | 0.26249 | 12.1015 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.0457
Q15323 | Keratin, type I cuticular Ha1 | KRT31 | 1:1 | Block_2 | 3.9644 | 9.9212 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N1N4 | Keratin, type II cytoskeletal 78 | KRT78 | 1:1 | Block_2 | 1.0989 | 10.6943 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BQD3 | KxDL motif-containing protein 1 | KXD1 | 1:1 | Block_2 | 1.151 | 10.5765 | TRUE | NA | NA | NA | NA | 3.0713 | NA | NA | 1.2613
Q5T7N2 | LINE-1 type transposase domain-containing protein 1 | L1TD1 | 1:1 | Block_2 | 2 | 11.9796 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q969R5 | Lethal(3)malignant brain tumor-like protein 2 | L3MBTL2 | 1:1 | Block_2 | 0.51285 | 14.861 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6IAA8 | Ragulator complex protein LAMTOR1 | LAMTOR1 | 1:1 | Block_2 | -1.2524 | 11.1682 | TRUE | NA | NA | NA | NA | NA | NA | NA | -0.98697
Q659C4 | La-related protein 1B | LARP1B | 1:1 | Block_2 | 1.0077 | 11.3765 | FALSE | 15.5414 | 20.6083 | NA | 1.3071 | 2.6943 | NA | 2.4672 | 4.3194
Q71RC2 | La-related protein 4 | LARP4 | 1:1 | Block_2 | 3.6254 | 11.223 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y4W2 | Ribosomal biogenesis protein LAS1L | LAS1L | 1:1 | Block_2 | 1.738 | 15.1036 | TRUE | 10.1348 | 36.9889 | NA | NA | NA | NA | NA | NA
Q6XYB7 | Transcription factor LBX2 | LBX2 | 1:1 | Block_2 | 0.34238 | 10.6087 | FALSE | 10.9039 | NA | NA | NA | 0.37755 | NA | NA | 1.3131
O95447 | Lebercilin-like protein | LCA5L | 1:1 | Block_2 | 0.56037 | 10.9443 | TRUE | 11.5191 | 19.631 | NA | NA | 1.6609 | NA | NA | 1.8645
P06239 | Tyrosine-protein kinase Lck | LCK | 1:1 | Block_2 | 0.39141 | 10.1066 | FALSE | 16.9683 | 16.2318 | NA | 1.0227 | 1.8494 | NA | 1.0502 | 2.7404
Q96BZ8 | Leukocyte receptor cluster member 1 | LENG1 | 1:1 | Block_2 | 0.52507 | 12.117 | TRUE | NA | NA | NA | 0.90205 | 3.0537 | 0.89946 | 3.263 | 5.2444
Q96PV6 | Leukocyte receptor cluster member 8 | LENG8 | 1:1 | Block_2 | 0.012733 | 13.0532 | TRUE | NA | NA | NA | NA | 0.68214 | NA | NA | 1.1737
P48742 | LIM/homeobox protein Lhx1 | LHX1 | 1:1 | Block_2 | 0.76624 | 9.9689 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.0623
Q9UGP4 | LIM domain-containing protein 1 | LIMD1 | 1:1 | Block_2 | -0.31179 | 10.0456 | TRUE | 15.2417 | 10.9125 | NA | 1.082 | 3.261 | 0.8678 | 3.3764 | 5.2911
A2RU37 | Uncharacterized protein encoded by LINC02872 | LINC02872 | 1:1 | Block_2 | 1.5 | 10.4873 | TRUE | NA | NA | NA | NA | 3.5928 | NA | NA | 1.7393
E5RJ46 | Putative uncharacterized protein LINC02906 | LINC02906 | 1:1 | Block_2 | 1.1793 | 11.685 | TRUE | NA | NA | NA | NA | 2.9546 | NA | NA | 1.1967
Q5VXI9 | Lipase member N | LIPN | 1:1 | Block_2 | 1.3471 | 7.4579 | TRUE | NA | NA | NA | NA | 1.4654 | NA | NA | 1.3696
Q9NZU5 | LIM and cysteine-rich domains protein 1 | LMCD1 | 1:1 | Block_2 | 1.0679 | 8.8565 | TRUE | NA | 4.6428 | NA | NA | NA | NA | NA | NA
Q8IXW0 | Lamin tail domain-containing protein 2 | LMNTD2 | 1:1 | Block_2 | 1.8468 | 16.4435 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.1714
Q96KN1 | Protein LRATD2 | LRATD2 | 1:1 | Block_2 | 1.1847 | 10.9994 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.3528
Q9Y2L9 | Leucine-rich repeat and calponin homology domain-containing protein 1 | LRCH1 | 1:1 | Block_2 | 2.3703 | 14.1605 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.7872
Q9BTN0 | Leucine-rich repeat and fibronectin type-III domain-containing protein 3 | LRFN3 | 1:1 | Block_2 | 3.5948 | 10.6026 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BY71 | Leucine-rich repeat-containing protein 3 | LRRC3 | 1:1 | Block_2 | 1.011 | 6.2968 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.0722
Q1X8D7 | Leucine-rich repeat-containing protein 36 | LRRC36 | 1:1 | Block_2 | 1.5776 | 14.0265 | TRUE | 9.277 | 0.024414 | NA | NA | NA | NA | NA | 2.7442
Q9HBW1 | Leucine-rich repeat-containing protein 4 | LRRC4 | 1:1 | Block_2 | 0.21591 | 9.4059 | TRUE | 25.4948 | 11.2345 | NA | NA | 0.87386 | NA | 0.68378 | 1.487
Q9Y546 | Leucine-rich repeat-containing protein 42 | LRRC42 | 1:1 | Block_2 | -0.033694 | 12.3701 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.25972
Q8IUZ0 | Leucine-rich repeat-containing protein 49 | LRRC49 | 1:1 | Block_2 | -2.5707 | 9.9631 | TRUE | NA | NA | NA | NA | NA | NA | NA | -1.2959
Q9HCJ2 | Leucine-rich repeat-containing protein 4C | LRRC4C | 1:1 | Block_2 | 2.051 | 9.725 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N4P6 | Leucine-rich repeat-containing protein 71 | LRRC71 | 1:1 | Block_2 | 0.82386 | 13.172 | TRUE | NA | NA | NA | NA | 1.0156 | NA | NA | 1.4243
Q9C099 | Leucine-rich repeat and coiled-coil domain-containing protein 1 | LRRCC1 | 1:1 | Block_2 | 0.72092 | 9.4746 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86VH4 | Leucine-rich repeat transmembrane neuronal protein 4 | LRRTM4 | 1:1 | Block_2 | 1.2123 | 11.7015 | TRUE | NA | NA | NA | NA | 1.3615 | NA | NA | 2.2401
Q3MHD2 | Protein LSM12 | LSM12 | 1:1 | Block_2 | 2.5 | 13.2897 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P29376 | Leukocyte tyrosine kinase receptor | LTK | 1:1 | Block_2 | -0.90639 | 8.5681 | FALSE | 12.9578 | 7.9572 | -0.62251 | 0.26127 | 1.111 | 0.061887 | 0.7844 | 1.9482
Q96GA3 | Protein LTV1 homolog | LTV1 | 1:1 | Block_2 | 3.0167 | 15.0988 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.1406
Q96LR2 | Leucine rich adaptor protein 1 | LURAP1 | 1:1 | Block_2 | 0.031594 | 7.3491 | TRUE | 16.9552 | 3.9013 | NA | 0.16465 | 0.82722 | NA | 0.52703 | 1.7791
Q86SG7 | Lysozyme g-like protein 2 | LYG2 | 1:1 | Block_2 | 0.001 | 7.7553 | FALSE | NA | NA | NA | NA | 0.37245 | NA | NA | 0.45364
Q96S90 | LysM and putative peptidoglycan-binding domain-containing protein 1 | LYSMD1 | 1:1 | Block_2 | 0.99442 | 13.0771 | TRUE | NA | 1.1435 | NA | NA | 2.1154 | NA | 2.5714 | 4.0558
O60299 | Leucine zipper putative tumor suppressor 3 | LZTS3 | 1:1 | Block_2 | 1.6335 | 11.2974 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q15013 | MAD2L1-binding protein | MAD2L1BP | 1:1 | Block_2 | 0.64761 | 5.9407 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.0719
Q9H063 | Repressor of RNA polymerase III transcription MAF1 homolog | MAF1 | 1:1 | Block_2 | 1.2795 | 14.0585 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.665
Q9Y5Q3 | Transcription factor MafB | MAFB | 1:1 | Block_2 | 0.5148 | 13.0088 | TRUE | 14.1386 | 10.3994 | NA | NA | NA | NA | NA | 2.1213
A6NCF6 | Putative MAGE domain-containing protein MAGEA13P | MAGEA13P | 1:1 | Block_2 | 0.4151 | 11.5127 | TRUE | 8.806 | 36.5719 | NA | NA | 1.2898 | NA | NA | 1.281
P43358 | Melanoma-associated antigen 4 | MAGEA4 | 1:1 | Block_2 | 2.3245 | 13.5283 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P43362 | Melanoma-associated antigen 9 | MAGEA9 | 1:1 | Block_2 | 1.1165 | 9.4422 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96M61 | Melanoma-associated antigen B18 | MAGEB18 | 1:1 | Block_2 | 2.8342 | 10.8988 | TRUE | 8.32 | 15.5371 | NA | NA | NA | NA | NA | NA
Q9H6Y5 | PDZ domain-containing protein MAGIX | MAGIX | 1:1 | Block_2 | -0.72267 | 11.6288 | TRUE | NA | NA | NA | NA | -0.58216 | NA | NA | 0.064672
Q9P2G4 | Microtubule-associated protein 10 | MAP10 | 1:1 | Block_2 | 1.8022 | 12.1256 | FALSE | NA | NA | NA | NA | 1.8364 | NA | NA | NA
Q92918 | Mitogen-activated protein kinase kinase kinase kinase 1 | MAP4K1 | 1:1 | Block_2 | 3.3424 | 12.8421 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.7145
Q9H9H5 | MAP6 domain-containing protein 1 | MAP6D1 | 1:1 | Block_2 | 1.5343 | 13.4272 | TRUE | NA | NA | NA | NA | 2.1572 | NA | NA | 1.9802
Q96T17 | MAP7 domain-containing protein 2 | MAP7D2 | 1:1 | Block_2 | 1.8839 | 12.322 | TRUE | NA | NA | NA | NA | 2.0886 | NA | 2.4232 | 4.2978
Q9UPT6 | C-Jun-amino-terminal kinase-interacting protein 3 | MAPK8IP3 | 1:1 | Block_2 | 2.1546 | 9.8607 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.3918
Q9BPZ7 | Target of rapamycin complex 2 subunit MAPKAP1 | MAPKAP1 | 1:1 | Block_2 | 2.1855 | 14.4013 | TRUE | NA | NA | NA | NA | 2.6199 | NA | NA | 3.7834
Q9H992 | E3 ubiquitin-protein ligase MARCHF7 | MARCHF7 | 1:1 | Block_2 | 1.5078 | 13.3948 | TRUE | NA | NA | NA | NA | 1.9905 | NA | NA | 1.575
P27448 | MAP/microtubule affinity-regulating kinase 3 | MARK3 | 1:1 | Block_2 | 1.5322 | 12.4665 | FALSE | NA | NA | NA | NA | 2.0375 | NA | NA | 2.4485
Q96L34 | MAP/microtubule affinity-regulating kinase 4 | MARK4 | 1:1 | Block_2 | 1.0636 | 13.7377 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6P0Q8 | Microtubule-associated serine/threonine-protein kinase 2 | MAST2 | 1:1 | Block_2 | -0.92886 | 11.1447 | TRUE | NA | NA | NA | NA | NA | NA | NA | -0.13349
O15021 | Microtubule-associated serine/threonine-protein kinase 4 | MAST4 | 1:1 | Block_2 | 1.3737 | 11.8507 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.3161
Q9NS73 | MAP3K12-binding inhibitory protein 1 | MBIP | 1:1 | Block_2 | 1.8783 | 13.3242 | TRUE | NA | NA | NA | NA | 2.3013 | NA | NA | NA
Q8IVS2 | Malonyl-CoA-acyl carrier protein transacylase, mitochondrial | MCAT | 1:1 | Block_2 | 2.5299 | 14.6881 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O60318 | Germinal-center associated nuclear protein | MCM3AP | 1:1 | Block_2 | 0.74206 | 11.0069 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.0131
Q96EZ8 | Microspherule protein 1 | MCRS1 | 1:1 | Block_2 | 0.73572 | 9.1116 | FALSE | NA | 6.1017 | NA | NA | NA | NA | NA | 1.8626
Q71F56 | Mediator of RNA polymerase II transcription subunit 13-like | MED13L | 1:1 | Block_2 | 1.5 | 12.222 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.0527
O75586 | Mediator of RNA polymerase II transcription subunit 6 | MED6 | 1:1 | Block_2 | 0.46251 | 12.1198 | FALSE | 8.0863 | 5.5235 | NA | NA | 0.64331 | NA | 0.62581 | 2.5978
Q06413 | Myocyte-specific enhancer factor 2C | MEF2C | 1:1 | Block_2 | 1.6991 | 8.9425 | FALSE | NA | NA | NA | NA | NA | NA | 1.9647 | 3.7342
O15553 | Pyrin | MEFV | 1:1 | Block_2 | -0.1067 | 7.7284 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.77412
A2RUB1 | Meiosis-specific coiled-coil domain-containing protein MEIOC | MEIOC | 1:1 | Block_2 | -2.1569 | 8.5223 | TRUE | 13.3585 | 12.9652 | NA | NA | -1.096 | NA | -2.1267 | -0.82606
Q96IZ6 | tRNA N(3)-methylcytidine methyltransferase METTL2A | METTL2A | 1:1 | Block_2 | 1.9615 | 10.5813 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TCB7 | tRNA N(3)-methylcytidine methyltransferase METTL6 | METTL6 | 1:1 | Block_2 | 0.16697 | 9.3624 | FALSE | 7.8092 | 1.6531 | NA | NA | 1.7995 | NA | NA | 0.88057
Q6ZN04 | RNA-binding protein MEX3B | MEX3B | 1:1 | Block_2 | 1.6978 | 11.333 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.598
Q86XN8 | RNA-binding protein MEX3D | MEX3D | 1:1 | Block_2 | 1.9622 | 13.1147 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.3547
P55082 | Microfibril-associated glycoprotein 3 | MFAP3 | 1:1 | Block_2 | 5.6374 | 14.6079 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BY79 | Membrane frizzled-related protein | MFRP | 1:1 | Block_2 | 1.4477 | 9.0744 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IWI9 | MAX gene-associated protein | MGA | 1:1 | Block_2 | 0.65291 | 11.9894 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5JRA6 | Transport and Golgi organization protein 1 homolog | MIA3 | 1:1 | Block_2 | 1.1325 | 9.9446 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TDZ2 | [F-actin]-monooxygenase MICAL1 | MICAL1 | 1:1 | Block_2 | 0.064325 | 7.9619 | TRUE | 16.8222 | 6.5634 | NA | 0.54478 | 2.8006 | 0.16864 | 3.0777 | 4.808
Q7RTP6 | [F-actin]-monooxygenase MICAL3 | MICAL3 | 1:1 | Block_2 | 0.70488 | 11.321 | FALSE | NA | NA | NA | NA | 1.5085 | NA | 0.75109 | 2.436
Q504T8 | Midnolin | MIDN | 1:1 | Block_2 | 1.5864 | 14.7183 | TRUE | NA | NA | NA | NA | 1.9309 | NA | 2.3976 | 4.3861
Q9UGB7 | Inositol oxygenase | MIOX | 1:1 | Block_2 | 2.8475 | 13.1724 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6P0N0 | Mis18-binding protein 1 | MIS18BP1 | 1:1 | Block_2 | 1.0142 | 9.3624 | TRUE | 6.1125 | 7.199 | NA | NA | NA | NA | NA | 1.048
Q8IVT2 | Mitotic interactor and substrate of PLK1 | MISP | 1:1 | Block_2 | 0.16643 | 12.9232 | TRUE | 11.488 | 12.946 | NA | 0.30916 | 1.6461 | 0.18655 | 1.37 | 2.75
Q13434 | Putative E3 ubiquitin-protein ligase makorin-4 | MKRN4P | 1:1 | Block_2 | 1.5676 | 12.5493 | FALSE | NA | NA | NA | NA | 1.9061 | NA | NA | 2.5538
Q16655 | Melanoma antigen recognized by T-cells 1 | MLANA | 1:1 | Block_2 | 5.436 | 11.2175 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q15773 | Myeloid leukemia factor 2 | MLF2 | 1:1 | Block_2 | 3.3318 | 11.9016 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5VWP3 | Muscular LMNA-interacting protein | MLIP | 1:1 | Block_2 | 1.5765 | 14.6055 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.6098
O95396 | Adenylyltransferase and sulfurtransferase MOCS3 | MOCS3 | 1:1 | Block_2 | 1 | 11.1645 | TRUE | 5.5588 | 15.251 | NA | 1.2733 | 2.4161 | NA | 2.2787 | 4.0683
Q7Z3U7 | Protein MON2 homolog | MON2 | 1:1 | Block_2 | 1.4032 | 11.1701 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q15014 | Mortality factor 4-like protein 2 | MORF4L2 | 1:1 | Block_2 | 2.389 | 10.3358 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q99547 | M-phase phosphoprotein 6 | MPHOSPH6 | 1:1 | Block_2 | 1.8588 | 11.5625 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.887
B2RBV5 | Putative MORF4 family-associated protein 1-like protein UPP | MR1L2 | 1:1 | Block_2 | 0.12826 | 12.1204 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96EL3 | 39S ribosomal protein L53, mitochondrial | MRPL53 | 1:1 | Block_2 | 0.78815 | 6.5875 | FALSE | NA | NA | NA | NA | 1.0498 | NA | NA | 2.0295
Q92552 | 28S ribosomal protein S27, mitochondrial | MRPS27 | 1:1 | Block_2 | 2.2531 | 11.5854 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.6705
Q96H12 | Myb/SANT-like DNA-binding domain-containing protein 3 | MSANTD3 | 1:1 | Block_2 | 1.5648 | 12.5446 | TRUE | NA | NA | NA | NA | 1.5761 | NA | NA | 3.4504
Q9HCI7 | E3 ubiquitin-protein ligase MSL2 | MSL2 | 1:1 | Block_2 | 1.8188 | 12.1931 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.7708
Q04912 | Macrophage-stimulating protein receptor | MST1R | 1:1 | Block_2 | 1.2176 | 7.4935 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.4492
Q13126 | S-methyl-5'-thioadenosine phosphorylase | MTAP | 1:1 | Block_2 | 1 | 5.1698 | FALSE | NA | NA | NA | NA | 1.7036 | NA | NA | 2.3066
Q9Y4B5 | Microtubule cross-linking factor 1 | MTCL1 | 1:1 | Block_2 | 0.74593 | 12.7319 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q49AM1 | Transcription termination factor 2, mitochondrial | MTERF2 | 1:1 | Block_2 | -0.34268 | 10.0386 | FALSE | 36.1813 | 18.1893 | NA | NA | 0.0025165 | NA | NA | 0.28677
Q96E29 | Transcription termination factor 3, mitochondrial | MTERF3 | 1:1 | Block_2 | 3.0693 | 14.8344 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.1313
Q14872 | Metal regulatory transcription factor 1 | MTF1 | 1:1 | Block_2 | 0.77986 | 14.2773 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.1658
Q9H019 | Mitochondrial fission regulator 1-like | MTFR1L | 1:1 | Block_2 | 0.4478 | 8.7786 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.48981
Q9Y217 | Myotubularin-related protein 6 | MTMR6 | 1:1 | Block_2 | 1.6817 | 13.2145 | TRUE | NA | NA | NA | NA | NA | NA | 1.7296 | 3.3908
Q9Y216 | Myotubularin-related protein 7 | MTMR7 | 1:1 | Block_2 | -0.8626 | 5.6337 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P0P8 | Mitochondrial transcription rescue factor 1 | MTRES1 | 1:1 | Block_2 | 2.2093 | 14.3555 | TRUE | NA | NA | NA | NA | 4.972 | NA | 3.4911 | 6.6811
Q13203 | Myosin-binding protein H | MYBPH | 1:1 | Block_2 | 1.508 | 11.5749 | TRUE | NA | NA | NA | NA | 2.5352 | NA | NA | 3.435
P12882 | Myosin-1 | MYH1 | 1:1 | Block_2 | -1.8461 | 7.5796 | TRUE | 17.9113 | 16.1046 | NA | NA | 0.50569 | NA | -0.34024 | 2.8328
P35580 | Myosin-10 | MYH10 | 1:1 | Block_2 | 0.37519 | 13.1 | TRUE | NA | NA | NA | 0.72488 | 1.7565 | NA | 1.1539 | 3.0889
P12883 | Myosin-7 | MYH7 | 1:1 | Block_2 | -0.2388 | 9.7487 | TRUE | 23.9252 | 7.5084 | NA | NA | 0.81962 | NA | 0.17616 | 2.9782
Q02045 | Myosin light chain 5 | MYL5 | 1:1 | Block_2 | 2.5786 | 11.9814 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.3558
Q8IUG5 | Unconventional myosin-XVIIIb | MYO18B | 1:1 | Block_2 | 1.0726 | 10.5439 | FALSE | 18.8752 | 9.4889 | NA | NA | 1.6236 | NA | NA | 1.7342
Q12965 | Unconventional myosin-Ie | MYO1E | 1:1 | Block_2 | 0.51229 | 11.8995 | TRUE | 13.086 | 18.9178 | NA | 1.3959 | 4.3781 | 1.8484 | 3.779 | 6.7001
B0I1T2 | Unconventional myosin-Ig | MYO1G | 1:1 | Block_2 | 1.3379 | 7.8759 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.8045
Q8WXR4 | Myosin-IIIb | MYO3B | 1:1 | Block_2 | 1.2165 | 10.6452 | TRUE | NA | NA | NA | NA | 1.3342 | NA | NA | 1.3678
Q9Y4I1 | Unconventional myosin-Va | MYO5A | 1:1 | Block_2 | 3.03 | 10.3145 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.0625
Q9ULV0 | Unconventional myosin-Vb | MYO5B | 1:1 | Block_2 | 1.2287 | 9.3411 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5VVJ2 | Deubiquitinase MYSM1 | MYSM1 | 1:1 | Block_2 | 0.38003 | 9.1303 | FALSE | NA | 16.0442 | NA | NA | 0.49591 | NA | NA | 1.0128
Q86UW6 | NEDD4-binding protein 2 | N4BP2 | 1:1 | Block_2 | 0.95315 | 12.3339 | FALSE | 27.1892 | 7.7454 | NA | NA | 1.0145 | NA | NA | 2.3607
Q92802 | NEDD4-binding protein 2-like 2 | N4BP2L2 | 1:1 | Block_2 | 2.1485 | 9.973 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BSU3 | N-alpha-acetyltransferase 11 | NAA11 | 1:1 | Block_2 | 3.0172 | 14.7198 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H115 | Beta-soluble NSF attachment protein | NAPB | 1:1 | Block_2 | 1.0768 | 7.9663 | TRUE | NA | NA | NA | NA | 1.0979 | NA | NA | 1.8705
Q8IVL1 | Neuron navigator 2 | NAV2 | 1:1 | Block_2 | 1.1916 | 13.3638 | TRUE | NA | NA | NA | NA | 1.3956 | NA | NA | 2.2761
Q6ZNJ1 | Neurobeachin-like protein 2 | NBEAL2 | 1:1 | Block_2 | -3.5753 | 8.1358 | FALSE | NA | NA | NA | -3.4674 | -2.5292 | NA | -2.3501 | 0.25927
Q86T75 | Neuroblastoma breakpoint family member 11 | NBPF11 | 1:1 | Block_2 | 2.3472 | 10.4372 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q15003 | Condensin complex subunit 2 | NCAPH | 1:1 | Block_2 | 2.5573 | 11.2776 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A8MVU1_P14598 | Neutrophil cytosol factor 1/1C | NCF1C_NCF1 | 1:1 | Block_2 | -2.7019 | 8.7816 | FALSE | 13.3922 | 17.9047 | -2.2918 | -1.2443 | 1.7718 | -1.9203 | 0.64375 | 8.2165
Q9HCH0 | Nck-associated protein 5-like | NCKAP5L | 1:1 | Block_2 | 0.92213 | 10.9956 | TRUE | NA | NA | NA | NA | 1.4015 | NA | 1.0846 | 2.5065
Q15788 | Nuclear receptor coactivator 1 | NCOA1 | 1:1 | Block_2 | 0.8683 | 12.7473 | TRUE | NA | NA | NA | NA | 1.7621 | NA | 1.0546 | 2.9806
Q8NI08 | Nuclear receptor coactivator 7 | NCOA7 | 1:1 | Block_2 | 2.854 | 10.4754 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O14931 | Natural cytotoxicity triggering receptor 3 | NCR3 | 1:1 | Block_2 | 0.54706 | 8.6692 | TRUE | 11.6296 | 7.6226 | NA | NA | 1.2293 | NA | NA | 1.5002
Q9NXR1 | Nuclear distribution protein nudE homolog 1 | NDE1 | 1:1 | Block_2 | 1.2935 | 5.415 | TRUE | NA | NA | NA | NA | 1.3347 | NA | NA | 1.7639
Q92597 | Protein NDRG1 | NDRG1 | 1:1 | Block_2 | -0.21737 | 7.8766 | FALSE | 25.8612 | 16.0741 | -0.12148 | 0.88203 | 1.6818 | 0.51479 | 1.74 | 3.3295
O95299 | NADH dehydrogenase [ubiquinone] 1 alpha subcomplex subunit 10, mitochondrial | NDUFA10 | 1:1 | Block_2 | 1.5458 | 8.1021 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5857
Q9NRX3 | NADH dehydrogenase [ubiquinone] 1 alpha subcomplex subunit 4-like 2 | NDUFA4L2 | 1:1 | Block_2 | 1.2786 | 8.4536 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y6M9 | NADH dehydrogenase [ubiquinone] 1 beta subcomplex subunit 9 | NDUFB9 | 1:1 | Block_2 | 1.4239 | 10.8326 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4419
P51955 | Serine/threonine-protein kinase Nek2 | NEK2 | 1:1 | Block_2 | 1.9266 | 13.4379 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P51957 | Serine/threonine-protein kinase Nek4 | NEK4 | 1:1 | Block_2 | 1 | 11.6684 | TRUE | NA | 10.3683 | NA | NA | 3.6691 | NA | 1.4941 | 3.4364
Q8IXH7 | Negative elongation factor C/D | NELFCD | 1:1 | Block_2 | -1 | 10.0011 | TRUE | NA | NA | NA | NA | -0.9466 | NA | NA | -0.33436
O76050 | E3 ubiquitin-protein ligase NEURL1 | NEURL1 | 1:1 | Block_2 | 1.1475 | 8.2273 | TRUE | 9.2703 | 14.4797 | NA | NA | NA | NA | NA | NA
A8MQ27 | E3 ubiquitin-protein ligase NEURL1B | NEURL1B | 1:1 | Block_2 | 0.89612 | 9.2337 | FALSE | NA | NA | NA | NA | 0.9371 | NA | NA | 0.96924
Q8NBZ9 | Putative uncharacterized protein NEXN-AS1 | NEXN-AS1 | 1:1 | Block_2 | 2.2337 | 13.3909 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NET5 | NFAT activation molecule 1 | NFAM1 | 1:1 | Block_2 | 0.68137 | 5.1356 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NEJ9 | Neuroguidin | NGDN | 1:1 | Block_2 | 0.55678 | 11.3227 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5392
Q6T4R5 | Actin remodeling regulator NHS | NHS | 1:1 | Block_2 | 1.3054 | 11.5186 | TRUE | 19.6179 | 11.8344 | NA | NA | 1.7014 | NA | NA | 2.1053
Q5SYE7 | NHS-like protein 1 | NHSL1 | 1:1 | Block_2 | 2 | 7.3773 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O15226 | NF-kappa-B-repressing factor | NKRF | 1:1 | Block_2 | 1.5091 | 3.4632 | FALSE | NA | NA | NA | NA | 1.5535 | NA | NA | NA
Q86W24 | NACHT, LRR and PYD domains-containing protein 14 | NLRP14 | 1:1 | Block_2 | 1.4854 | 11.0494 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NWW6 | Nicotinamide riboside kinase 1 | NMRK1 | 1:1 | Block_2 | 1.7109 | 7.4677 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NPI5 | Nicotinamide riboside kinase 2 | NMRK2 | 1:1 | Block_2 | 3.2212 | 9.3135 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.8123
Q9Y314 | Nitric oxide synthase-interacting protein | NOSIP | 1:1 | Block_2 | 0.001 | 11.8654 | TRUE | 22.2914 | 14.9022 | NA | NA | 1.1919 | NA | 1.2178 | 3.6383
P51513 | RNA-binding protein Nova-1 | NOVA1 | 1:1 | Block_2 | 0.85327 | 11.7685 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.1168
Q9UNW9 | RNA-binding protein Nova-2 | NOVA2 | 1:1 | Block_2 | 0.72993 | 9.0533 | TRUE | 12.5615 | 9.0724 | NA | NA | NA | NA | NA | 1.2805
E9PIF3 | Nuclear pore complex-interacting protein family member A2 | NPIPA2 | 1:1 | Block_2 | 2.5349 | 9.5735 | TRUE | NA | NA | NA | NA | 2.7227 | NA | NA | 3.2618
E9PJ23 | Nuclear pore complex-interacting protein family member B6 | NPIPB6 | 1:1 | Block_2 | -1.3261 | 9.4849 | TRUE | NA | 4.1351 | NA | -0.7634 | 0.21606 | NA | -1.1673 | 1.3834
Q14995 | Nuclear receptor subfamily 1 group D member 2 | NR1D2 | 1:1 | Block_2 | 0.1491 | 12.1469 | TRUE | 6.5855 | 0.74204 | NA | NA | NA | NA | NA | NA
P49116 | Nuclear receptor subfamily 2 group C member 2 | NR2C2 | 1:1 | Block_2 | 2.9896 | 12.7955 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q86VF7 | Nebulin-related-anchoring protein | NRAP | 1:1 | Block_2 | 2.6812 | 13.3782 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O14511 | Pro-neuregulin-2, membrane-bound isoform | NRG2 | 1:1 | Block_2 | 1.4569 | 9.4824 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P56975 | Pro-neuregulin-3, membrane-bound isoform | NRG3 | 1:1 | Block_2 | 1.689 | 9.7682 | TRUE | NA | NA | NA | NA | 2.3266 | NA | NA | 1.834
P42857 | Neuronal vesicle trafficking-associated protein 1 | NSG1 | 1:1 | Block_2 | 2.1184 | 14.6274 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P78549 | Endonuclease III-like protein 1 | NTHL1 | 1:1 | Block_2 | 1.0608 | 13.4198 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.1831
Q9BSD7 | Cancer-related nucleoside-triphosphatase | NTPCR | 1:1 | Block_2 | 2.1395 | 12.9841 | FALSE | NA | NA | NA | NA | 2.5211 | NA | 3.1199 | 5.0152
P04629 | High affinity nerve growth factor receptor | NTRK1 | 1:1 | Block_2 | 0.1114 | 6.2173 | TRUE | 12.1738 | 16.04 | NA | 0.23949 | 0.88026 | NA | 0.56029 | 1.2967
O95848 | Uridine diphosphate glucose pyrophosphatase NUDT14 | NUDT14 | 1:1 | Block_2 | 2.7665 | 7.8155 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P0C024 | Peroxisomal coenzyme A diphosphatase NUDT7 | NUDT7 | 1:1 | Block_2 | 1 | 12.2209 | TRUE | 17.2013 | 14.775 | NA | 1.1438 | 2.3679 | NA | 1.646 | 3.5264
Q7Z417 | FMR1-interacting protein NUFIP2 | NUFIP2 | 1:1 | Block_2 | 1.5 | 11.7264 | TRUE | 21.8454 | NA | NA | NA | 1.6563 | NA | NA | 3.1836
O60356 | Nuclear protein 1 | NUPR1 | 1:1 | Block_2 | 0.78924 | 7.3199 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96CM4 | Nucleoredoxin-like protein 1 | NXNL1 | 1:1 | Block_2 | -0.59136 | 12.8271 | TRUE | NA | NA | NA | NA | NA | NA | NA | -0.39384
Q5VST9 | Obscurin | OBSCN | 1:1 | Block_2 | 3.1556 | 12.4946 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96NG3 | Outer dynein arm-docking complex subunit 4 | ODAD4 | 1:1 | Block_2 | 3.5278 | 14.1009 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O15527 | N-glycosylase/DNA lyase | OGG1 | 1:1 | Block_2 | -0.35659 | 9.8334 | TRUE | NA | NA | NA | NA | -0.19708 | NA | NA | 0.57734
O95897 | Noelin-2 | OLFM2 | 1:1 | Block_2 | 1.6553 | 11.4708 | TRUE | 4.3321 | 9.385 | NA | NA | NA | NA | NA | 2.0082
Q13438 | Protein OS-9 | OS9 | 1:1 | Block_2 | -2.9838 | 6.5862 | TRUE | NA | 4.7411 | NA | NA | -2.1686 | NA | -2.7874 | -1.3779
Q9NX31 | Oxidative stress-responsive serine-rich protein 1 | OSER1 | 1:1 | Block_2 | 2.4917 | 12.0086 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96FW1 | Ubiquitin thioesterase OTUB1 | OTUB1 | 1:1 | Block_2 | 2.7198 | 8.7172 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q01804 | OTU domain-containing protein 4 | OTUD4 | 1:1 | Block_2 | 1.9143 | 13.0166 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.0862
Q12889 | Oviduct-specific glycoprotein | OVGP1 | 1:1 | Block_2 | -0.70045 | 11.4759 | TRUE | NA | NA | NA | NA | 0.21044 | NA | NA | 0.73488
Q92791 | Endoplasmic reticulum protein SC65 | P3H4 | 1:1 | Block_2 | 5 | 14.123 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P0CB38 | Polyadenylate-binding protein 4-like | PABPC4L | 1:1 | Block_2 | 0.23939 | 7.7037 | FALSE | NA | NA | NA | NA | 0.31394 | NA | NA | 0.90486
Q8N7B6 | PACRG-like protein | PACRGL | 1:1 | Block_2 | 0.45482 | 10.3313 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.2138
Q9UKS6 | Protein kinase C and casein kinase substrate in neurons protein 3 | PACSIN3 | 1:1 | Block_2 | 1 | 12.702 | TRUE | 0.34242 | 13.554 | NA | 1.0231 | 2.3772 | NA | 2.2967 | 3.7527
Q5JUK9 | P antigen family member 3 | PAGE3 | 1:1 | Block_2 | 2.5083 | 15.2603 | FALSE | NA | NA | NA | NA | 3.0248 | NA | NA | NA
Q8TEW8 | Partitioning defective 3 homolog B | PARD3B | 1:1 | Block_2 | -0.45619 | 12.2035 | FALSE | NA | 1.2833 | NA | NA | 0.087944 | NA | NA | 1.6984
Q9BYG5 | Partitioning defective 6 homolog beta | PARD6B | 1:1 | Block_2 | 1.6042 | 11.1372 | TRUE | 16.7533 | 32.5583 | NA | NA | NA | NA | NA | 1.7884
O95453 | Poly(A)-specific ribonuclease PARN | PARN | 1:1 | Block_2 | 1.7529 | 13.5866 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q53GL7 | Protein mono-ADP-ribosyltransferase PARP10 | PARP10 | 1:1 | Block_2 | 0.001 | 13.7663 | TRUE | 39.3678 | 39.3172 | NA | 1.7894 | 3.1015 | 1.2381 | 3.2233 | 4.7918
Q8N3A8 | Protein mono-ADP-ribosyltransferase PARP8 | PARP8 | 1:1 | Block_2 | 2 | 13.4636 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NWS1 | PCNA-interacting partner | PARPBP | 1:1 | Block_2 | 2.9801 | 12.6666 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.6628
Q8NI35 | InaD-like protein | PATJ | 1:1 | Block_2 | 2.5 | 13.2476 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.5962
Q96KB5 | Lymphokine-activated killer T-cell-originated protein kinase | PBK | 1:1 | Block_2 | 2 | 14.1628 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P40424 | Pre-B-cell leukemia transcription factor 1 | PBX1 | 1:1 | Block_2 | 0.65973 | 11.4368 | TRUE | NA | NA | NA | NA | 0.79222 | NA | NA | NA
Q9P2E7 | Protocadherin-10 | PCDH10 | 1:1 | Block_2 | 1 | 7.4644 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7147
Q9Y5I3 | Protocadherin alpha-1 | PCDHA1 | 1:1 | Block_2 | 2 | 10.3541 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P12004 | Proliferating cell nuclear antigen | PCNA | 1:1 | Block_2 | 3 | 8.7278 | FALSE | NA | NA | NA | NA | 3.1364 | NA | NA | NA
P49585 | Choline-phosphate cytidylyltransferase A | PCYT1A | 1:1 | Block_2 | 1.9608 | 6.3895 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P20941 | Phosducin | PDC | 1:1 | Block_2 | 2.0415 | 12.7061 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q16342 | Programmed cell death protein 2 | PDCD2 | 1:1 | Block_2 | 1.4094 | 12.2931 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7552
Q13371 | Phosducin-like protein | PDCL | 1:1 | Block_2 | 1 | 12.56 | TRUE | NA | NA | NA | NA | 1.7076 | NA | 1.4007 | 2.6798
P54750 | Dual specificity calcium/calmodulin-dependent 3',5'-cyclic nucleotide phosphodiesterase 1A | PDE1A | 1:1 | Block_2 | 0.17953 | 6.0797 | FALSE | NA | 19.5295 | NA | NA | 0.29747 | NA | NA | 0.61263
O00408 | cGMP-dependent 3',5'-cyclic phosphodiesterase | PDE2A | 1:1 | Block_2 | 1.3881 | 11.0165 | TRUE | NA | NA | NA | NA | 2.469 | NA | 2.9269 | 4.9214
Q5VU43 | Myomegalin | PDE4DIP | 1:1 | Block_2 | 1.4111 | 10.828 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P35913 | Rod cGMP-specific 3',5'-cyclic phosphodiesterase subunit beta | PDE6B | 1:1 | Block_2 | 2.4295 | 10.0895 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6P474 | Putative pyridoxal-dependent decarboxylase domain-containing protein 2 | PDXDC2P | 1:1 | Block_2 | 1.0649 | 10.9953 | TRUE | NA | 5.1915 | NA | NA | 1.2204 | NA | NA | 2.2051
Q76G19 | PDZ domain-containing protein 4 | PDZD4 | 1:1 | Block_2 | 1.1352 | 8.7525 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZMN7 | PDZ domain-containing RING finger protein 4 | PDZRN4 | 1:1 | Block_2 | 0.37446 | 12.0433 | TRUE | 17.8457 | 24.0462 | NA | NA | NA | NA | NA | NA
Q9H792 | Inactive tyrosine-protein kinase PEAK1 | PEAK1 | 1:1 | Block_2 | 1.4059 | 10.2977 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.552
O95394 | Phosphoacetylglucosamine mutase | PGM3 | 1:1 | Block_2 | 2.3481 | 5.7685 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9C0D0 | Phosphatase and actin regulator 1 | PHACTR1 | 1:1 | Block_2 | 2.2718 | 10.604 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.905
Q8IXK0 | Polyhomeotic-like protein 2 | PHC2 | 1:1 | Block_2 | 1.7517 | 11.1614 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96QT6 | PHD finger protein 12 | PHF12 | 1:1 | Block_2 | 2.328 | 12.3673 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A8MW92 | PHD finger protein 20-like protein 1 | PHF20L1 | 1:1 | Block_2 | 0.84061 | 12.8769 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.3359
Q96EK2 | PHD finger protein 21B | PHF21B | 1:1 | Block_2 | 1.5693 | 11.1328 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UPV7 | PHD finger protein 24 | PHF24 | 1:1 | Block_2 | 2.604 | 13.0844 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y5J5 | Pleckstrin homology-like domain family A member 3 | PHLDA3 | 1:1 | Block_2 | -4 | 5.0095 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O14813 | Paired mesoderm homeobox protein 2A | PHOX2A | 1:1 | Block_2 | 1.1254 | 8.5976 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q7Z3Z4 | Piwi-like protein 4 | PIWIL4 | 1:1 | Block_2 | 1 | 13.2154 | TRUE | NA | 10.8382 | NA | NA | NA | NA | NA | 1.487
Q8NG27 | E3 ubiquitin-protein ligase Praja-1 | PJA1 | 1:1 | Block_2 | 1.644 | 12.0582 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P51178 | 1-phosphatidylinositol 4,5-bisphosphate phosphodiesterase delta-1 | PLCD1 | 1:1 | Block_2 | 1.9745 | 10.9893 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BRC7 | 1-phosphatidylinositol 4,5-bisphosphate phosphodiesterase delta-4 | PLCD4 | 1:1 | Block_2 | 0.29464 | 12.5937 | TRUE | NA | NA | NA | NA | 0.46274 | NA | NA | 1.5717
P19174 | 1-phosphatidylinositol 4,5-bisphosphate phosphodiesterase gamma-1 | PLCG1 | 1:1 | Block_2 | 0.001 | 7.5212 | FALSE | 12.7747 | 17.7563 | NA | 0.32551 | 1.3711 | 0.040061 | 1.2703 | 2.6075
Q4KWH8 | 1-phosphatidylinositol 4,5-bisphosphate phosphodiesterase eta-1 | PLCH1 | 1:1 | Block_2 | 1 | 14.1014 | TRUE | 21.5004 | 28.103 | NA | NA | 2.1638 | NA | NA | 3.4006
Q15111 | Inactive phospholipase C-like protein 1 | PLCL1 | 1:1 | Block_2 | -0.46909 | 9.3623 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H4M7 | Pleckstrin homology domain-containing family A member 4 | PLEKHA4 | 1:1 | Block_2 | -0.38525 | 10.6622 | TRUE | NA | NA | NA | NA | NA | NA | NA | -0.0076593
A6NEE1 | Pleckstrin homology domain-containing family D member 1 | PLEKHD1 | 1:1 | Block_2 | -2.9576 | 7.6752 | TRUE | 14.1548 | NA | NA | NA | NA | NA | NA | -2.2724
Q96S99 | Pleckstrin homology domain-containing family F member 1 | PLEKHF1 | 1:1 | Block_2 | 2.3398 | 11.4054 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H7P9 | Pleckstrin homology domain-containing family G member 2 | PLEKHG2 | 1:1 | Block_2 | 1 | 8.9997 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4947
O94827 | Pleckstrin homology domain-containing family G member 5 | PLEKHG5 | 1:1 | Block_2 | 0.87596 | 10.7221 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9ULM0 | Pleckstrin homology domain-containing family H member 1 | PLEKHH1 | 1:1 | Block_2 | 1.2636 | 14.0013 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.3049
Q8IVE3 | Pleckstrin homology domain-containing family H member 2 | PLEKHH2 | 1:1 | Block_2 | 1.6152 | 14.12 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.9994
Q9NYY3 | Serine/threonine-protein kinase PLK2 | PLK2 | 1:1 | Block_2 | 3 | 13.6824 | TRUE | NA | NA | NA | NA | 4.1753 | NA | NA | 5.6315
O60568 | Multifunctional procollagen lysine hydroxylase and glycosyltransferase LH3 | PLOD3 | 1:1 | Block_2 | 0.001 | 7.4275 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O43660 | Pleiotropic regulator 1 | PLRG1 | 1:1 | Block_2 | 1.76 | 10.1176 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.9398
P13797 | Plastin-3 | PLS3 | 1:1 | Block_2 | 1.8133 | 10.5391 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O75051 | Plexin-A2 | PLXNA2 | 1:1 | Block_2 | 1.195 | 5.9829 | TRUE | NA | NA | NA | NA | 3.5658 | NA | NA | 1.5138
Q86V59 | Paraneoplastic antigen-like protein 8A | PNMA8A | 1:1 | Block_2 | -0.73577 | 11.5974 | TRUE | 18.6622 | 7.6673 | NA | -0.62051 | 0.64851 | NA | -0.35671 | 1.2371
Q8TC44 | POC1 centriolar protein homolog B | POC1B | 1:1 | Block_2 | 1.0445 | 10.9399 | TRUE | NA | NA | NA | NA | 1.1533 | NA | NA | 2.9952
Q8NBL1 | Protein O-glucosyltransferase 1 | POGLUT1 | 1:1 | Block_2 | 1.6527 | 10.2669 | FALSE | NA | NA | NA | NA | 1.7339 | NA | NA | NA
P06746 | DNA polymerase beta | POLB | 1:1 | Block_2 | 1.3742 | 9.9215 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.8839
Q15054 | DNA polymerase delta subunit 3 | POLD3 | 1:1 | Block_2 | 1 | 14.6224 | TRUE | NA | NA | NA | NA | 1.2068 | NA | NA | 2.6433
Q9HCU8 | DNA polymerase delta subunit 4 | POLD4 | 1:1 | Block_2 | 2.161 | 8.8511 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BY77 | Polymerase delta-interacting protein 3 | POLDIP3 | 1:1 | Block_2 | 2.2351 | 11.8587 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P05423 | DNA-directed RNA polymerase III subunit RPC4 | POLR3D | 1:1 | Block_2 | 1.3852 | 8.9877 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.6623
Q6S8J7 | POTE ankyrin domain family member A | POTEA | 1:1 | Block_2 | 1.7266 | 15.276 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.6989
B2RU33 | POTE ankyrin domain family member C | POTEC | 1:1 | Block_2 | 0.45783 | 12.5112 | TRUE | NA | NA | NA | NA | 0.47008 | NA | NA | 0.9249
Q9BYX7 | Putative beta-actin-like protein 3 | POTEKP | 1:1 | Block_2 | 1.1552 | 10.1231 | TRUE | NA | NA | NA | NA | 1.2978 | NA | NA | 1.3027
A6NI47 | Putative POTE ankyrin domain family member M | POTEM | 1:1 | Block_2 | 0.001 | 11.252 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UKI9 | POU domain, class 2, transcription factor 3 | POU2F3 | 1:1 | Block_2 | 2.0641 | 13.2104 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H2U2 | Inorganic pyrophosphatase 2, mitochondrial | PPA2 | 1:1 | Block_2 | 1 | 6.098 | FALSE | NA | NA | NA | NA | 2.3951 | NA | 2.0167 | 3.339
Q13136 | Liprin-alpha-1 | PPFIA1 | 1:1 | Block_2 | 0.001 | 9.3691 | TRUE | 12.1807 | 12.1756 | NA | 0.32987 | 1.2499 | 0.54288 | 1.6559 | 3.034
O75335 | Liprin-alpha-4 | PPFIA4 | 1:1 | Block_2 | 2.5604 | 14.7818 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.1184
Q86W92 | Liprin-beta-1 | PPFIBP1 | 1:1 | Block_2 | 1 | 13.6673 | TRUE | 36.9628 | 6.5926 | NA | 1.0765 | 2.9702 | NA | 2.403 | 5.0926
Q8ND30 | Liprin-beta-2 | PPFIBP2 | 1:1 | Block_2 | 0.78393 | 11.1462 | TRUE | 8.2071 | 8.9144 | NA | NA | 1.0547 | NA | NA | 1.8512
Q96KQ4 | Apoptosis-stimulating of p53 protein 1 | PPP1R13B | 1:1 | Block_2 | 0.0078667 | 10.022 | FALSE | 12.2283 | 6.9239 | NA | 0.11954 | 0.82709 | NA | 0.57942 | 1.5415
Q13522 | Protein phosphatase 1 regulatory subunit 1A | PPP1R1A | 1:1 | Block_2 | 0.47916 | 12.0102 | TRUE | 26.5073 | 17.8151 | NA | NA | 1.3131 | NA | 1.2562 | 3.2616
Q8WVI7 | Protein phosphatase 1 regulatory subunit 1C | PPP1R1C | 1:1 | Block_2 | 2.3062 | 12.6536 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96PQ5 | Putative protein phosphatase inhibitor 2-like protein 1 | PPP1R2P1 | 1:1 | Block_2 | 0.44769 | 11.1563 | TRUE | 9.8709 | 10.0928 | NA | 0.57728 | 1.713 | 0.4569 | 1.7319 | 2.8299
Q8TAP8 | Protein phosphatase 1 regulatory subunit 35 | PPP1R35 | 1:1 | Block_2 | 0.5 | 15.1896 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5386
Q86XI6 | Protein phosphatase 1 regulatory subunit 3B | PPP1R3B | 1:1 | Block_2 | 1.5865 | 14.2198 | FALSE | NA | NA | NA | NA | 2.521 | NA | NA | 2.8847
Q15435 | Protein phosphatase 1 regulatory subunit 7 | PPP1R7 | 1:1 | Block_2 | 4.1481 | 13.9618 | TRUE | NA | NA | NA | NA | 4.6203 | NA | NA | NA
Q6NUP7 | Serine/threonine-protein phosphatase 4 regulatory subunit 4 | PPP4R4 | 1:1 | Block_2 | 1.1614 | 15.2489 | TRUE | NA | NA | NA | NA | 2.1644 | NA | NA | 1.6495
O75170 | Serine/threonine-protein phosphatase 6 regulatory subunit 2 | PPP6R2 | 1:1 | Block_2 | 0.26743 | 8.2266 | TRUE | 8.7671 | 5.2242 | NA | NA | 0.36148 | NA | NA | 1.5433
Q9NQV5 | PR domain-containing protein 11 | PRDM11 | 1:1 | Block_2 | 2.8203 | 13.535 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q13029 | PR domain zinc finger protein 2 | PRDM2 | 1:1 | Block_2 | -0.43179 | 10.861 | TRUE | NA | NA | NA | NA | 0.33176 | NA | NA | 0.96456
P49642 | DNA primase small subunit | PRIM1 | 1:1 | Block_2 | 2.1251 | 10.5852 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P41743 | Protein kinase C iota type | PRKCI | 1:1 | Block_2 | 0.50118 | 7.2045 | TRUE | NA | NA | NA | NA | 0.55561 | NA | NA | 1.0891
O60260 | E3 ubiquitin-protein ligase parkin | PRKN | 1:1 | Block_2 | 1.3582 | 14.6189 | FALSE | NA | NA | NA | NA | 1.4499 | NA | NA | 2.0434
O60678 | Protein arginine N-methyltransferase 3 | PRMT3 | 1:1 | Block_2 | 1.0683 | 9.1073 | FALSE | NA | NA | NA | NA | 1.5191 | NA | 1.1208 | 2.3793
Q5THK1 | Protein PRR14L | PRR14L | 1:1 | Block_2 | 1.1679 | 12.5746 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BU68 | Proline-rich protein 15-like protein | PRR15L | 1:1 | Block_2 | 1.0067 | 12.9925 | TRUE | 7.405 | 1.1026 | NA | NA | 1.8292 | NA | NA | 2.1837
B1ATL7 | Proline-rich protein 32 | PRR32 | 1:1 | Block_2 | 0.78305 | 14.1268 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P0CG20 | Proline-rich protein 35 | PRR35 | 1:1 | Block_2 | 1.564 | 15.2675 | TRUE | NA | 3.4491 | NA | NA | NA | NA | NA | NA
O14668 | Transmembrane gamma-carboxyglutamic acid protein 1 | PRRG1 | 1:1 | Block_2 | 0.93245 | 11.1285 | TRUE | NA | 11.8174 | NA | NA | 1.5263 | NA | NA | 2.0767
Q06124 | Tyrosine-protein phosphatase non-receptor type 11 | PTPN11 | 1:1 | Block_2 | -0.14585 | 6.8601 | FALSE | 21.2791 | 18.0098 | -0.13278 | 1.9432 | 3.746 | 1.4853 | 4.2504 | 6.0798
Q15678 | Tyrosine-protein phosphatase non-receptor type 14 | PTPN14 | 1:1 | Block_2 | 0.62709 | 12.5166 | FALSE | NA | 17.3675 | NA | NA | 1.3703 | NA | NA | 2.6704
P17706 | Tyrosine-protein phosphatase non-receptor type 2 | PTPN2 | 1:1 | Block_2 | 2.6237 | 11.3903 | FALSE | NA | NA | NA | NA | 3.0272 | NA | NA | 3.2777
Q4JDL3 | Tyrosine-protein phosphatase non-receptor type 20 | PTPN20 | 1:1 | Block_2 | -0.81965 | 12.83 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P18433 | Receptor-type tyrosine-protein phosphatase alpha | PTPRA | 1:1 | Block_2 | 1.5597 | 11.5635 | TRUE | NA | NA | NA | NA | 2.4036 | NA | NA | 1.7449
Q5TGL8 | PX domain-containing protein 1 | PXDC1 | 1:1 | Block_2 | 5 | 16.9054 | TRUE | NA | NA | NA | NA | NA | NA | NA | 5.085
Q96C36 | Pyrroline-5-carboxylate reductase 2 | PYCR2 | 1:1 | Block_2 | 1.994 | 9.6964 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.3374
Q9H0J4 | Glutamine-rich protein 2 | QRICH2 | 1:1 | Block_2 | 0.58287 | 11.5623 | TRUE | NA | NA | NA | NA | 0.7351 | NA | NA | 0.92651
Q2KHR3 | Glutamine and serine-rich protein 1 | QSER1 | 1:1 | Block_2 | 2 | 12.3652 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZRP7 | Sulfhydryl oxidase 2 | QSOX2 | 1:1 | Block_2 | 1.3852 | 9.4857 | FALSE | NA | NA | NA | NA | 1.5851 | NA | NA | 1.7873
Q9Y3T6 | R3H and coiled-coil domain-containing protein 1 | R3HCC1 | 1:1 | Block_2 | 0.57356 | 11.7239 | FALSE | NA | 11.2905 | NA | 0.7679 | 1.6724 | 0.59657 | 1.8067 | 3.8377
Q7Z5L2 | Coiled-coil domain-containing protein R3HCC1L | R3HCC1L | 1:1 | Block_2 | 0.61973 | 10.6231 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.7497
P62491 | Ras-related protein Rab-11A | RAB11A | 1:1 | Block_2 | 1 | 5.1307 | FALSE | 2.0992 | 5.5067 | NA | NA | 1.4931 | NA | NA | 1.6785
Q9BXF6 | Rab11 family-interacting protein 5 | RAB11FIP5 | 1:1 | Block_2 | -0.19188 | 11.5509 | TRUE | 4.5643 | 14.9336 | NA | NA | 0.68502 | NA | 0.02865 | 1.344
Q14964 | Ras-related protein Rab-39A | RAB39A | 1:1 | Block_2 | 1.4694 | 10.9665 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96E17 | Ras-related protein Rab-3C | RAB3C | 1:1 | Block_2 | 1.2897 | 12.8936 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UNT1 | Rab-like protein 2B | RABL2B | 1:1 | Block_2 | 1.4374 | 11.8266 | FALSE | NA | NA | NA | 1.5758 | 3.4078 | 1.7338 | 3.8655 | 5.3964
Q96B01 | RAD51-associated protein 1 | RAD51AP1 | 1:1 | Block_2 | 1.7722 | 12.1054 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q09MP3 | RAD51-associated protein 2 | RAD51AP2 | 1:1 | Block_2 | 2.4115 | 10.0524 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.8413
Q96JH8 | Ras-associating and dilute domain-containing protein | RADIL | 1:1 | Block_2 | -0.14443 | 12.3144 | TRUE | NA | NA | NA | NA | 0.61736 | NA | 0.25273 | 1.4875
Q86X27 | Ras-specific guanine nucleotide-releasing factor RalGPS2 | RALGPS2 | 1:1 | Block_2 | -0.91056 | 9.0027 | TRUE | NA | NA | NA | NA | NA | NA | NA | -0.19127
Q9HD47 | Ran guanine nucleotide release factor | RANGRF | 1:1 | Block_2 | 0.73831 | 10.2475 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TEU7 | Rap guanine nucleotide exchange factor 6 | RAPGEF6 | 1:1 | Block_2 | 2.1299 | 12.6391 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6NK89 | Ras association domain-containing protein 10 | RASSF10 | 1:1 | Block_2 | 1.1155 | 7.2523 | FALSE | NA | NA | NA | NA | 1.1195 | NA | NA | 1.4507
Q8WWW0 | Ras association domain-containing protein 5 | RASSF5 | 1:1 | Block_2 | 1.2303 | 9.968 | FALSE | NA | NA | NA | NA | NA | NA | NA | 5.0072
Q02833 | Ras association domain-containing protein 7 | RASSF7 | 1:1 | Block_2 | 0.90801 | 12.1841 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y2V3 | Retinal homeobox protein Rx | RAX | 1:1 | Block_2 | 2.4683 | 11.1009 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q15291 | Retinoblastoma-binding protein 5 | RBBP5 | 1:1 | Block_2 | 1.4825 | 11.6964 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4881
Q5T481 | RNA-binding protein 20 | RBM20 | 1:1 | Block_2 | 2.8645 | 11.6167 | TRUE | 21.9489 | 26.4933 | NA | NA | NA | NA | NA | NA
Q9NW64 | Pre-mRNA-splicing factor RBM22 | RBM22 | 1:1 | Block_2 | 2.9558 | 12.0859 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P2N5 | RNA-binding protein 27 | RBM27 | 1:1 | Block_2 | 1.4027 | 11.2349 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8972
Q96EV2 | RNA-binding protein 33 | RBM33 | 1:1 | Block_2 | 1.9457 | 11.3897 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6XE24 | RNA-binding motif, single-stranded-interacting protein 3 | RBMS3 | 1:1 | Block_2 | 1.2529 | 7.0078 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6JBY9 | CapZ-interacting protein | RCSD1 | 1:1 | Block_2 | 2.3376 | 7.907 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IYK8 | GTP-binding protein REM 2 | REM2 | 1:1 | Block_2 | 1.4767 | 9.4356 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NFH8 | RalBP1-associated Eps domain-containing protein 2 | REPS2 | 1:1 | Block_2 | 1.7912 | 10.7797 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.1845
P35251 | Replication factor C subunit 1 | RFC1 | 1:1 | Block_2 | 1.813 | 7.6825 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q14699 | Raftlin | RFTN1 | 1:1 | Block_2 | 1.5154 | 13.9679 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P48382 | DNA-binding protein RFX5 | RFX5 | 1:1 | Block_2 | 2.2328 | 10.608 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P49796 | Regulator of G-protein signaling 3 | RGS3 | 1:1 | Block_2 | 1.1103 | 12.3536 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.9466
P49798 | Regulator of G-protein signaling 4 | RGS4 | 1:1 | Block_2 | 2.5254 | 11.2889 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P49758 | Regulator of G-protein signaling 6 | RGS6 | 1:1 | Block_2 | 1.792 | 13.7363 | TRUE | NA | NA | NA | 2.2137 | 5.638 | 2.0756 | 5.9267 | 8.0116
Q6ZWK4 | Regulator of hemoglobinization and erythroid cell expansion protein | RHEX | 1:1 | Block_2 | 2.2019 | 13.3786 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.2239
O94844 | Rho-related BTB domain-containing protein 1 | RHOBTB1 | 1:1 | Block_2 | 2 | 11.3462 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N443 | RIB43A-like with coiled-coils protein 1 | RIBC1 | 1:1 | Block_2 | -0.02858 | 11.6116 | TRUE | NA | NA | NA | NA | 0.10531 | NA | NA | 0.44216
Q6NUQ1 | RAD50-interacting protein 1 | RINT1 | 1:1 | Block_2 | 0.50186 | 13.437 | TRUE | NA | NA | NA | 0.87896 | 2.6377 | 0.80458 | 2.3706 | 4.4328
Q9BVS4 | Serine/threonine-protein kinase RIO2 | RIOK2 | 1:1 | Block_2 | 1.3505 | 11.4609 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4732
Q9Y4F9 | Rho family-interacting cell polarization regulator 2 | RIPOR2 | 1:1 | Block_2 | 2.0129 | 10.1419 | TRUE | NA | NA | NA | NA | NA | NA | NA | 4.0714
Q96K30 | RBPJ-interacting and tubulin-associated protein 1 | RITA1 | 1:1 | Block_2 | 0.001 | 10.3549 | TRUE | 24.2685 | 22.0574 | NA | 0.10392 | 1.6403 | NA | 0.47543 | 2.4699
P12271 | Retinaldehyde-binding protein 1 | RLBP1 | 1:1 | Block_2 | 2 | 7.9422 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WXF3 | Relaxin-3 | RLN3 | 1:1 | Block_2 | 0.29971 | 11.9169 | TRUE | NA | NA | NA | NA | 0.3862 | NA | NA | 0.86383
P60153 | Inactive ribonuclease-like protein 9 | RNASE9 | 1:1 | Block_2 | 1.5137 | 12.4778 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P61587 | Rho-related GTP-binding protein RhoE | RND3 | 1:1 | Block_2 | 3.4895 | 8.4111 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N5U6 | RING finger protein 10 | RNF10 | 1:1 | Block_2 | 1 | 13.3415 | TRUE | NA | NA | NA | NA | 2.2069 | NA | 1.6562 | 2.9071
Q8WVD5 | RING finger protein 141 | RNF141 | 1:1 | Block_2 | 2.4859 | 11.4028 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q63HN8 | E3 ubiquitin-protein ligase RNF213 | RNF213 | 1:1 | Block_2 | 1.0491 | 11.9048 | TRUE | NA | NA | NA | NA | 1.3343 | NA | NA | 2.3254
Q8ND24 | RING finger protein 214 | RNF214 | 1:1 | Block_2 | 0.11243 | 11.0106 | TRUE | NA | NA | NA | NA | 0.70593 | NA | 0.55545 | 2.4858
A6NIN4 | RING finger protein 227 | RNF227 | 1:1 | Block_2 | 1 | 11.2723 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96LT9 | RNA-binding region-containing protein 3 | RNPC3 | 1:1 | Block_2 | 1 | 6.3433 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96MS0 | Roundabout homolog 3 | ROBO3 | 1:1 | Block_2 | 1.2054 | 12.0792 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P56715 | Oxygen-regulated protein 1 | RP1 | 1:1 | Block_2 | 2.1494 | 12.4988 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.344
Q2QD12 | Ribulose-phosphate 3-epimerase-like protein 1 | RPEL1 | 1:1 | Block_2 | -0.41344 | 10.3227 | TRUE | NA | NA | NA | -0.21504 | 1.4886 | NA | 1.3324 | 3.139
Q9H9Y2 | Ribosome production factor 1 | RPF1 | 1:1 | Block_2 | 1 | 9.2551 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q68CZ1 | Protein fantom | RPGRIP1L | 1:1 | Block_2 | 0.91206 | 14.0617 | TRUE | NA | NA | NA | NA | 0.98486 | NA | NA | NA
Q9UNE2 | Rab effector Noc2 | RPH3AL | 1:1 | Block_2 | 0.23403 | 8.3627 | TRUE | NA | 2.6393 | NA | NA | 0.56632 | NA | NA | 1.5997
P78346 | Ribonuclease P protein subunit p30 | RPP30 | 1:1 | Block_2 | 1.5248 | 11.0642 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.5503
Q96P16 | Regulation of nuclear pre-mRNA domain-containing protein 1A | RPRD1A | 1:1 | Block_2 | 0.001 | 10.7518 | TRUE | NA | 4.8796 | NA | NA | 0.47196 | NA | NA | 1.0522
P25398 | 40S ribosomal protein S12 | RPS12 | 1:1 | Block_2 | -2.3912 | 8.3218 | FALSE | NA | NA | NA | NA | -1.9904 | NA | NA | -1.2838
Q9NYV6 | RNA polymerase I-specific transcription initiation factor RRN3 | RRN3 | 1:1 | Block_2 | 1.0364 | 10.2337 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WXG1 | Radical S-adenosyl methionine domain-containing protein 2 | RSAD2 | 1:1 | Block_2 | 1.1917 | 11.5652 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.326
Q5TD94 | Radial spoke head protein 4 homolog A | RSPH4A | 1:1 | Block_2 | -0.22724 | 10.8997 | TRUE | NA | NA | NA | NA | NA | NA | NA | -0.092665
Q96IZ7 | Serine/Arginine-related protein 53 | RSRC1 | 1:1 | Block_2 | 0.059758 | 14.8069 | TRUE | 4.3488 | 8.6509 | NA | NA | 0.64906 | NA | 0.46235 | 3.2302
Q9BUV0 | Arginine/serine-rich protein 1 | RSRP1 | 1:1 | Block_2 | 1 | 13.8594 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5671
Q9BY42 | Replication termination factor 2 | RTF2 | 1:1 | Block_2 | 1.7259 | 14.3855 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.7518
Q86UN3 | Reticulon-4 receptor-like 2 | RTN4RL2 | 1:1 | Block_2 | 0.27211 | 8.0435 | FALSE | NA | 12.5413 | NA | NA | 0.77885 | NA | NA | 1.6743
Q96DX8 | Receptor-transporting protein 4 | RTP4 | 1:1 | Block_2 | 1.2979 | 12.5971 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.5401
Q9H714 | Protein associated with UVRAG as autophagy enhancer | RUBCNL | 1:1 | Block_2 | 1.3283 | 9.9509 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q59EK9 | RUN domain-containing protein 3A | RUNDC3A | 1:1 | Block_2 | 1.4935 | 9.6337 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BVN2 | AP-4 complex accessory subunit RUSC1 | RUSC1 | 1:1 | Block_2 | 2.1979 | 12.8278 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.2263
Q9Y3V2 | RWD domain-containing protein 3 | RWDD3 | 1:1 | Block_2 | -0.62622 | 7.3115 | FALSE | 15.2408 | 7.6929 | NA | NA | -0.06531 | NA | NA | 1.3463
P34925 | Tyrosine-protein kinase RYK | RYK | 1:1 | Block_2 | -1.7861 | 7.6412 | TRUE | 11.033 | 14.9026 | -1.3774 | -0.72446 | -0.22987 | -0.88049 | -0.22402 | 0.62271
P33763 | Protein S100-A5 | S100A5 | 1:1 | Block_2 | -0.1322 | 11.5553 | TRUE | NA | NA | NA | NA | 0.39104 | NA | NA | 0.68871
Q8WXG8 | Protein S100-Z | S100Z | 1:1 | Block_2 | 0.88899 | 13.4244 | TRUE | 10.1792 | 11.4676 | NA | NA | 0.97969 | NA | NA | 4.0369
Q9UJQ4 | Sal-like protein 4 | SALL4 | 1:1 | Block_2 | 1.3553 | 9.4954 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.3633
Q8N8I0 | Sterile alpha motif domain-containing protein 12 | SAMD12 | 1:1 | Block_2 | 0.53007 | 10.095 | FALSE | NA | NA | NA | NA | 0.65905 | NA | NA | 0.73779
Q5VXD3 | Sterile alpha motif domain-containing protein 13 | SAMD13 | 1:1 | Block_2 | 2.0103 | 15.0217 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96LT4 | Sphingomyelin synthase-related protein 1 | SAMD8 | 1:1 | Block_2 | -1.4163 | 10.6219 | TRUE | 6.3704 | 6.843 | NA | NA | -0.62558 | NA | NA | 0.31536
Q6NSI8 | SANT and BTB domain regulator of class switch recombination | SANBR | 1:1 | Block_2 | -0.42871 | 10.071 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y3A5 | Ribosome maturation protein SBDS | SBDS | 1:1 | Block_2 | 1 | 11.7139 | TRUE | NA | NA | NA | 1.8042 | 3.7186 | 1.2587 | 4.1837 | 6.7583
Q86WG5 | Myotubularin-related protein 13 | SBF2 | 1:1 | Block_2 | 1.4541 | 13.2638 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.6254
Q8IX30 | Signal peptide, CUB and EGF-like domain-containing protein 3 | SCUBE3 | 1:1 | Block_2 | 2.9613 | 10.0526 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P31040 | Succinate dehydrogenase [ubiquinone] flavoprotein subunit, mitochondrial | SDHA | 1:1 | Block_2 | 1.4693 | 7.9355 | TRUE | NA | NA | NA | NA | 2.1325 | NA | NA | 2.7711
Q9NRP4 | Succinate dehydrogenase assembly factor 3, mitochondrial | SDHAF3 | 1:1 | Block_2 | 1 | 10.7 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.3778
Q5VUM1 | Succinate dehydrogenase assembly factor 4, mitochondrial | SDHAF4 | 1:1 | Block_2 | 0.92482 | 9.0117 | TRUE | NA | NA | NA | NA | 1.7425 | NA | 1.3895 | 3.0266
Q8N3Y7 | Epidermal retinol dehydrogenase 2 | SDR16C5 | 1:1 | Block_2 | 1 | 12.6171 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.3233
Q96GA7 | Serine dehydratase-like | SDSL | 1:1 | Block_2 | 1.5 | 11.1217 | TRUE | 0.85358 | NA | NA | NA | 2.3852 | NA | NA | 1.7149
P55735 | Protein SEC13 homolog | SEC13 | 1:1 | Block_2 | -0.26811 | 9.6992 | TRUE | 18.397 | 14.0018 | -0.23924 | 1.2124 | 2.8299 | 1.1466 | 2.9484 | 3.999
O76054 | SEC14-like protein 2 | SEC14L2 | 1:1 | Block_2 | 0.44787 | 10.9488 | TRUE | NA | NA | NA | NA | 1.1508 | NA | NA | 1.4663
Q9H3T3 | Semaphorin-6B | SEMA6B | 1:1 | Block_2 | 2 | 10.9962 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.6473
Q9HC62 | Sentrin-specific protease 2 | SENP2 | 1:1 | Block_2 | 0.98306 | 10.5121 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.389
Q96LD8 | Sentrin-specific protease 8 | SENP8 | 1:1 | Block_2 | 1 | 8.8537 | FALSE | 4.1425 | 1.737 | NA | NA | NA | NA | NA | 1.7223
Q9P0V9 | Septin-10 | SEPTIN10 | 1:1 | Block_2 | 0.47815 | 11.4721 | TRUE | NA | 1.5433 | NA | 0.53755 | 1.3726 | NA | 1.1714 | 2.3234
Q6ZU15 | Septin-14 | SEPTIN14 | 1:1 | Block_2 | 1 | 11.6929 | TRUE | 7.5915 | 11.7683 | NA | NA | 1.4467 | NA | NA | 1.3581
O75920 | Small EDRK-rich factor 1 | SERF1A_SERF1B | 1:1 | Block_2 | 0.39081 | 10.6796 | FALSE | NA | 11.5919 | NA | NA | 0.64717 | NA | NA | 1.5687
Q9UHV2 | SERTA domain-containing protein 1 | SERTAD1 | 1:1 | Block_2 | 1.005 | 12.5443 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.3682
Q9NUC0 | SERTA domain-containing protein 4 | SERTAD4 | 1:1 | Block_2 | 0.001 | 11.9734 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86VW0 | SEC14 domain and spectrin repeat-containing protein 1 | SESTD1 | 1:1 | Block_2 | 2.53 | 8.2387 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.8739
Q9Y6X0 | SET-binding protein | SETBP1 | 1:1 | Block_2 | 0.43316 | 9.5789 | TRUE | 21.4707 | 3.5471 | NA | NA | 0.56864 | NA | NA | 1.3189
Q9C0A6 | Histone-lysine N-methyltransferase SETD5 | SETD5 | 1:1 | Block_2 | 2.3096 | 13.3239 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q13326 | Gamma-sarcoglycan | SGCG | 1:1 | Block_2 | 1.5 | 7.7788 | FALSE | 1.0018 | 11.57 | NA | NA | NA | NA | NA | 1.9692
Q5SQS7 | SH2 domain-containing protein 4B | SH2D4B | 1:1 | Block_2 | 1.4232 | 15.3349 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8925
A6NKC9 | SH2 domain-containing protein 7 | SH2D7 | 1:1 | Block_2 | 2.0395 | 13.0543 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A4FU49 | SH3 domain-containing protein 21 | SH3D21 | 1:1 | Block_2 | 1.9869 | 11.7692 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q99961 | Endophilin-A2 | SH3GL1 | 1:1 | Block_2 | 3.8414 | 9.7409 | TRUE | NA | NA | NA | NA | NA | NA | NA | 4.4812
Q9Y566 | SH3 and multiple ankyrin repeat domains protein 1 | SHANK1 | 1:1 | Block_2 | 0.30892 | 10.5257 | TRUE | 7.4804 | NA | NA | NA | 0.4505 | NA | NA | 0.65499
Q5VXU9 | Protein shortage in chiasmata 1 ortholog | SHOC1 | 1:1 | Block_2 | 1.3956 | 12.8449 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O15266 | Short stature homeobox protein | SHOX | 1:1 | Block_2 | -0.42242 | 11.8221 | FALSE | 8.6049 | 2.0851 | NA | NA | 1.8836 | NA | NA | 1.7009
O60902 | Short stature homeobox protein 2 | SHOX2 | 1:1 | Block_2 | 1.4975 | 11.4226 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.5233
Q2M3G4 | Protein Shroom1 | SHROOM1 | 1:1 | Block_2 | 2.1554 | 12.7373 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.1862
Q13796 | Protein Shroom2 | SHROOM2 | 1:1 | Block_2 | 0.53149 | 11.1393 | TRUE | 15.2321 | 12.2167 | NA | NA | 0.73855 | NA | NA | 0.63535
Q8TF72 | Protein Shroom3 | SHROOM3 | 1:1 | Block_2 | 1.1098 | 7.5366 | TRUE | NA | NA | NA | NA | 1.179 | NA | NA | 1.3477
Q9ULL8 | Protein Shroom4 | SHROOM4 | 1:1 | Block_2 | 0.65757 | 12.4943 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.416
Q9BRV8 | Suppressor of IKBKE 1 | SIKE1 | 1:1 | Block_2 | 1.5685 | 12.8577 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.8641
P81133 | Single-minded homolog 1 | SIM1 | 1:1 | Block_2 | 1.177 | 11.7776 | TRUE | NA | NA | NA | NA | 2.0805 | NA | NA | 1.4583
Q8NDZ2 | SUMO-interacting motif-containing protein 1 | SIMC1 | 1:1 | Block_2 | 1.1887 | 9.1453 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O75182 | Paired amphipathic helix protein Sin3b | SIN3B | 1:1 | Block_2 | 0.98295 | 11.1574 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P1W8 | Signal-regulatory protein gamma | SIRPG | 1:1 | Block_2 | 1.0676 | 10.4482 | TRUE | NA | NA | NA | NA | 1.6704 | NA | NA | 2.218
Q96BD8 | Spindle and kinetochore-associated protein 1 | SKA1 | 1:1 | Block_2 | 1.1189 | 12.2871 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P12757 | Ski-like protein | SKIL | 1:1 | Block_2 | -0.18347 | 12.8244 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.45489
P84550 | SKI family transcriptional corepressor 1 | SKOR1 | 1:1 | Block_2 | 1.7696 | 12.0337 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BXP2 | Solute carrier family 12 member 9 | SLC12A9 | 1:1 | Block_2 | 0.88777 | 10.1107 | TRUE | NA | NA | NA | NA | 0.98888 | NA | NA | 1.7114
Q13183 | Solute carrier family 13 member 2 | SLC13A2 | 1:1 | Block_2 | 1.8794 | 13.2743 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.2182
P41440 | Reduced folate transporter | SLC19A1 | 1:1 | Block_2 | 1.1685 | 12.1407 | TRUE | NA | NA | NA | NA | 1.3229 | NA | NA | 1.8979
Q92911 | Sodium/iodide cotransporter | SLC5A5 | 1:1 | Block_2 | -0.50145 | 11.514 | FALSE | NA | NA | NA | NA | NA | NA | NA | 0.40614
O94991 | SLIT and NTRK-like protein 5 | SLITRK5 | 1:1 | Block_2 | 0.30459 | 9.1397 | TRUE | 1.5891 | 9.627 | NA | NA | 1.22 | NA | 0.84475 | 1.7253
Q5VYV7 | Protein SLX4IP | SLX4IP | 1:1 | Block_2 | 0.53952 | 11.6013 | TRUE | NA | NA | NA | NA | 0.82196 | NA | NA | 1.5899
P28370 | Probable global transcription activator SNF2L1 | SMARCA1 | 1:1 | Block_2 | 3.1406 | 11.3725 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P51532 | Transcription activator BRG1 | SMARCA4 | 1:1 | Block_2 | 0.53599 | 9.5438 | FALSE | NA | NA | NA | 0.73924 | 2.0403 | 0.78055 | 2.1946 | 3.7965
Q9NTJ3 | Structural maintenance of chromosomes protein 4 | SMC4 | 1:1 | Block_2 | 0.76422 | 10.8142 | TRUE | 16.1663 | 11.199 | NA | NA | 1.0232 | NA | NA | 1.0499
A6NFE2 | Single-pass membrane and coiled-coil domain-containing protein 2 | SMCO2 | 1:1 | Block_2 | 1.608 | 15.9879 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.828
A2RU48 | Single-pass membrane and coiled-coil domain-containing protein 3 | SMCO3 | 1:1 | Block_2 | -0.35299 | 7.1484 | TRUE | NA | NA | NA | NA | 0.64583 | NA | NA | NA
Q9NWM0 | Spermine oxidase | SMOX | 1:1 | Block_2 | 0.5 | 11.797 | TRUE | 17.3858 | 14.4743 | NA | 1.343 | 2.6224 | NA | 2.111 | 3.8719
Q9HCE7 | E3 ubiquitin-protein ligase SMURF1 | SMURF1 | 1:1 | Block_2 | -0.06328 | 12.396 | TRUE | NA | 9.5932 | NA | NA | 0.089951 | NA | -0.06204 | 0.88244
Q9H7B4 | Histone-lysine N-methyltransferase SMYD3 | SMYD3 | 1:1 | Block_2 | 1.1441 | 2.224 | FALSE | NA | NA | NA | NA | 2.5099 | NA | 2.0051 | 4.4111
O60641 | Clathrin coat assembly protein AP180 | SNAP91 | 1:1 | Block_2 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y6H5 | Synphilin-1 | SNCAIP | 1:1 | Block_2 | 0.23413 | 12.2547 | TRUE | 5.6385 | NA | NA | NA | 1.458 | NA | NA | 1.7711
Q9NRH2 | SNF-related serine/threonine-protein kinase | SNRK | 1:1 | Block_2 | 1.9897 | 12.7683 | TRUE | NA | NA | NA | NA | 2.3923 | NA | 2.33 | 4.3662
P62316 | Small nuclear ribonucleoprotein Sm D2 | SNRPD2 | 1:1 | Block_2 | 0.72307 | 13.7319 | TRUE | NA | NA | NA | NA | 0.76314 | NA | NA | 2.532
Q92543 | Sorting nexin-19 | SNX19 | 1:1 | Block_2 | 2 | 11.6982 | TRUE | NA | NA | NA | NA | 2.0961 | NA | NA | 2.7304
Q9NX45 | Spermatogenesis- and oogenesis-specific basic helix-loop-helix-containing protein 2 | SOHLH2 | 1:1 | Block_2 | 2.1712 | 12.597 | FALSE | 16.3363 | 16.4709 | NA | NA | NA | NA | NA | NA
Q8WY21 | VPS10 domain-containing receptor SorCS1 | SORCS1 | 1:1 | Block_2 | -0.8873 | 8.8196 | TRUE | 7.0101 | 10.0833 | 0.1215 | 0.70774 | 1.325 | 0.32268 | 1.1574 | 2.1509
Q9UPU3 | VPS10 domain-containing receptor SorCS3 | SORCS3 | 1:1 | Block_2 | 1.147 | 10.717 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6NEL2 | Ankyrin repeat domain-containing protein SOWAHB | SOWAHB | 1:1 | Block_2 | 2.1678 | 10.5303 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UN79 | Transcription factor SOX-13 | SOX13 | 1:1 | Block_2 | 1.4626 | 14.4329 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.2092
O94993 | Transcription factor SOX-30 | SOX30 | 1:1 | Block_2 | 1.0809 | 9.7273 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q06945 | Transcription factor SOX-4 | SOX4 | 1:1 | Block_2 | -0.19012 | 10.7864 | TRUE | NA | NA | NA | NA | -0.031936 | NA | NA | 0.65873
P57073 | Transcription factor SOX-8 | SOX8 | 1:1 | Block_2 | -0.16686 | 9.4294 | FALSE | 20.7003 | 14.1618 | NA | NA | NA | NA | NA | 0.1807
Q13342 | Nuclear body protein SP140 | SP140 | 1:1 | Block_2 | 1.1074 | 10.3696 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q02447 | Transcription factor Sp3 | SP3 | 1:1 | Block_2 | 1 | 10.1113 | TRUE | NA | NA | NA | NA | 1.1213 | NA | NA | 1.5519
Q8TDD2 | Transcription factor Sp7 | SP7 | 1:1 | Block_2 | 1.1928 | 9.52 | FALSE | 15.464 | 13.4329 | NA | NA | 1.83 | NA | NA | 1.3148
Q8N0X2 | Sperm-associated antigen 16 protein | SPAG16 | 1:1 | Block_2 | 1.6462 | 11.7394 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96R06 | Sperm-associated antigen 5 | SPAG5 | 1:1 | Block_2 | 1 | 15.6315 | TRUE | 18.0915 | 16.8555 | NA | NA | 1.3962 | NA | NA | 2.3272
O75391 | Sperm-associated antigen 7 | SPAG7 | 1:1 | Block_2 | 0.96239 | 7.8225 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8529
Q99932 | Sperm-associated antigen 8 | SPAG8 | 1:1 | Block_2 | 0.88571 | 11.3171 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.1967
Q9NS25 | Sperm protein associated with the nucleus on the X chromosome B1 | SPANXB1 | 1:1 | Block_2 | 1.1875 | 12.4918 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.062
Q5VX52 | Spermatogenesis-associated protein 1 | SPATA1 | 1:1 | Block_2 | 1 | 11.4113 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q7Z572 | Spermatogenesis-associated protein 21 | SPATA21 | 1:1 | Block_2 | -1.4472 | 9.7868 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NHX4 | Spermatogenesis-associated protein 3 | SPATA3 | 1:1 | Block_2 | 1.8426 | 13.2858 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.0836
Q8N4H0 | Spermatogenesis associated 6-like protein | SPATA6L | 1:1 | Block_2 | -1.405 | 12.8387 | TRUE | NA | NA | NA | NA | -1.24 | NA | NA | -0.63489
Q9P0W8 | Spermatogenesis-associated protein 7 | SPATA7 | 1:1 | Block_2 | 2.1959 | 14.3218 | TRUE | 13.3775 | 14.3047 | NA | NA | 2.2404 | NA | NA | 2.4335
Q76KD6 | Speriolin | SPATC1 | 1:1 | Block_2 | 0.99913 | 13.8101 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.0402
Q496A3 | Spermatogenesis-associated serine-rich protein 1 | SPATS1 | 1:1 | Block_2 | 4.4632 | 12.3931 | TRUE | 29.2762 | 8.5852 | NA | NA | NA | NA | NA | 5.3104
Q96EA4 | Protein Spindly | SPDL1 | 1:1 | Block_2 | 1.1376 | 13.8598 | TRUE | 10.8058 | 14.9984 | NA | NA | 1.8393 | NA | NA | 1.8321
Q9Y4P9 | Sperm flagellar protein 1 | SPEF1 | 1:1 | Block_2 | 3.7402 | 14.0859 | TRUE | 9.0225 | 48.8989 | NA | NA | NA | NA | NA | NA
Q96T58 | Msx2-interacting protein | SPEN | 1:1 | Block_2 | 0.001 | 7.4631 | FALSE | 10.0956 | 15.3737 | NA | NA | 0.43629 | NA | NA | 1.2109
P16150 | Leukosialin | SPN | 1:1 | Block_2 | 0.25743 | 10.0051 | TRUE | 12.2966 | 4.4592 | NA | 0.52048 | 1.1393 | 0.491 | 1.2178 | 2.5444
Q5T280 | Putative methyltransferase C9orf114 | SPOUT1 | 1:1 | Block_2 | 0.81775 | 7.3281 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H040 | DNA-dependent metalloprotease SPRTN | SPRTN | 1:1 | Block_2 | 2.04 | 11.36 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H254 | Spectrin beta chain, non-erythrocytic 4 | SPTBN4 | 1:1 | Block_2 | 1 | 10.7119 | TRUE | NA | NA | NA | NA | 2.351 | NA | NA | 2.5259
Q9NRC6 | Spectrin beta chain, non-erythrocytic 5 | SPTBN5 | 1:1 | Block_2 | 2.3843 | 12.3041 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N5C6 | S1 RNA-binding domain-containing protein 1 | SRBD1 | 1:1 | Block_2 | 1 | 11.5949 | FALSE | NA | NA | NA | NA | 1.5781 | NA | NA | NA
Q8NEF9 | Serum response factor-binding protein 1 | SRFBP1 | 1:1 | Block_2 | 1 | 12.562 | TRUE | 17.3754 | 12.501 | NA | NA | 1.3638 | NA | NA | 1.8201
O43295 | SLIT-ROBO Rho GTPase-activating protein 3 | SRGAP3 | 1:1 | Block_2 | 1 | 12.2467 | TRUE | 20.4488 | 20.0901 | NA | NA | 1.4651 | NA | 1.6233 | 4.0093
P08240 | Signal recognition particle receptor subunit alpha | SRPRA | 1:1 | Block_2 | 1 | 13.6537 | TRUE | NA | NA | NA | NA | 2.0432 | NA | 1.7665 | 4.1236
Q9Y5M8 | Signal recognition particle receptor subunit beta | SRPRB | 1:1 | Block_2 | 0.71561 | 5.7932 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BXP5 | Serrate RNA effector molecule homolog | SRRT | 1:1 | Block_2 | 1.7827 | 9.9502 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.5488
Q05066 | Sex-determining region Y protein | SRY | 1:1 | Block_2 | -1.9501 | 11.5076 | FALSE | 4.9739 | 6.3005 | NA | NA | -1.7365 | NA | NA | 0.39287
Q76I76 | Protein phosphatase Slingshot homolog 2 | SSH2 | 1:1 | Block_2 | 1.0324 | 9.3111 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7746
P43307 | Translocon-associated protein subunit alpha | SSR1 | 1:1 | Block_2 | 1.5384 | 9.5704 | FALSE | NA | 6.5539 | NA | NA | 1.7571 | NA | NA | 2.8339
P32745 | Somatostatin receptor type 3 | SSTR3 | 1:1 | Block_2 | 1.1125 | 9.9335 | FALSE | 11.992 | 8.4674 | NA | NA | 1.2042 | NA | NA | NA
Q9Y2M2 | Protein SSUH2 homolog | SSUH2 | 1:1 | Block_2 | -0.27537 | 10.1375 | FALSE | NA | NA | NA | NA | -0.093325 | NA | NA | 1.1797
O60225 | Protein SSX5 | SSX5 | 1:1 | Block_2 | -0.89877 | 12.9885 | TRUE | NA | NA | NA | NA | NA | NA | NA | -0.33689
Q8NFI4 | Putative protein FAM10A5 | ST13P5 | 1:1 | Block_2 | 0.5 | 11.7522 | TRUE | NA | NA | NA | NA | 0.57917 | NA | NA | 1.6127
Q8WVM7 | Cohesin subunit SA-1 | STAG1 | 1:1 | Block_2 | 1.4996 | 10.4303 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.6784
Q9UJ98 | Cohesin subunit SA-3 | STAG3 | 1:1 | Block_2 | 0.5 | 11.5776 | TRUE | NA | 37.8229 | NA | NA | 1.5201 | NA | 0.78701 | 2.4695
O75886 | Signal transducing adapter molecule 2 | STAM2 | 1:1 | Block_2 | 1 | 11.0331 | TRUE | 19.494 | 10.3895 | NA | NA | 2.2981 | NA | 2.171 | 4.008
P49675 | Steroidogenic acute regulatory protein, mitochondrial | STAR | 1:1 | Block_2 | 2.0283 | 13.226 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.0711
Q9Y365 | START domain-containing protein 10 | STARD10 | 1:1 | Block_2 | 0.91583 | 8.3983 | TRUE | NA | NA | NA | NA | 1.4557 | NA | NA | 2.1055
Q9Y3M8 | StAR-related lipid transfer protein 13 | STARD13 | 1:1 | Block_2 | 1.5433 | 10.911 | TRUE | NA | NA | NA | NA | 1.636 | NA | NA | 1.9358
Q14849 | StAR-related lipid transfer protein 3 | STARD3 | 1:1 | Block_2 | 1.6494 | 10.9666 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q92502 | StAR-related lipid transfer protein 8 | STARD8 | 1:1 | Block_2 | 0.94805 | 7.8806 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.031
Q8N1F8 | Serine/threonine-protein kinase 11-interacting protein | STK11IP | 1:1 | Block_2 | 1.0753 | 11.1854 | TRUE | NA | NA | NA | NA | 1.1134 | NA | NA | 2.3048
Q9UEE5 | Serine/threonine-protein kinase 17A | STK17A | 1:1 | Block_2 | 1.561 | 8.1682 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P289 | Serine/threonine-protein kinase 26 | STK26 | 1:1 | Block_2 | 1 | 11.6822 | TRUE | 8.6597 | 8.2395 | NA | NA | 1.7959 | NA | NA | 1.6937
Q86UX6 | Serine/threonine-protein kinase 32C | STK32C | 1:1 | Block_2 | 1.094 | 12.9005 | TRUE | 0.72211 | 21.8033 | NA | NA | 2.1734 | NA | 1.1337 | 2.4684
Q9BYT3 | Serine/threonine-protein kinase 33 | STK33 | 1:1 | Block_2 | 0.63102 | 11.736 | TRUE | 0.86162 | 5.8933 | NA | NA | 1.4093 | NA | NA | 1.8386
Q9NRP7 | Serine/threonine-protein kinase 36 | STK36 | 1:1 | Block_2 | 0.5 | 13.1201 | TRUE | 18.2765 | 32.3563 | NA | NA | 1.6446 | NA | NA | 2.2692
Q9UEW8 | STE20/SPS1-related proline-alanine-rich protein kinase | STK39 | 1:1 | Block_2 | 2.0976 | 15.0237 | TRUE | NA | NA | NA | NA | 3.5377 | NA | 3.7195 | 5.511
Q9H169 | Stathmin-4 | STMN4 | 1:1 | Block_2 | 1.0301 | 11.8245 | TRUE | 2.2658 | 4.7127 | NA | NA | NA | NA | NA | 1.3542
Q9P2F5 | Storkhead-box protein 2 | STOX2 | 1:1 | Block_2 | 1.6337 | 12.1977 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5TH74 | O(6)-methylguanine-induced apoptosis 2 | STPG1 | 1:1 | Block_2 | 1.3851 | 12.4737 | TRUE | NA | NA | NA | NA | 2.2834 | NA | NA | 1.9145
P56962 | Syntaxin-17 | STX17 | 1:1 | Block_2 | 0.59125 | 6.7956 | FALSE | NA | NA | NA | NA | 2.7464 | NA | 2.4012 | 4.5792
Q9P2R7 | Succinate--CoA ligase [ADP-forming] subunit beta, mitochondrial | SUCLA2 | 1:1 | Block_2 | 0.7937 | 10.4236 | TRUE | NA | NA | NA | NA | 1.9323 | NA | NA | 2.4269
Q8IWU5 | Extracellular sulfatase Sulf-2 | SULF2 | 1:1 | Block_2 | 1.7179 | 10.4973 | TRUE | NA | NA | NA | NA | 2.1921 | NA | NA | 1.889
P63165 | Small ubiquitin-related modifier 1 | SUMO1 | 1:1 | Block_2 | 2 | 11.9738 | TRUE | NA | NA | NA | NA | NA | NA | 2.0811 | 4.067
Q8NEM7 | Transcription factor SPT20 homolog | SUPT20H | 1:1 | Block_2 | 1 | 12.4804 | TRUE | NA | NA | NA | NA | 1.3617 | NA | NA | 1.3695
Q8IZU3 | Synaptonemal complex protein 3 | SYCP3 | 1:1 | Block_2 | 2.0275 | 8.9049 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P43405 | Tyrosine-protein kinase SYK | SYK | 1:1 | Block_2 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZMZ3 | Nesprin-3 | SYNE3 | 1:1 | Block_2 | 0.61112 | 10.7771 | FALSE | NA | NA | NA | NA | 1.0692 | NA | NA | 2.2713
Q8IYJ3 | Synaptotagmin-like protein 1 | SYTL1 | 1:1 | Block_2 | 1 | 13.7019 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.7694
Q9HCH5 | Synaptotagmin-like protein 2 | SYTL2 | 1:1 | Block_2 | -2 | 10.6886 | TRUE | NA | NA | NA | NA | NA | NA | NA | -1.573
O95359 | Transforming acidic coiled-coil-containing protein 2 | TACC2 | 1:1 | Block_2 | 1.3796 | 9.9563 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.6918
Q15544 | Transcription initiation factor TFIID subunit 11 | TAF11 | 1:1 | Block_2 | 1.125 | 8.8013 | FALSE | NA | NA | NA | NA | 1.3434 | NA | NA | 2.3159
Q7Z5A9 | Chemokine-like protein TAFA-1 | TAFA1 | 1:1 | Block_2 | 0.19524 | 11.5629 | TRUE | NA | 3.0934 | NA | NA | 0.68652 | NA | NA | 1.0783
Q7Z5A8 | Chemokine-like protein TAFA-3 | TAFA3 | 1:1 | Block_2 | 0.62488 | 9.4151 | TRUE | NA | NA | NA | NA | 0.68193 | NA | NA | 1.067
Q9C0D5 | Protein TANC1 | TANC1 | 1:1 | Block_2 | 2.0422 | 12.0378 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.2865
Q13148 | TAR DNA-binding protein 43 | TARDBP | 1:1 | Block_2 | 2.1211 | 12.5876 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.1836
A2RTX5 | Threonine--tRNA ligase 2, cytoplasmic | TARS3 | 1:1 | Block_2 | 0.82556 | 11.3328 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NVG8 | TBC1 domain family member 13 | TBC1D13 | 1:1 | Block_2 | 0.5 | 11.1023 | TRUE | 12.1879 | 3.4408 | NA | 1.6096 | 5.1495 | 1.8211 | 4.5169 | 6.9955
Q9NU19 | TBC1 domain family member 22B | TBC1D22B | 1:1 | Block_2 | -0.82733 | 10.3361 | FALSE | 13.9199 | 18.0577 | -0.54037 | 0.011779 | 1.1397 | -0.49149 | 0.32889 | 1.4583
Q2M2D7 | TBC1 domain family member 28 | TBC1D28 | 1:1 | Block_2 | 1.0995 | 12.7727 | TRUE | NA | NA | NA | NA | 1.14 | NA | NA | 1.5626
Q96DN5 | TBC1 domain family member 31 | TBC1D31 | 1:1 | Block_2 | 0.44625 | 10.6943 | TRUE | NA | 1.2675 | NA | NA | 0.44839 | NA | NA | 0.91153
Q96NH3 | Protein broad-minded | TBC1D32 | 1:1 | Block_2 | 2.3191 | 13.123 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q13207 | T-box transcription factor TBX2 | TBX2 | 1:1 | Block_2 | 1.8112 | 11.3539 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q15560 | Transcription elongation factor A protein 2 | TCEA2 | 1:1 | Block_2 | 1.0753 | 10.3265 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.4034
Q15170 | Transcription elongation factor A protein-like 1 | TCEAL1 | 1:1 | Block_2 | 2.2912 | 12.3463 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.7777
Q9BQ70 | Transcription factor 25 | TCF25 | 1:1 | Block_2 | -0.36987 | 12.6209 | TRUE | 34.1596 | 7.2323 | NA | -0.074455 | 1.7514 | 0.20978 | 2.0131 | 4.6794
Q07283 | Trichohyalin | TCHH | 1:1 | Block_2 | -2.8729 | 10.6458 | TRUE | 6.6489 | 18.8957 | NA | -2.7202 | -0.97527 | NA | NA | -1.1801
Q5QJ38 | Trichohyalin-like protein 1 | TCHHL1 | 1:1 | Block_2 | 0.5 | 11.545 | TRUE | 8.5187 | 13.2208 | NA | NA | 1.0084 | NA | NA | 1.7004
Q9BT92 | Trichoplein keratin filament-binding protein | TCHP | 1:1 | Block_2 | 1.3061 | 11.9593 | TRUE | NA | NA | NA | NA | 1.733 | NA | NA | 2.6446
O95988 | T-cell leukemia/lymphoma protein 1B | TCL1B | 1:1 | Block_2 | 1 | 12.1337 | FALSE | 7.0142 | 16.9054 | NA | NA | 1.1152 | NA | NA | 1.8325
Q96PL2 | Beta-tectorin | TECTB | 1:1 | Block_2 | 1.2831 | 13.6348 | TRUE | NA | NA | NA | NA | 1.517 | NA | NA | 1.6237
Q7L2K0 | Tubulin epsilon and delta complex protein 2 | TEDC2 | 1:1 | Block_2 | 0.79657 | 12.476 | TRUE | NA | NA | NA | NA | 0.98945 | NA | NA | NA
Q969V4 | Tektin-1 | TEKT1 | 1:1 | Block_2 | 0.97262 | 10.1798 | TRUE | 1.7884 | 7.7 | NA | NA | NA | NA | NA | 1.2141
Q9BXF9 | Tektin-3 | TEKT3 | 1:1 | Block_2 | 0.26911 | 11.9207 | FALSE | 10.977 | 5.5741 | NA | NA | 1.5529 | NA | NA | 0.63717
Q8IYF3 | Testis-expressed protein 11 | TEX11 | 1:1 | Block_2 | 0.91413 | 12.9992 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5158
Q9BXU0 | Testis-expressed protein 12 | TEX12 | 1:1 | Block_2 | 1 | 9.4166 | FALSE | NA | NA | NA | NA | 2.1131 | NA | NA | 1.6694
Q9BXU3 | Testis-expressed protein 13A | TEX13A | 1:1 | Block_2 | 1.6284 | 13.1282 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.0696
Q53QW1 | Testis-expressed protein 44 | TEX44 | 1:1 | Block_2 | 1.0384 | 10.5056 | TRUE | 34.7314 | NA | NA | NA | 1.3284 | NA | NA | 1.6642
Q8N6V9 | Testis-expressed protein 9 | TEX9 | 1:1 | Block_2 | 1.9486 | 13.9148 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5H9I0 | Transcription factor Dp family member 3 | TFDP3 | 1:1 | Block_2 | 0.056441 | 9.9989 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P19484 | Transcription factor EB | TFEB | 1:1 | Block_2 | -1.1632 | 10.4278 | FALSE | 12.1618 | 15.5456 | NA | -1.0571 | -0.028813 | NA | -0.2902 | 3.0576
Q96RS0 | Trimethylguanosine synthase | TGS1 | 1:1 | Block_2 | 1.627 | 6.476 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WY91 | Peroxynitrite isomerase THAP4 | THAP4 | 1:1 | Block_2 | 1.6277 | 10.7309 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UPZ6 | Thrombospondin type-1 domain-containing protein 7A | THSD7A | 1:1 | Block_2 | 1.7927 | 13.4118 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8915
P31314 | T-cell leukemia homeobox protein 1 | TLX1 | 1:1 | Block_2 | -0.46484 | 11.4464 | FALSE | NA | NA | NA | NA | -0.2862 | NA | NA | 0.42998
A2VDJ0 | Transmembrane protein 131-like | TMEM131L | 1:1 | Block_2 | 2.7456 | 10.282 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NZR1 | Tropomodulin-2 | TMOD2 | 1:1 | Block_2 | 0.83053 | 12.7996 | TRUE | 14.5902 | 9.5108 | NA | NA | 1.191 | NA | NA | 1.9946
Q9UPQ9 | Trinucleotide repeat-containing gene 6B protein | TNRC6B | 1:1 | Block_2 | 1.82 | 13.1105 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.9088
Q9Y4F4 | TOG array regulator of axonemal microtubules protein 1 | TOGARAM1 | 1:1 | Block_2 | 0.74472 | 10.3747 | TRUE | NA | 1.1679 | NA | NA | NA | NA | NA | NA
Q15785 | Mitochondrial import receptor subunit TOM34 | TOMM34 | 1:1 | Block_2 | 0.001 | 12.0823 | FALSE | 33.2401 | 8.0885 | NA | 0.83416 | 3.5151 | 0.33088 | 3.4071 | 5.5197
Q13625 | Apoptosis-stimulating of p53 protein 2 | TP53BP2 | 1:1 | Block_2 | 0.57771 | 11.8368 | FALSE | NA | NA | NA | 0.70098 | 2.286 | NA | 1.7323 | 3.1542
Q96J77 | Tumor protein D55 | TPD52L3 | 1:1 | Block_2 | 3.0183 | 9.8996 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q68CL5 | Tubulin polyglutamylase complex subunit 2 | TPGS2 | 1:1 | Block_2 | 1 | 11.251 | TRUE | 12.8496 | 17.661 | NA | NA | 1.4865 | NA | NA | 2.2728
Q9ULW0 | Targeting protein for Xklp2 | TPX2 | 1:1 | Block_2 | 2.2987 | 12.6351 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O14545 | TRAF-type zinc finger domain-containing protein 1 | TRAFD1 | 1:1 | Block_2 | 1 | 8.6987 | FALSE | NA | 10.0245 | NA | NA | 2.1427 | NA | 1.8198 | 3.7379
Q96PN7 | Transcriptional-regulating factor 1 | TRERF1 | 1:1 | Block_2 | 0.5 | 12.0352 | FALSE | 14.0078 | 28.0783 | NA | NA | 0.57773 | NA | NA | 1.8043
Q14142 | Tripartite motif-containing protein 14 | TRIM14 | 1:1 | Block_2 | 1.146 | 7.7211 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.3639
Q8WV44 | E3 ubiquitin-protein ligase TRIM41 | TRIM41 | 1:1 | Block_2 | 0.5 | 12.7049 | TRUE | NA | NA | NA | NA | 1.519 | NA | 0.53443 | 2.5069
Q96LD4 | E3 ubiquitin-protein ligase TRIM47 | TRIM47 | 1:1 | Block_2 | 0.5437 | 12.2559 | TRUE | 6.5215 | 6.1569 | NA | NA | 1.2702 | NA | 0.73948 | 1.9562
Q9BYV2 | Tripartite motif-containing protein 54 | TRIM54 | 1:1 | Block_2 | 1.8803 | 12.0105 | TRUE | NA | NA | NA | NA | 2.5649 | NA | NA | 1.993
Q495X7 | Tripartite motif-containing protein 60 | TRIM60 | 1:1 | Block_2 | 0.14289 | 9.4174 | TRUE | 5.9009 | 6.8845 | NA | NA | 0.63145 | NA | 0.2091 | 0.80905
Q9H2D6 | TRIO and F-actin-binding protein | TRIOBP | 1:1 | Block_2 | -0.078805 | 8.5584 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q15642 | Cdc42-interacting protein 4 | TRIP10 | 1:1 | Block_2 | 1.7565 | 10.7249 | TRUE | NA | NA | NA | NA | 3.1202 | NA | 2.4413 | 4.7134
Q15645 | Pachytene checkpoint protein 2 homolog | TRIP13 | 1:1 | Block_2 | 3.3259 | 11.5056 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.8084
Q9P272 | Probable tRNA methyltransferase 9B | TRMT9B | 1:1 | Block_2 | 0.5 | 13.6574 | FALSE | 36.0647 | 19.9337 | NA | NA | 1.9961 | NA | NA | 2.0849
Q8TEL6 | Short transient receptor potential channel 4-associated protein | TRPC4AP | 1:1 | Block_2 | 1 | 12.1678 | FALSE | NA | 26.1354 | NA | NA | 1.2648 | NA | NA | 1.7234
A6NMA1 | Putative uncharacterized protein TRPC5OS | TRPC5OS | 1:1 | Block_2 | 1 | 9.9346 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.9451
Q8NER1 | Transient receptor potential cation channel subfamily V member 1 | TRPV1 | 1:1 | Block_2 | 1.2103 | 9.8134 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q99576 | TSC22 domain family protein 3 | TSC22D3 | 1:1 | Block_2 | -1.2377 | 11.1107 | TRUE | 8.2663 | 14.3594 | -0.67227 | 1.4052 | 4.4205 | 1.5125 | 4.1662 | 6.3056
Q9BZW7 | Testis-specific gene 10 protein | TSGA10 | 1:1 | Block_2 | 0.9387 | 12.6823 | TRUE | NA | NA | NA | NA | 0.98413 | NA | NA | 0.9426
Q9NRE2 | Teashirt homolog 2 | TSHZ2 | 1:1 | Block_2 | 0.68695 | 11.2559 | FALSE | NA | NA | NA | NA | NA | NA | NA | 0.71928
O95857 | Tetraspanin-13 | TSPAN13 | 1:1 | Block_2 | 0.5 | 9.8391 | TRUE | 13.0966 | 3.5825 | NA | NA | 1.0541 | NA | NA | 1.4742
O95153 | Peripheral-type benzodiazepine receptor-associated protein 1 | TSPOAP1 | 1:1 | Block_2 | 0.93106 | 10.6023 | FALSE | 17.3662 | 12.382 | NA | NA | 0.97248 | NA | NA | 1.4936
P0CV98_P0CW01 | Testis-specific Y-encoded protein 3/10 | TSPY3_TSPY10 | 1:1 | Block_2 | 0.81841 | 13.3313 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.82122
Q86VY4 | Testis-specific Y-encoded-like protein 5 | TSPYL5 | 1:1 | Block_2 | 0.68097 | 10.9902 | TRUE | 8.6759 | 7.4082 | NA | NA | 0.79121 | NA | NA | 1.1102
Q9Y5U2 | Protein TSSC4 | TSSC4 | 1:1 | Block_2 | 0.18464 | 9.7889 | TRUE | NA | 9.8812 | NA | 0.84792 | 2.2863 | 0.63822 | 2.8436 | 4.4865
Q9H892 | Tetratricopeptide repeat protein 12 | TTC12 | 1:1 | Block_2 | 2.172 | 11.9018 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96N46 | Tetratricopeptide repeat protein 14 | TTC14 | 1:1 | Block_2 | 1.882 | 13.2188 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NEE8 | Tetratricopeptide repeat protein 16 | TTC16 | 1:1 | Block_2 | 1.3308 | 9.7686 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4309
Q8NDW8 | Tetratricopeptide repeat protein 21A | TTC21A | 1:1 | Block_2 | 0.93139 | 11.8935 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N0Z6 | Tetratricopeptide repeat protein 5 | TTC5 | 1:1 | Block_2 | 1.3326 | 8.8323 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q14679 | Tubulin monoglutamylase TTLL4 | TTLL4 | 1:1 | Block_2 | 0.21246 | 13.0942 | TRUE | 12.9313 | 8.7261 | NA | NA | 0.78785 | NA | 0.28782 | 1.1756
Q9BQE3 | Tubulin alpha-1C chain | TUBA1C | 1:1 | Block_2 | 0.5 | 9.3256 | TRUE | 6.4918 | 18.611 | NA | NA | 0.97403 | NA | NA | 1.3809
Q9NRH3 | Tubulin gamma-2 chain | TUBG2 | 1:1 | Block_2 | 1.3435 | 8.5186 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NRJ4 | Tubby-related protein 4 | TULP4 | 1:1 | Block_2 | 0.59581 | 11.3418 | TRUE | NA | NA | NA | NA | 0.98115 | NA | NA | 1.1151
Q96RR1 | Twinkle mtDNA helicase | TWNK | 1:1 | Block_2 | 1.2989 | 11.2648 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BZA5 | Putative gamma-taxilin 2 | TXLNGY | 1:1 | Block_2 | 0.5 | 13.9733 | TRUE | 11.0937 | 31.1327 | NA | NA | 1.8765 | NA | 0.9232 | 4.7164
Q6F5E7 | Putative uncharacterized protein TXNRD3NB | TXNRD3NB | 1:1 | Block_2 | 0.55787 | 8.5057 | FALSE | NA | NA | NA | NA | 0.83 | NA | NA | 1.0668
Q9NV66 | S-adenosyl-L-methionine-dependent tRNA 4-demethylwyosine synthase TYW1 | TYW1 | 1:1 | Block_2 | 0.5 | 12.5383 | FALSE | NA | 20.1008 | NA | NA | 1.5692 | NA | NA | 1.8409
A2RUC4 | tRNA wybutosine-synthesizing protein 5 | TYW5 | 1:1 | Block_2 | 0.29227 | 9.8282 | TRUE | 15.7161 | 31.2448 | NA | 0.9684 | 2.251 | NA | 2.6996 | 4.4317
P22314 | Ubiquitin-like modifier-activating enzyme 1 | UBA1 | 1:1 | Block_2 | 1 | 8.1251 | TRUE | 4.6895 | 10.6623 | NA | NA | 1.1054 | NA | NA | 1.911
Q5T6F2 | Ubiquitin-associated protein 2 | UBAP2 | 1:1 | Block_2 | 0.001 | 12.2979 | TRUE | 12.6616 | 17.7483 | NA | 0.31101 | 1.6888 | NA | 1.1319 | 2.5962
P57075 | Ubiquitin-associated and SH3 domain-containing protein A | UBASH3A | 1:1 | Block_2 | 1.0507 | 10.2289 | FALSE | NA | 15.7743 | NA | NA | NA | NA | NA | 2.0291
P62837 | Ubiquitin-conjugating enzyme E2 D2 | UBE2D2 | 1:1 | Block_2 | 1.4071 | 10.5652 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5037
O95155 | Ubiquitin conjugation factor E4 B | UBE4B | 1:1 | Block_2 | 1.2877 | 11.3405 | TRUE | NA | NA | NA | NA | 1.4652 | NA | NA | 1.6727
O94941 | RING finger protein 37 | UBOX5 | 1:1 | Block_2 | 0.75231 | 8.5358 | FALSE | NA | NA | NA | NA | 1.1943 | NA | 0.97975 | 2.6556
Q8IYU4 | Ubiquilin-like protein | UBQLNL | 1:1 | Block_2 | 0.09833 | 12.4284 | TRUE | NA | 3.1358 | NA | NA | 0.28865 | NA | NA | 1.1778
Q8IWV8 | E3 ubiquitin-protein ligase UBR2 | UBR2 | 1:1 | Block_2 | 1 | 12.1331 | TRUE | NA | 45.4133 | NA | NA | 2.5095 | NA | 2.5479 | 4.3649
Q5T124 | UBX domain-containing protein 11 | UBXN11 | 1:1 | Block_2 | 1.8704 | 13.5803 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8854
Q9BZV1 | UBX domain-containing protein 6 | UBXN6 | 1:1 | Block_2 | 0.01775 | 7.0008 | TRUE | NA | NA | NA | NA | 0.62902 | NA | 0.54391 | 2.4916
Q8WVF2 | Unique cartilage matrix-associated protein | UCMA | 1:1 | Block_2 | 1.4255 | 8.0513 | TRUE | NA | 21.4178 | NA | NA | NA | NA | NA | NA
Q8IYT8 | Serine/threonine-protein kinase ULK2 | ULK2 | 1:1 | Block_2 | 0.67681 | 11.1834 | TRUE | 25.8833 | 15.3522 | NA | NA | 1.3 | NA | NA | 1.1559
Q96C45 | Serine/threonine-protein kinase ULK4 | ULK4 | 1:1 | Block_2 | 0.02568 | 10.6101 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.85369
Q9H3U1 | Protein unc-45 homolog A | UNC45A | 1:1 | Block_2 | 1.4962 | 13.0034 | TRUE | 8.5432 | 15.4571 | NA | NA | 3.1148 | NA | 3.2833 | 5.3439
Q9H9P5 | Putative E3 ubiquitin-protein ligase UNKL | UNKL | 1:1 | Block_2 | 1.646 | 12.2892 | FALSE | NA | NA | NA | NA | 2.262 | NA | NA | 2.8609
Q9BRT2 | Ubiquinol-cytochrome-c reductase complex assembly factor 2 | UQCC2 | 1:1 | Block_2 | 2.3989 | 11.7671 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6NGE7 | Putative 2-oxo-4-hydroxy-4-carboxy-5-ureidoimidazoline decarboxylase | URAD | 1:1 | Block_2 | 1 | 11.7663 | TRUE | 17.7804 | 11.1475 | NA | NA | 1.6862 | NA | NA | 1.4077
Q9Y6N9 | Harmonin | USH1C | 1:1 | Block_2 | 0.64794 | 15.2747 | TRUE | NA | 2.9537 | NA | NA | 0.78098 | NA | NA | 1.7381
Q495M9 | pre-mRNA splicing regulator USH1G | USH1G | 1:1 | Block_2 | 2.2779 | 16.1445 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.2927
Q70CQ4 | Ubiquitin carboxyl-terminal hydrolase 31 | USP31 | 1:1 | Block_2 | 0.5 | 12.9504 | TRUE | 12.0097 | 15.6357 | NA | NA | 2.4489 | NA | 1.1411 | 3.1117
Q8TEY7 | Ubiquitin carboxyl-terminal hydrolase 33 | USP33 | 1:1 | Block_2 | 2.4986 | 13.8717 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q70CQ2 | Ubiquitin carboxyl-terminal hydrolase 34 | USP34 | 1:1 | Block_2 | 2.2258 | 11.2541 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86T82 | Ubiquitin carboxyl-terminal hydrolase 37 | USP37 | 1:1 | Block_2 | -0.73933 | 11.3273 | TRUE | 20.3688 | 8.5744 | NA | NA | -0.72804 | NA | NA | 0.3062
Q93008 | Probable ubiquitin carboxyl-terminal hydrolase FAF-X | USP9X | 1:1 | Block_2 | 1.7089 | 7.351 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O75379 | Vesicle-associated membrane protein 4 | VAMP4 | 1:1 | Block_2 | 0.57999 | 8.9345 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.0092
Q5SQQ9 | Ventral anterior homeobox 1 | VAX1 | 1:1 | Block_2 | 0.001 | 11.9122 | TRUE | NA | NA | NA | NA | 0.77464 | NA | NA | 0.46104
Q9UIW0 | Ventral anterior homeobox 2 | VAX2 | 1:1 | Block_2 | 0.61902 | 8.6919 | TRUE | 19.8979 | 20.2743 | NA | NA | NA | NA | NA | 2.0779
P09327 | Villin-1 | VIL1 | 1:1 | Block_2 | 0.5 | 9.9797 | TRUE | 13.2547 | 36.8501 | NA | 0.63623 | 2.6832 | NA | 2.2178 | 4.1906
P08670 | Vimentin | VIM | 1:1 | Block_2 | -1.79 | 2.227 | FALSE | 24.0541 | 28.2749 | NA | -1.4669 | -0.19742 | -1.5252 | -0.25252 | 1.5518
Q9H269 | Vacuolar protein sorting-associated protein 16 homolog | VPS16 | 1:1 | Block_2 | 1.7153 | 10.6254 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.6236
Q9P1Q0 | Vacuolar protein sorting-associated protein 54 | VPS54 | 1:1 | Block_2 | 0.35697 | 12.4067 | TRUE | NA | NA | NA | NA | 0.49159 | NA | 0.98603 | 3.526
Q99986 | Serine/threonine-protein kinase VRK1 | VRK1 | 1:1 | Block_2 | 0.4419 | 10.1681 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.2529
Q8IV63 | Inactive serine/threonine-protein kinase VRK3 | VRK3 | 1:1 | Block_2 | -1.6676 | 9.1883 | TRUE | NA | NA | NA | NA | NA | NA | NA | -0.8325
Q8IW00 | V-set and transmembrane domain-containing protein 4 | VSTM4 | 1:1 | Block_2 | -0.4185 | 6.099 | FALSE | 13.8204 | 22.5528 | NA | -0.1763 | 0.73456 | NA | 0.12426 | 1.4255
Q6ICG8 | Postacrosomal sheath WW domain-binding protein | WBP2NL | 1:1 | Block_2 | 2 | 8.0723 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H6R7 | WD repeat and coiled-coil-containing protein | WDCP | 1:1 | Block_2 | 1.0664 | 12.145 | TRUE | 29.232 | 21.9974 | NA | NA | 1.1301 | NA | NA | 2.9975
O75717 | WD repeat and HMG-box DNA-binding protein 1 | WDHD1 | 1:1 | Block_2 | 1.6371 | 11.1701 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TBZ3 | WD repeat-containing protein 20 | WDR20 | 1:1 | Block_2 | -2.1358 | 5.2591 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96S15 | GATOR complex protein WDR24 | WDR24 | 1:1 | Block_2 | 1.4148 | 12.5555 | TRUE | 20.4017 | 19.3046 | NA | NA | NA | NA | NA | 2.4362
Q7Z5U6 | WD repeat-containing protein 53 | WDR53 | 1:1 | Block_2 | 1 | 14.192 | TRUE | 23.0206 | 38.213 | NA | NA | NA | NA | NA | 2.318
Q9NW82 | WD repeat-containing protein 70 | WDR70 | 1:1 | Block_2 | 1.2762 | 12.1921 | TRUE | NA | 0.81756 | NA | NA | NA | NA | NA | 2.0257
Q8IUB2 | WAP four-disulfide core domain protein 3 | WFDC3 | 1:1 | Block_2 | -1 | 9.9528 | TRUE | 17.8889 | 11.1359 | NA | NA | 0.67123 | NA | NA | 0.81026
Q9BYP7 | Serine/threonine-protein kinase WNK3 | WNK3 | 1:1 | Block_2 | 0.28761 | 12.1211 | TRUE | 9.8741 | NA | NA | NA | 0.43633 | NA | NA | 1.4539
Q9UBV4 | Protein Wnt-16 | WNT16 | 1:1 | Block_2 | 1 | 14.0869 | FALSE | 19.2837 | 11.5362 | NA | NA | 1.6993 | NA | NA | 2.0541
Q14191 | Bifunctional 3'-5' exonuclease/ATP-dependent helicase WRN | WRN | 1:1 | Block_2 | 0.43916 | 12.729 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.152
Q8IX03 | Protein KIBRA | WWC1 | 1:1 | Block_2 | 0.4811 | 12.7075 | TRUE | NA | NA | NA | NA | 1.1726 | NA | NA | 1.9893
Q6AWC2 | Protein WWC2 | WWC2 | 1:1 | Block_2 | 1.3091 | 12.6618 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H0M0 | NEDD4-like E3 ubiquitin-protein ligase WWP1 | WWP1 | 1:1 | Block_2 | 1.3545 | 9.2056 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A4UGR9 | Xin actin-binding repeat-containing protein 2 | XIRP2 | 1:1 | Block_2 | 0.98592 | 12.0366 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.3658
P18887 | DNA repair protein XRCC1 | XRCC1 | 1:1 | Block_2 | 1.0871 | 11.6513 | TRUE | NA | 5.6503 | NA | NA | NA | NA | NA | 1.4756
O43543 | DNA repair protein XRCC2 | XRCC2 | 1:1 | Block_2 | 1.9898 | 12.6704 | FALSE | 9.3827 | NA | NA | NA | NA | NA | NA | 2.3103
A8MUA0_A8MX80 | Putative UPF0607 protein ENSP00000383144/ENSP00000381514 | YB057_YM017 | 1:1 | Block_2 | 0.51 | 11.1533 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.6468
Q9Y2T7 | Y-box-binding protein 2 | YBX2 | 1:1 | Block_2 | 1 | 11.9832 | TRUE | NA | 12.7745 | NA | NA | 1.5962 | NA | NA | 1.9452
P13994 | Probable splicing factor YJU2B | YJU2B | 1:1 | Block_2 | -2.0165 | 7.1887 | FALSE | NA | NA | NA | NA | NA | NA | NA | -0.91361
O15498 | Synaptobrevin homolog YKT6 | YKT6 | 1:1 | Block_2 | 0.38604 | 10.18 | TRUE | NA | NA | NA | NA | 0.40356 | NA | NA | 2.4594
Q6ZSR3 | Putative uncharacterized protein FLJ45275, mitochondrial | YO027 | 1:1 | Block_2 | 1.79 | 12.7939 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.1151
Q6ZRN7 | Putative uncharacterized protein FLJ46214 | YP029 | 1:1 | Block_2 | 0.55387 | 11.5384 | TRUE | NA | NA | NA | NA | 1.1614 | NA | NA | 1.3706
A8MT66 | Putative uncharacterized protein ENSP00000383407 | YU005 | 1:1 | Block_2 | 0.51933 | 12.6408 | TRUE | 2.0427 | 6.1244 | NA | NA | 0.78402 | NA | NA | 1.696
Q9H869 | YY1-associated protein 1 | YY1AP1 | 1:1 | Block_2 | 2.2328 | 14.0819 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O15391 | Transcription factor YY2 | YY2 | 1:1 | Block_2 | 1.3615 | 13.191 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O75132 | Zinc finger BED domain-containing protein 4 | ZBED4 | 1:1 | Block_2 | 2.2536 | 14.4495 | TRUE | 12.3559 | NA | NA | NA | 2.6693 | NA | NA | NA
O43167 | Zinc finger and BTB domain-containing protein 24 | ZBTB24 | 1:1 | Block_2 | 3.5338 | 13.4852 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y2Y4 | Zinc finger and BTB domain-containing protein 32 | ZBTB32 | 1:1 | Block_2 | -2.1338 | 12.5896 | TRUE | NA | NA | NA | NA | -1.7028 | NA | NA | -1.3244
P10074 | Telomere zinc finger-associated protein | ZBTB48 | 1:1 | Block_2 | 1.3533 | 7.493 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZSB9 | Zinc finger and BTB domain-containing protein 49 | ZBTB49 | 1:1 | Block_2 | 2.0445 | 13.7764 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O15156 | Zinc finger and BTB domain-containing protein 7B | ZBTB7B | 1:1 | Block_2 | -0.88388 | 6.9886 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A1YPR0 | Zinc finger and BTB domain-containing protein 7C | ZBTB7C | 1:1 | Block_2 | 0.7521 | 12.3088 | TRUE | NA | NA | NA | NA | 1.0517 | NA | NA | 1.6544
A2A288 | Probable ribonuclease ZC3H12D | ZC3H12D | 1:1 | Block_2 | 0.5 | 13.0177 | FALSE | NA | 10.8849 | NA | NA | 1.071 | NA | 1.076 | 2.7256
P61129 | Zinc finger CCCH domain-containing protein 6 | ZC3H6 | 1:1 | Block_2 | 0.21386 | 13.124 | TRUE | NA | NA | NA | NA | 1.579 | NA | 0.85437 | 3.0417
P0CG32 | Zinc finger CCHC domain-containing protein 18 | ZCCHC18 | 1:1 | Block_2 | 0.98972 | 11.7309 | TRUE | 15.6976 | 16.757 | NA | NA | 1.9757 | NA | NA | 1.0951
Q9C0B9 | Zinc finger CCHC domain-containing protein 2 | ZCCHC2 | 1:1 | Block_2 | 1.5492 | 13.4149 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H8U3 | AN1-type zinc finger protein 3 | ZFAND3 | 1:1 | Block_2 | 1 | 7.7563 | FALSE | 14.2377 | 9.9577 | NA | NA | 1.1307 | NA | NA | 1.9236
P17010 | Zinc finger X-chromosomal protein | ZFX | 1:1 | Block_2 | 0.63485 | 7.5896 | TRUE | 19.0559 | 11.0252 | NA | NA | NA | NA | NA | NA
Q7Z3T8 | Zinc finger FYVE domain-containing protein 16 | ZFYVE16 | 1:1 | Block_2 | 1.0278 | 13.4515 | TRUE | NA | NA | NA | NA | 1.2971 | NA | 1.1179 | 3.2754
O95405 | Zinc finger FYVE domain-containing protein 9 | ZFYVE9 | 1:1 | Block_2 | 0.9499 | 11.752 | TRUE | NA | 6.9697 | NA | NA | 2.3066 | NA | NA | 2.227
O60481 | Zinc finger protein ZIC 3 | ZIC3 | 1:1 | Block_2 | 3 | 10.4008 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9HA38 | Zinc finger matrin-type protein 3 | ZMAT3 | 1:1 | Block_2 | 1.2806 | 9.3878 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q15326 | Zinc finger MYND domain-containing protein 11 | ZMYND11 | 1:1 | Block_2 | -0.32248 | 12.5872 | TRUE | 3.2618 | NA | NA | NA | 0.25247 | NA | 0.2712 | 2.3381
Q9Y473 | Zinc finger protein 175 | ZNF175 | 1:1 | Block_2 | 0.001 | 12.745 | TRUE | 14.7231 | 13.9158 | NA | NA | 2.7345 | NA | NA | 1.3508
O95125 | Zinc finger protein 202 | ZNF202 | 1:1 | Block_2 | 0.18039 | 10.8446 | TRUE | NA | NA | NA | NA | 1.0091 | NA | NA | 0.26719
O75362 | Zinc finger protein 217 | ZNF217 | 1:1 | Block_2 | 0.56837 | 9.5609 | FALSE | NA | NA | NA | NA | 1.7217 | NA | NA | 1.3212
O43296 | Zinc finger protein 264 | ZNF264 | 1:1 | Block_2 | 0.5 | 11.9869 | TRUE | NA | 4.6592 | NA | NA | NA | NA | NA | 1.2421
Q9UDV7 | Zinc finger protein 282 | ZNF282 | 1:1 | Block_2 | -0.28289 | 14.2598 | TRUE | 28.4214 | 19.544 | NA | NA | 0.81937 | NA | -0.10115 | 1.3161
P13682 | Zinc finger protein 35 | ZNF35 | 1:1 | Block_2 | -0.059027 | 7.0317 | TRUE | NA | NA | NA | NA | 0.3189 | NA | NA | 0.69573
Q96PM9 | Zinc finger protein 385A | ZNF385A | 1:1 | Block_2 | 1.2524 | 12.7277 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.3668
Q569K4 | Zinc finger protein 385B | ZNF385B | 1:1 | Block_2 | -0.081208 | 13.1215 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.0525
Q53GI3 | Zinc finger protein 394 | ZNF394 | 1:1 | Block_2 | 0.001 | 5.4803 | FALSE | NA | NA | NA | NA | 0.64509 | NA | NA | 0.75515
Q8NF99 | Zinc finger protein 397 | ZNF397 | 1:1 | Block_2 | 1.1219 | 11.3287 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.291
Q9ULD9 | Zinc finger protein 608 | ZNF608 | 1:1 | Block_2 | 1.2172 | 14.3927 | TRUE | 10.721 | 4.3236 | NA | NA | NA | NA | NA | 1.3733
Q8N1G0 | Zinc finger protein 687 | ZNF687 | 1:1 | Block_2 | 1.8215 | 12.179 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.0156
Q6NUN9 | Zinc finger protein 746 | ZNF746 | 1:1 | Block_2 | 0.001 | 13.5066 | TRUE | 15.8332 | 18.8406 | NA | NA | 0.39478 | NA | 0.026996 | 1.5458
Q96N20 | Zinc finger protein 75A | ZNF75A | 1:1 | Block_2 | 0.7114 | 13.2591 | TRUE | NA | NA | NA | NA | 2.0974 | NA | NA | 1.7966
Q5JPB2 | Zinc finger protein 831 | ZNF831 | 1:1 | Block_2 | 0.33356 | 11.8628 | TRUE | 9.285 | 10.6982 | NA | 0.61667 | 2.6843 | NA | NA | 1.6152
A8K0R7 | Zinc finger protein 839 | ZNF839 | 1:1 | Block_2 | 0.5 | 10.9551 | FALSE | 17.8149 | 4.5942 | NA | NA | 1.2881 | NA | NA | 1.2026
Q8IUF1 | Zinc-regulated GTPase metalloprotein activator 1B | ZNG1B | 1:1 | Block_2 | 1.3757 | 12.1824 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UGI0 | Ubiquitin thioesterase ZRANB1 | ZRANB1 | 1:1 | Block_2 | 1 | 16.1401 | TRUE | 39.894 | 16.3101 | NA | NA | 2.1197 | NA | 1.5108 | 3.169
Q5FWF4 | DNA annealing helicase and endonuclease ZRANB3 | ZRANB3 | 1:1 | Block_2 | 3.163 | 11.7496 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q15696 | U2 small nuclear ribonucleoprotein auxiliary factor 35 kDa subunit-related protein 2 | ZRSR2 | 1:1 | Block_2 | 0.0531 | 12.3545 | TRUE | 4.2578 | 13.6631 | NA | NA | NA | NA | NA | 2.034
Q9Y5A6 | Zinc finger and SCAN domain-containing protein 21 | ZSCAN21 | 1:1 | Block_2 | 2.6926 | 13.2634 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.6221
Q96LW9 | Zinc finger and SCAN domain-containing protein 31 | ZSCAN31 | 1:1 | Block_2 | 2.8514 | 7.8452 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.8952
Q96MP5 | Zinc finger SWIM domain-containing protein 3 | ZSWIM3 | 1:1 | Block_2 | 1.08 | 12.0751 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NUQ8 | ATP-binding cassette sub-family F member 3 | ABCF3 | 1:1 | Block_1 | 1.0512 | 11.1501 | TRUE | NA | 0.12889 | NA | NA | 1.246 | NA | NA | 1.74
Q8WTS1 | 1-acylglycerol-3-phosphate O-acyltransferase ABHD5 | ABHD5 | 1:1 | Block_1 | 1.2048 | 7.8471 | FALSE | 9.1241 | 14.4632 | NA | NA | 1.6313 | NA | NA | 1.5082
Q96I13 | Protein ABHD8 | ABHD8 | 1:1 | Block_1 | 5.9385 | 13.719 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H845 | Complex I assembly factor ACAD9, mitochondrial | ACAD9 | 1:1 | Block_1 | 4.1695 | 16.4552 | FALSE | NA | NA | NA | NA | NA | NA | NA | 4.1795
Q6NUN0 | Acyl-coenzyme A synthetase ACSM5, mitochondrial | ACSM5 | 1:1 | Block_1 | 2.8745 | 12.3493 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IZF7 | Adhesion G-protein coupled receptor F2 | ADGRF2 | 1:1 | Block_1 | 1.2135 | 11.1012 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ULP2 | Aftiphilin | AFTPH | 1:1 | Block_1 | 0.9293 | 8.4815 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.93032
Q8N157 | Jouberin | AHI1 | 1:1 | Block_1 | 0.57152 | 12.5081 | TRUE | NA | NA | NA | NA | 1.5023 | NA | NA | 1.6091
C9J069 | Apical junction component 1 homolog | AJM1 | 1:1 | Block_1 | 0.65386 | 10.8956 | FALSE | NA | NA | NA | NA | 0.84072 | NA | NA | 1.3361
P24588 | A-kinase anchor protein 5 | AKAP5 | 1:1 | Block_1 | 1.6803 | 10.3171 | TRUE | NA | NA | NA | NA | 1.6911 | NA | NA | NA
Q86TB3 | Alpha-protein kinase 2 | ALPK2 | 1:1 | Block_1 | 1.623 | 10.8479 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q99218 | Amelogenin, Y isoform | AMELY | 1:1 | Block_1 | 1.8955 | 10.0221 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UM13 | Anaphase-promoting complex subunit 10 | ANAPC10 | 1:1 | Block_1 | 1.8654 | 14.0306 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5TYW2 | Ankyrin repeat domain-containing protein 20A1 | ANKRD20A1 | 1:1 | Block_1 | 1.8444 | 13.0137 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6NSI1 | Putative ankyrin repeat domain-containing protein 26-like protein | ANKRD26P1 | 1:1 | Block_1 | 0.87436 | 10.8655 | TRUE | NA | 0.80767 | NA | NA | 1.1792 | NA | NA | 1.0942
O15084 | Serine/threonine-protein phosphatase 6 regulatory ankyrin repeat subunit A | ANKRD28 | 1:1 | Block_1 | 1.2231 | 7.3487 | TRUE | NA | NA | NA | NA | 1.2458 | NA | NA | 1.7112
A5PLL1 | Ankyrin repeat domain-containing protein 34B | ANKRD34B | 1:1 | Block_1 | 1.1896 | 9.3515 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P0C6C1 | Ankyrin repeat domain-containing protein 34C | ANKRD34C | 1:1 | Block_1 | 1.627 | 13.8178 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96IX9 | Putative ankyrin repeat domain-containing protein 26-like 1 | ANKRD36BP1 | 1:1 | Block_1 | 0.067124 | 12.8766 | TRUE | 7.3326 | 10.8452 | NA | NA | 0.33555 | NA | NA | 0.44202
Q5TZF3 | Ankyrin repeat domain-containing protein 45 | ANKRD45 | 1:1 | Block_1 | 3.0272 | 9.1719 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NQW6 | Anillin | ANLN | 1:1 | Block_1 | 1.2366 | 12.9894 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O00213 | Amyloid beta precursor protein binding family B member 1 | APBB1 | 1:1 | Block_1 | 0.70513 | 10.357 | FALSE | NA | NA | NA | NA | 0.91017 | NA | NA | 1.282
Q9UBZ4 | DNA-(apurinic or apyrimidinic site) endonuclease 2 | APEX2 | 1:1 | Block_1 | 4 | 10.8794 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IW19 | Aprataxin and PNK-like factor | APLF | 1:1 | Block_1 | 2.9117 | 14.4049 | TRUE | NA | NA | NA | NA | 4.0504 | NA | 3.8067 | 5.8665
Q9Y235 | C->U-editing enzyme APOBEC-2 | APOBEC2 | 1:1 | Block_1 | 0.89711 | 7.0167 | TRUE | NA | 5.4184 | NA | NA | NA | NA | NA | 1.3836
Q7Z2E3 | Aprataxin | APTX | 1:1 | Block_1 | 1.2295 | 11.1306 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.563
P29972 | Aquaporin-1 | AQP1 | 1:1 | Block_1 | 2.9936 | 13.1111 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O15033 | Apoptosis-resistant E3 ubiquitin protein ligase 1 | AREL1 | 1:1 | Block_1 | 2.9835 | 12.268 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y6D5 | Brefeldin A-inhibited guanine nucleotide-exchange protein 2 | ARFGEF2 | 1:1 | Block_1 | 1.3597 | 9.3098 | TRUE | NA | NA | NA | NA | 1.8757 | NA | 1.4075 | 2.5237
Q9P2F6 | Rho GTPase-activating protein 20 | ARHGAP20 | 1:1 | Block_1 | 0.5771 | 10.8155 | TRUE | NA | NA | NA | NA | 0.95352 | NA | NA | NA
A8MVX0 | Rho guanine nucleotide exchange factor 33 | ARHGEF33 | 1:1 | Block_1 | 2.9115 | 13.6191 | TRUE | 21.44 | NA | NA | NA | NA | NA | NA | NA
Q8TER5 | Rho guanine nucleotide exchange factor 40 | ARHGEF40 | 1:1 | Block_1 | 1.7819 | 10.0372 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N8R7 | ARL14 effector protein | ARL14EP | 1:1 | Block_1 | 2.0917 | 13.6392 | TRUE | NA | 4.7427 | NA | NA | 2.3109 | NA | NA | 2.9576
Q9UBL0 | cAMP-regulated phosphoprotein 21 | ARPP21 | 1:1 | Block_1 | 3.6428 | 10.7283 | TRUE | 7.9355 | NA | NA | NA | NA | NA | NA | NA
P36575 | Arrestin-C | ARR3 | 1:1 | Block_1 | 2.0236 | 10.7927 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O00192 | Splicing regulator ARVCF | ARVCF | 1:1 | Block_1 | 2.9226 | 12.076 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TDY4 | Arf-GAP with SH3 domain, ANK repeat and PH domain-containing protein 3 | ASAP3 | 1:1 | Block_1 | 3.1428 | 12.759 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.6478
Q8WXH4 | Ankyrin repeat and SOCS box protein 11 | ASB11 | 1:1 | Block_1 | 0.55326 | 10.1157 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NWX5 | Ankyrin repeat and SOCS box protein 6 | ASB6 | 1:1 | Block_1 | 3.1192 | 13.3452 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9C0F0 | Putative Polycomb group protein ASXL3 | ASXL3 | 1:1 | Block_1 | 1.7106 | 12.234 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WWH4 | Ankyrin repeat, SAM and basic leucine zipper domain-containing protein 1 | ASZ1 | 1:1 | Block_1 | 0.49717 | 9.6389 | FALSE | 21.8474 | 6.9418 | NA | NA | 0.71244 | NA | NA | 0.92196
P18847 | Cyclic AMP-dependent transcription factor ATF-3 | ATF3 | 1:1 | Block_1 | 6.5461 | 15.713 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q15904 | V-type proton ATPase subunit S1 | ATP6AP1 | 1:1 | Block_1 | 1.5555 | 9.8399 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P46100 | Transcriptional regulator ATRX | ATRX | 1:1 | Block_1 | 1.5943 | 11.964 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.9114
P54253 | Ataxin-1 | ATXN1 | 1:1 | Block_1 | 1.0266 | 7.9878 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O15265 | Ataxin-7 | ATXN7 | 1:1 | Block_1 | 2.9134 | 13.7873 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5T6C5 | Ataxin-7-like protein 2 | ATXN7L2 | 1:1 | Block_1 | 2.4211 | 14.6838 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H7T9 | Aurora kinase A and ninein-interacting protein | AUNIP | 1:1 | Block_1 | 4.88 | 10.9305 | TRUE | 0.25894 | 15.7827 | NA | NA | NA | NA | NA | NA
Q8N9N5 | Protein BANP | BANP | 1:1 | Block_1 | 2.1675 | 11.4966 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NRL2 | Bromodomain adjacent to zinc finger domain protein 1A | BAZ1A | 1:1 | Block_1 | 3.014 | 10.2056 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O43521 | Bcl-2-like protein 11 | BCL2L11 | 1:1 | Block_1 | 0.87821 | 8.4685 | FALSE | NA | NA | NA | NA | 1.102 | NA | NA | NA
Q7RTS1 | Class A basic helix-loop-helix protein 15 | BHLHA15 | 1:1 | Block_1 | 2.8076 | 10.9359 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96G01 | Protein bicaudal D homolog 1 | BICD1 | 1:1 | Block_1 | 0.73135 | 6.1999 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZP65 | BICD family-like cargo adapter 1 | BICDL1 | 1:1 | Block_1 | 2.5 | 15.6428 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6AI39 | BRD4-interacting chromatin-remodeling complex-associated protein-like | BICRAL | 1:1 | Block_1 | 2.5 | 12.155 | TRUE | NA | NA | NA | NA | 2.6203 | NA | NA | 2.8242
Q13323 | Bcl-2-interacting killer | BIK | 1:1 | Block_1 | 2.8278 | 11.0037 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O00499 | Myc box-dependent-interacting protein 1 | BIN1 | 1:1 | Block_1 | 2.4263 | 8.2411 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q13489 | Baculoviral IAP repeat-containing protein 3 | BIRC3 | 1:1 | Block_1 | 4.1895 | 9.1112 | FALSE | NA | NA | NA | NA | NA | NA | NA | 4.2382
Q96CA5 | Baculoviral IAP repeat-containing protein 7 | BIRC7 | 1:1 | Block_1 | 2.8638 | 9.4406 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P54132 | RecQ-like DNA helicase BLM | BLM | 1:1 | Block_1 | 1.5211 | 9.0257 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H2G9 | Golgin-45 | BLZF1 | 1:1 | Block_1 | 1.5364 | 11.3692 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.8694
P35226 | Polycomb complex protein BMI-1 | BMI1 | 1:1 | Block_1 | 1.1773 | 12.083 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O00238 | Bone morphogenetic protein receptor type-1B | BMPR1B | 1:1 | Block_1 | 1.9627 | 11.8732 | TRUE | NA | NA | NA | NA | 2.3047 | NA | NA | NA
Q8NFC6 | Biorientation of chromosomes in cell division protein 1-like 1 | BOD1L1 | 1:1 | Block_1 | 0.80584 | 9.0406 | FALSE | NA | NA | NA | NA | 0.82056 | NA | NA | NA
Q9BQP9 | BPI fold-containing family A member 3 | BPIFA3 | 1:1 | Block_1 | 2.198 | 10.7203 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P55201 | Peregrin | BRPF1 | 1:1 | Block_1 | 5.3328 | 14.2366 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NY30 | Protein BTG4 | BTG4 | 1:1 | Block_1 | 0.56 | 12.5948 | TRUE | NA | NA | NA | NA | 0.89895 | NA | NA | 1.8341
Q13895 | Bystin | BYSL | 1:1 | Block_1 | 3.6815 | 12.184 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5T681 | Uncharacterized protein C10orf62 | C10orf62 | 1:1 | Block_1 | 3.5939 | 11.8342 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IYJ2 | Uncharacterized protein C10orf67, mitochondrial | C10orf67 | 1:1 | Block_1 | 6.66 | 16.3311 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96M02 | (E2-independent) E3 ubiquitin-conjugating enzyme FATS | C10orf90 | 1:1 | Block_1 | 0.30816 | 11.1276 | TRUE | NA | NA | NA | NA | 0.31603 | NA | NA | NA
Q96LP6 | Uncharacterized protein C12orf42 | C12orf42 | 1:1 | Block_1 | 2.8464 | 12.5715 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N769 | Uncharacterized protein C14orf178 | C14orf178 | 1:1 | Block_1 | 0.80188 | 8.8647 | TRUE | 8.8449 | 7.9385 | NA | NA | 0.94829 | NA | NA | 0.8213
Q5SNV9 | Uncharacterized protein C1orf167 | C1orf167 | 1:1 | Block_1 | 1.7791 | 12.3839 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BZE7 | UPF0193 protein EVG1 | C22orf23 | 1:1 | Block_1 | 7.2607 | 16.5291 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TF44 | C2 calcium-dependent domain-containing protein 4C | C2CD4C | 1:1 | Block_1 | 1.009 | 11.8411 | TRUE | NA | NA | NA | NA | 1.7751 | NA | NA | 1.2724
Q9NWW7 | Uncharacterized protein C2orf42 | C2orf42 | 1:1 | Block_1 | 2.6747 | 12.6188 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q53S99 | Folate transporter-like protein C2orf83 | C2orf83 | 1:1 | Block_1 | 2 | 14.098 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q53FE4 | Uncharacterized protein C4orf17 | C4orf17 | 1:1 | Block_1 | 1.5673 | 11.4856 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5T0Z8 | Uncharacterized protein C6orf132 | C6orf132 | 1:1 | Block_1 | 2.558 | 12.0583 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5VYM1 | Uncharacterized protein C9orf131 | C9orf131 | 1:1 | Block_1 | 4.0235 | 15.1093 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P57796 | Calcium-binding protein 4 | CABP4 | 1:1 | Block_1 | 0.41352 | 9.2786 | FALSE | 13.8087 | 8.5777 | NA | NA | 1.4136 | NA | NA | 0.90792
O00305 | Voltage-dependent L-type calcium channel subunit beta-4 | CACNB4 | 1:1 | Block_1 | 1.9179 | 5.4972 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WUQ7 | Cactin | CACTIN | 1:1 | Block_1 | 1.0084 | 10.512 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TC20 | Cancer-associated gene 1 protein | CAGE1 | 1:1 | Block_1 | 0.85627 | 12.4742 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.4765
Q96NX5 | Calcium/calmodulin-dependent protein kinase type 1G | CAMK1G | 1:1 | Block_1 | 7.0429 | 16.8981 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WWF8 | Calcyphosin-like protein | CAPSL | 1:1 | Block_1 | 4.1673 | 12.3676 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WXD9 | Caskin-1 | CASKIN1 | 1:1 | Block_1 | 1.2129 | 11.6461 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NA61 | Protein chibby homolog 2 | CBY2 | 1:1 | Block_1 | 1.43 | 13.6683 | TRUE | NA | NA | NA | NA | 1.7066 | NA | NA | 1.4422
Q6ZUS5 | Coiled-coil domain-containing protein 121 | CCDC121 | 1:1 | Block_1 | 3.0273 | 13.2921 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P0C7W6 | Coiled-coil domain-containing protein 172 | CCDC172 | 1:1 | Block_1 | 1.8462 | 10.53 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5TID7 | Coiled-coil domain-containing protein 181 | CCDC181 | 1:1 | Block_1 | 1.132 | 9.6051 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5VVM6 | Coiled-coil domain-containing protein 30 | CCDC30 | 1:1 | Block_1 | 2.6355 | 14.5014 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N5R6 | Coiled-coil domain-containing protein 33 | CCDC33 | 1:1 | Block_1 | 1.2652 | 12.3302 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6NI79 | Coiled-coil domain-containing protein 69 | CCDC69 | 1:1 | Block_1 | 1.8867 | 9.6444 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P24385 | G1/S-specific cyclin-D1 | CCND1 | 1:1 | Block_1 | 1.3471 | 13.0616 | FALSE | NA | NA | NA | NA | 2.2708 | NA | NA | 3.3809
Q6UXZ3 | CMRF35-like molecule 5 | CD300LD | 1:1 | Block_1 | 0.76992 | 7.609 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P30307 | M-phase inducer phosphatase 3 | CDC25C | 1:1 | Block_1 | 1.6778 | 13.7905 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q69YH5 | Cell division cycle-associated protein 2 | CDCA2 | 1:1 | Block_1 | 1.1242 | 11.0886 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q99618 | Cell division cycle-associated protein 3 | CDCA3 | 1:1 | Block_1 | -2 | 7.9731 | TRUE | NA | NA | NA | NA | NA | NA | NA | -2
Q53HL2 | Borealin | CDCA8 | 1:1 | Block_1 | 0.87243 | 8.7313 | TRUE | NA | 1.4427 | NA | NA | 1.2516 | NA | NA | NA
O75871 | Carcinoembryonic antigen-related cell adhesion molecule 4 | CEACAM4 | 1:1 | Block_1 | -0.40796 | 5.2696 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N2Z9 | Centromere protein S | CENPS | 1:1 | Block_1 | 0.21873 | 11.1189 | FALSE | 12.2253 | 9.9738 | NA | NA | 0.31371 | NA | NA | 0.91258
Q5JTW2 | Centrosomal protein of 78 kDa | CEP78 | 1:1 | Block_1 | 1.8761 | 11.5682 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y592 | Centrosomal protein of 83 kDa | CEP83 | 1:1 | Block_1 | 1.6585 | 9.2881 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O95813 | Cerberus | CER1 | 1:1 | Block_1 | 1.2599 | 10.5658 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q12798 | Centrin-1 | CETN1 | 1:1 | Block_1 | 2.3382 | 10.1794 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5VTH2 | Protein Flattop | CFAP126 | 1:1 | Block_1 | 3.4867 | 11.3938 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZVT6 | Protein CFAP20DC | CFAP20DC | 1:1 | Block_1 | 2.6422 | 14.0055 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NDM7 | Cilia- and flagella-associated protein 43 | CFAP43 | 1:1 | Block_1 | 3.0233 | 13.9841 | FALSE | NA | 15.8716 | NA | NA | NA | NA | NA | NA
Q9C0B2 | Cilia- and flagella-associated protein 74 | CFAP74 | 1:1 | Block_1 | -0.7 | 8.8034 | TRUE | 15.0344 | 12.8446 | NA | NA | NA | NA | NA | -0.18153
Q9P2B7 | Cilia- and flagella-associated protein 97 | CFAP97 | 1:1 | Block_1 | 1.6385 | 11.7095 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q13112 | Chromatin assembly factor 1 subunit B | CHAF1B | 1:1 | Block_1 | 2.6729 | 14.4907 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96JM3 | Chromosome alignment-maintaining phosphoprotein 1 | CHAMP1 | 1:1 | Block_1 | 1.2317 | 9.5325 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O14647 | Chromodomain-helicase-DNA-binding protein 2 | CHD2 | 1:1 | Block_1 | 2.3084 | 10.2302 | TRUE | NA | NA | NA | NA | 2.3354 | NA | NA | NA
O43916 | Carbohydrate sulfotransferase 1 | CHST1 | 1:1 | Block_1 | 1.6634 | 9.1396 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O76071 | Probable cytosolic iron-sulfur protein assembly protein CIAO1 | CIAO1 | 1:1 | Block_1 | 0.79227 | 5.6207 | TRUE | NA | NA | NA | NA | 0.87546 | NA | NA | NA
Q9C0C6 | CLOCK-interacting pacemaker | CIPC | 1:1 | Block_1 | 3.3855 | 14.4449 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N5K1 | CDGSH iron-sulfur domain-containing protein 2 | CISD2 | 1:1 | Block_1 | 1.2408 | 6.3017 | FALSE | 13.1965 | 12.8435 | NA | NA | NA | NA | NA | NA
Q96F83 | Uncharacterized protein CLBA1 | CLBA1 | 1:1 | Block_1 | 3.9787 | 15.449 | TRUE | NA | NA | NA | NA | 3.9978 | NA | NA | 4.0344
Q9HAE3 | Calaxin | CLXN | 1:1 | Block_1 | 5.821 | 13.0777 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q969H4 | Connector enhancer of kinase suppressor of ras 1 | CNKSR1 | 1:1 | Block_1 | 0.13171 | 10.5712 | FALSE | 12.8746 | 13.3859 | NA | NA | 0.80689 | NA | NA | 0.77951
Q9UKZ1 | CCR4-NOT transcription complex subunit 11 | CNOT11 | 1:1 | Block_1 | 0.90381 | 7.9222 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q86Y22 | Collagen alpha-1(XXIII) chain | COL23A1 | 1:1 | Block_1 | 1.21 | 12.2636 | TRUE | NA | NA | NA | NA | 1.4292 | NA | NA | 1.8812
Q14031 | Collagen alpha-6(IV) chain | COL4A6 | 1:1 | Block_1 | 1.2607 | 8.3221 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y215 | Acetylcholinesterase collagenic tail peptide | COLQ | 1:1 | Block_1 | 0.95343 | 8.7302 | FALSE | 12.5753 | 11.789 | NA | NA | 0.96866 | NA | NA | NA
Q86VX2 | COMM domain-containing protein 7 | COMMD7 | 1:1 | Block_1 | 1.2704 | 9.0172 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P2I0 | Cleavage and polyadenylation specificity factor subunit 2 | CPSF2 | 1:1 | Block_1 | 2.0886 | 9.8178 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.5912
Q9UKF6 | Cleavage and polyadenylation specificity factor subunit 3 | CPSF3 | 1:1 | Block_1 | 1.2424 | 10.6315 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IUH2 | Protein CREG2 | CREG2 | 1:1 | Block_1 | 1.1818 | 5.0111 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q86T23 | Putative ciliary rootlet coiled-coil protein-like 1 protein | CROCCP2 | 1:1 | Block_1 | 2.7467 | 10.016 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IVE0 | Putative ciliary rootlet coiled-coil protein-like 2 protein | CROCCP3 | 1:1 | Block_1 | 1.4 | 13.5339 | TRUE | 7.3195 | 14.4744 | NA | NA | 2.1481 | NA | NA | 1.7269
P02489 | Alpha-crystallin A chain | CRYAA | 1:1 | Block_1 | 2.3423 | 11.0837 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.4571
P78368 | Casein kinase I isoform gamma-2 | CSNK1G2 | 1:1 | Block_1 | 1.9919 | 9.1044 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H114 | Cystatin-like 1 | CSTL1 | 1:1 | Block_1 | 0.4498 | 9.0503 | TRUE | NA | NA | NA | NA | 1.0011 | NA | NA | 1.5475
Q5HYN5 | Cancer/testis antigen family 45 member A1 | CT45A1 | 1:1 | Block_1 | 2.2301 | 13.7114 | TRUE | 11.6144 | 6.2118 | NA | NA | NA | NA | NA | NA
O15194 | CTD small phosphatase-like protein | CTDSPL | 1:1 | Block_1 | 1.6239 | 9.1881 | FALSE | NA | NA | NA | NA | 1.7922 | NA | NA | 2.4049
Q9UI47 | Catenin alpha-3 | CTNNA3 | 1:1 | Block_1 | 2.3718 | 10.2128 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.1389
O60716 | Catenin delta-1 | CTNND1 | 1:1 | Block_1 | 2.1198 | 11.5557 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.8318
Q5NE16 | Putative inactive cathepsin L-like protein CTSL3P | CTSL3P | 1:1 | Block_1 | 1.8223 | 10.7918 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P2B4 | CTTNBP2 N-terminal-like protein | CTTNBP2NL | 1:1 | Block_1 | 0.16415 | 11.3075 | TRUE | 2.4513 | NA | NA | NA | NA | NA | NA | 0.90169
Q14999 | Cullin-7 | CUL7 | 1:1 | Block_1 | 2.1061 | 8.5838 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q69YN2 | CWF19-like protein 1 | CWF19L1 | 1:1 | Block_1 | 2.2487 | 14.5569 | TRUE | 13.5428 | 27.0261 | NA | NA | NA | NA | NA | 2.3954
P0DPH9 | Uncharacterized protein CXorf51B | CXorf51B | 1:1 | Block_1 | 2.584 | 12.8119 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q14093 | Cylicin-2 | CYLC2 | 1:1 | Block_1 | 1.8808 | 11.2781 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
A8MQ03 | Cysteine-rich tail protein 1 | CYSRT1 | 1:1 | Block_1 | 6.1894 | 15.2639 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O75553 | Disabled homolog 1 | DAB1 | 1:1 | Block_1 | 1.192 | 11.2971 | FALSE | 13.061 | 14.0789 | NA | NA | NA | NA | NA | NA
Q5VWQ8 | Disabled homolog 2-interacting protein | DAB2IP | 1:1 | Block_1 | 1.3666 | 10.6959 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.5719
Q5T6F0 | DDB1- and CUL4-associated factor 12 | DCAF12 | 1:1 | Block_1 | 2.5907 | 12.4737 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
A2VCK2 | Doublecortin domain-containing protein 2B | DCDC2B | 1:1 | Block_1 | 0.68645 | 12.4146 | TRUE | 14.1965 | 14.3745 | NA | NA | 1.0408 | NA | NA | 1.9224
Q9C098 | Serine/threonine-protein kinase DCLK3 | DCLK3 | 1:1 | Block_1 | 3.2498 | 14.4931 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WTU0 | Protein DDI1 homolog 1 | DDI1 | 1:1 | Block_1 | 0.60977 | 9.9551 | TRUE | NA | 5.0663 | NA | NA | 0.87176 | NA | NA | 0.76209
P35638 | DNA damage-inducible transcript 3 protein | DDIT3 | 1:1 | Block_1 | 2.2489 | 10.7443 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZN54 | Differentially expressed in FDCP 8 homolog | DEF8 | 1:1 | Block_1 | 1.6344 | 12.2873 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.4236
Q5TB30 | DEP domain-containing protein 1A | DEPDC1 | 1:1 | Block_1 | 0.14 | 10.1884 | FALSE | NA | 28.1285 | NA | NA | 0.65682 | NA | NA | 0.46456
Q05D60 | Deuterosome assembly protein 1 | DEUP1 | 1:1 | Block_1 | 1.8088 | 10.4411 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y2E4 | Disco-interacting protein 2 homolog C | DIP2C | 1:1 | Block_1 | 3.21 | 13.3342 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IYB7 | DIS3-like exonuclease 2 | DIS3L2 | 1:1 | Block_1 | 1.5 | 7.7655 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q155Q3 | Dixin | DIXDC1 | 1:1 | Block_1 | 6.0131 | 14.9192 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6UYE1 | Leukemia-associated protein 7 | DLEU7 | 1:1 | Block_1 | 2.8283 | 13.6408 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O60479 | Homeobox protein DLX-3 | DLX3 | 1:1 | Block_1 | 2.1255 | 12.1047 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5HYR2 | Doublesex- and mab-3-related transcription factor C1 | DMRTC1 | 1:1 | Block_1 | 0.78026 | 12.5444 | FALSE | 0.90108 | 1.1505 | NA | NA | NA | NA | NA | NA
Q9NVR5 | Protein kintoun | DNAAF2 | 1:1 | Block_1 | 6.5492 | 13.2975 | TRUE | 17.1489 | 27.1596 | NA | NA | NA | NA | NA | NA
Q8WXU2 | Dynein axonemal assembly factor 4 | DNAAF4 | 1:1 | Block_1 | 0.41214 | 11.3064 | FALSE | 4.8496 | 6.9561 | NA | NA | 0.88067 | NA | NA | 0.67672
Q9NQM4 | Dynein axonemal assembly factor 6 | DNAAF6 | 1:1 | Block_1 | 3.2244 | 13.3057 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96DT5 | Dynein axonemal heavy chain 11 | DNAH11 | 1:1 | Block_1 | 0.44682 | 9.7717 | TRUE | NA | NA | NA | NA | 0.46261 | NA | NA | 0.76671
Q8WXX0 | Dynein axonemal heavy chain 7 | DNAH7 | 1:1 | Block_1 | 0.69242 | 16.3251 | TRUE | 0.67203 | 2.9722 | NA | NA | 1.509 | NA | NA | NA
Q9NYC9 | Dynein axonemal heavy chain 9 | DNAH9 | 1:1 | Block_1 | 1.5 | 12.8755 | TRUE | 1.7578 | NA | NA | NA | NA | NA | NA | 1.8632
Q9NZQ0 | DnaJ homolog subfamily C member 27 | DNAJC27 | 1:1 | Block_1 | 2.3883 | 11.811 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O14645 | Axonemal dynein light intermediate polypeptide 1 | DNALI1 | 1:1 | Block_1 | -0.1259 | 11.6937 | TRUE | NA | NA | NA | NA | NA | NA | NA | -0.12545
Q9BU89 | Deoxyhypusine hydroxylase | DOHH | 1:1 | Block_1 | 0.84471 | 10.8163 | FALSE | 46.7057 | 4.6863 | NA | NA | 1.652 | NA | 1.4356 | 2.6675
Q18PE1 | Protein Dok-7 | DOK7 | 1:1 | Block_1 | 1.6849 | 10.8286 | TRUE | 6.2127 | 13.6859 | NA | NA | NA | NA | NA | NA
Q9Y3R5 | Protein dopey-2 | DOP1B | 1:1 | Block_1 | 1.14 | 12.5638 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.2515
Q7L190 | Developmental pluripotency-associated protein 4 | DPPA4 | 1:1 | Block_1 | 2.1149 | 10.1212 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96MC2 | Dynein regulatory complex protein 1 | DRC1 | 1:1 | Block_1 | 5.8474 | 15.6221 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H069 | Dynein regulatory complex subunit 3 | DRC3 | 1:1 | Block_1 | 1.2128 | 11.299 | TRUE | 0.43191 | 15.0719 | NA | NA | 1.2287 | NA | NA | NA
A6NNA5 | Dorsal root ganglia homeobox protein | DRGX | 1:1 | Block_1 | 2.1149 | 13.895 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZMT9 | Death domain-containing protein 1 | DTHD1 | 1:1 | Block_1 | 1.8549 | 11.9054 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6P1R4 | tRNA-dihydrouridine(16/17) synthase [NAD(P)(+)]-like | DUS1L | 1:1 | Block_1 | 4.7055 | 13.5424 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NX74 | tRNA-dihydrouridine(20) synthase [NAD(P)+]-like | DUS2 | 1:1 | Block_1 | 1.0361 | 12.2946 | TRUE | NA | 29.4732 | NA | NA | NA | NA | NA | NA
Q96G46 | tRNA-dihydrouridine(47) synthase [NAD(P)(+)]-like | DUS3L | 1:1 | Block_1 | 0.58355 | 7.0967 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NR20 | Dual specificity tyrosine-phosphorylation-regulated kinase 4 | DYRK4 | 1:1 | Block_1 | 0.35911 | 11.9461 | TRUE | NA | NA | NA | NA | 1.1367 | NA | NA | 0.87796
Q9NVP4 | Double zinc ribbon and ankyrin repeat-containing protein 1 | DZANK1 | 1:1 | Block_1 | 1.4264 | 12.8246 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.4345
O95905 | Protein ecdysoneless homolog | ECD | 1:1 | Block_1 | 6.7699 | 14.3629 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5VYK3 | Proteasome adapter and scaffold protein ECM29 | ECPAS | 1:1 | Block_1 | 2.3123 | 10.1928 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WWZ3 | Ectodysplasin-A receptor-associated adapter protein | EDARADD | 1:1 | Block_1 | 4.1242 | 12.7166 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IY85 | EF-hand calcium-binding domain-containing protein 13 | EFCAB13 | 1:1 | Block_1 | 4.2392 | 11.1105 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H9B1 | Histone-lysine N-methyltransferase EHMT1 | EHMT1 | 1:1 | Block_1 | 0.80873 | 11.4475 | FALSE | NA | NA | NA | NA | 1.9506 | NA | NA | NA
Q9Y6B2 | EP300-interacting inhibitor of differentiation 1 | EID1 | 1:1 | Block_1 | 2.0812 | 11.8995 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BQI3 | Eukaryotic translation initiation factor 2-alpha kinase 1 | EIF2AK1 | 1:1 | Block_1 | 1.2376 | 12.9759 | FALSE | NA | NA | NA | NA | 1.7417 | NA | NA | 1.6798
Q13144 | Translation initiation factor eIF-2B subunit epsilon | EIF2B5 | 1:1 | Block_1 | 4.7903 | 13.7662 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q7L2H7 | Eukaryotic translation initiation factor 3 subunit M | EIF3M | 1:1 | Block_1 | 5.6067 | 17.3655 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6IS14 | Eukaryotic translation initiation factor 5A-1-like | EIF5AL1 | 1:1 | Block_1 | -0.78757 | 11.0147 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UKW6 | ETS-related transcription factor Elf-5 | ELF5 | 1:1 | Block_1 | 0.7432 | 11.0225 | FALSE | NA | NA | NA | NA | 0.85521 | NA | NA | NA
Q9NPA0 | ER membrane protein complex subunit 7 | EMC7 | 1:1 | Block_1 | 3.2763 | 8.4202 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96AY2 | Crossover junction endonuclease EME1 | EME1 | 1:1 | Block_1 | 1.8947 | 13.1926 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O00423 | Echinoderm microtubule-associated protein-like 1 | EML1 | 1:1 | Block_1 | -0.26597 | 11.8402 | FALSE | 13.7252 | 8.8773 | -0.24751 | 0.65628 | 1.5308 | 0.25768 | 1.0597 | 2.3705
P19622 | Homeobox protein engrailed-2 | EN2 | 1:1 | Block_1 | 1.6507 | 7.3452 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O94919 | Endonuclease domain-containing 1 protein | ENDOD1 | 1:1 | Block_1 | 2 | 11.4146 | FALSE | 22.2248 | 18.1294 | NA | NA | 3.2036 | NA | NA | 4.0538
Q96L91 | E1A-binding protein p400 | EP400 | 1:1 | Block_1 | 1.6472 | 12.11 | TRUE | NA | 7.092 | NA | NA | NA | NA | NA | NA
Q6ZTU2 | Putative EP400-like protein | EP400P1 | 1:1 | Block_1 | 2 | 13.1127 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q99814 | Endothelial PAS domain-containing protein 1 | EPAS1 | 1:1 | Block_1 | 2 | 14.9556 | TRUE | 36.9414 | 11.1584 | NA | NA | 2.0945 | NA | NA | NA
P29320 | Ephrin type-A receptor 3 | EPHA3 | 1:1 | Block_1 | 1 | 14.3845 | FALSE | 28.5695 | 4.4074 | NA | NA | 1.8453 | NA | NA | 1.7634
Q2NKX8 | DNA excision repair protein ERCC-6-like | ERCC6L | 1:1 | Block_1 | 2.379 | 11.5922 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P11308 | Transcriptional regulator ERG | ERG | 1:1 | Block_1 | 2.0415 | 4.2221 | FALSE | NA | NA | NA | NA | NA | NA | NA | 4.7503
A8K979 | ERI1 exoribonuclease 2 | ERI2 | 1:1 | Block_1 | 1.2947 | 12.1619 | FALSE | 4.0712 | 12.6083 | NA | NA | NA | NA | NA | 1.6192
P43268 | ETS translocation variant 4 | ETV4 | 1:1 | Block_1 | 1.0748 | 9.1291 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UI08 | Ena/VASP-like protein | EVL | 1:1 | Block_1 | 0.086801 | 7.1707 | FALSE | 24.3978 | 22.7928 | NA | 0.91179 | 1.851 | NA | 1.1617 | 2.9306
A8MZ36 | Envoplakin-like protein | EVPLL | 1:1 | Block_1 | -0.03751 | 14.209 | TRUE | NA | 20.7791 | NA | NA | 0.018468 | NA | NA | 0.44926
Q03828 | Homeobox even-skipped homolog protein 2 | EVX2 | 1:1 | Block_1 | 2.3296 | 12.9925 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NQT4 | Exosome complex component RRP46 | EXOSC5 | 1:1 | Block_1 | 1.8031 | 9.883 | TRUE | 16.7751 | 19.5384 | NA | NA | NA | NA | NA | 2.6624
Q8NEV8 | Exophilin-5 | EXPH5 | 1:1 | Block_1 | 1.3312 | 12.7108 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P0DMU3 | FAM231A/C-like protein LOC102723383 | F231L | 1:1 | Block_1 | 3.4091 | 13.5247 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6SJ93 | Serine protease FAM111B | FAM111B | 1:1 | Block_1 | 2.1798 | 12.9512 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NX05 | Constitutive coactivator of PPAR-gamma-like protein 2 | FAM120C | 1:1 | Block_1 | -3.3447 | 6.2254 | TRUE | NA | NA | NA | NA | -3.0535 | NA | NA | -1.7802
Q8NE31 | Protein FAM13C | FAM13C | 1:1 | Block_1 | 3.1475 | 11.7017 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UHL3 | Protein FAM153A | FAM153A | 1:1 | Block_1 | 2.5776 | 12.8407 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6J272 | Protein FAM166A | FAM166A | 1:1 | Block_1 | 1.2381 | 12.7812 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96KS9 | Protein FAM167A | FAM167A | 1:1 | Block_1 | 0.26754 | 12.2741 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
A6NMN3 | Protein FAM170B | FAM170B | 1:1 | Block_1 | 1.5109 | 10.8988 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6PVY3 | Protein FAM177B | FAM177B | 1:1 | Block_1 | 0.001 | 11.0369 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9ULE4 | Protein FAM184B | FAM184B | 1:1 | Block_1 | 2.0293 | 14.1869 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IYM0 | Protein FAM186B | FAM186B | 1:1 | Block_1 | 1.5 | 9.7311 | TRUE | 11.8696 | NA | NA | NA | 1.681 | NA | NA | 1.55
Q96PV7 | Protein FAM193B | FAM193B | 1:1 | Block_1 | -2.3436 | 9.0583 | TRUE | NA | NA | NA | NA | NA | NA | NA | -1.3332
Q8WU58 | Protein FAM222B | FAM222B | 1:1 | Block_1 | 0.70107 | 12.2842 | TRUE | 12.0809 | 16.677 | NA | NA | 2.1149 | NA | 1.8576 | 3.3036
F5H4B4 | Protein FAM227A | FAM227A | 1:1 | Block_1 | 4.6063 | 13.0143 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q86W67 | Protein FAM228A | FAM228A | 1:1 | Block_1 | 1.1004 | 13.2165 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.5154
Q96LP2 | Protein FAM81B | FAM81B | 1:1 | Block_1 | 2.9714 | 14.2944 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BQN1 | Protein FAM83C | FAM83C | 1:1 | Block_1 | 2.5 | 17.4375 | FALSE | NA | 37.7105 | NA | NA | 3.3332 | NA | NA | 4.3112
O94887 | FERM, ARHGEF and pleckstrin domain-containing protein 2 | FARP2 | 1:1 | Block_1 | 1.3223 | 10.1847 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6PCT2 | F-box/LRR-repeat protein 19 | FBXL19 | 1:1 | Block_1 | 2.158 | 9.8516 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UJT9 | F-box/LRR-repeat protein 7 | FBXL7 | 1:1 | Block_1 | 2 | 10.3824 | TRUE | NA | NA | NA | NA | 2.0602 | NA | NA | NA
Q5XUX0 | F-box only protein 31 | FBXO31 | 1:1 | Block_1 | 3.1532 | 13.2346 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NWN3 | F-box only protein 34 | FBXO34 | 1:1 | Block_1 | 2 | 12.5172 | TRUE | 40.4184 | NA | NA | NA | NA | NA | NA | 2.8875
Q9UH90 | F-box only protein 40 | FBXO40 | 1:1 | Block_1 | 1.8419 | 11.9555 | FALSE | NA | NA | NA | NA | 2.2066 | NA | NA | 3.1528
Q6P3S6 | F-box only protein 42 | FBXO42 | 1:1 | Block_1 | 2.5 | 8.5149 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O14526 | F-BAR domain only protein 1 | FCHO1 | 1:1 | Block_1 | 0.80938 | 11.002 | FALSE | 7.4831 | 6.585 | NA | NA | 1.2839 | NA | NA | 6.8636
O94868 | F-BAR and double SH3 domains protein 2 | FCHSD2 | 1:1 | Block_1 | 0.22546 | 9.1728 | FALSE | 4.03 | 5.5094 | NA | NA | NA | NA | NA | 1.1892
P22570 | NADPH:adrenodoxin oxidoreductase, mitochondrial | FDXR | 1:1 | Block_1 | 1.1132 | 8.4225 | FALSE | 7.2911 | 15.1548 | NA | NA | 1.443 | NA | NA | 1.3191
P05230 | Fibroblast growth factor 1 | FGF1 | 1:1 | Block_1 | 2 | 14.5047 | TRUE | NA | 59.6213 | NA | NA | 2.6658 | NA | NA | 2.9564
B1AJZ9 | Forkhead-associated domain-containing protein 1 | FHAD1 | 1:1 | Block_1 | 1.8543 | 11.6767 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q2V2M9 | FH1/FH2 domain-containing protein 3 | FHOD3 | 1:1 | Block_1 | 0.66261 | 12.0485 | TRUE | NA | NA | NA | NA | 2.5502 | NA | NA | 2.1868
Q7Z7B0 | Filamin-A-interacting protein 1 | FILIP1 | 1:1 | Block_1 | 1.0361 | 11.7623 | TRUE | 11.2127 | 4.6012 | NA | NA | NA | NA | NA | 1.0413
Q96Q35 | Flagellum-associated coiled-coil domain-containing protein 1 | FLACC1 | 1:1 | Block_1 | 0.74013 | 8.2222 | TRUE | NA | 2.6295 | NA | NA | NA | NA | NA | 0.74018
Q8NFG4 | Folliculin | FLCN | 1:1 | Block_1 | 1 | 11.9691 | FALSE | NA | NA | NA | NA | 1.0193 | NA | NA | NA
P55316 | Forkhead box protein G1 | FOXG1 | 1:1 | Block_1 | 1.2718 | 9.6138 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q12952 | Forkhead box protein L1 | FOXL1 | 1:1 | Block_1 | 0.41916 | 10.7372 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O15353 | Forkhead box protein N1 | FOXN1 | 1:1 | Block_1 | 0.001 | 11.8175 | TRUE | 20.9744 | 24.5293 | NA | NA | 0.35064 | NA | NA | 0.69218
Q8IVH2 | Forkhead box protein P4 | FOXP4 | 1:1 | Block_1 | 0.094413 | 8.7467 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.31837
Q96NE9 | FERM domain-containing protein 6 | FRMD6 | 1:1 | Block_1 | 3.0807 | 15.2839 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BZ67 | FERM domain-containing protein 8 | FRMD8 | 1:1 | Block_1 | 0.91072 | 12.625 | TRUE | 13.8958 | 7.5993 | NA | 1.5201 | 3.9262 | 1.1923 | 3.9148 | 5.7149
Q5SYB0 | FERM and PDZ domain-containing protein 1 | FRMPD1 | 1:1 | Block_1 | 2.0067 | 11.6353 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6IN97 | Putative protein FRMPD2-like | FRMPD2B | 1:1 | Block_1 | 2.5 | 16.1531 | TRUE | NA | NA | NA | NA | 3.9967 | NA | NA | 2.7917
Q5CZC0 | Fibrous sheath-interacting protein 2 | FSIP2 | 1:1 | Block_1 | 1.1458 | 9.6438 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96I24 | Far upstream element-binding protein 3 | FUBP3 | 1:1 | Block_1 | 1.7823 | 11.0823 | FALSE | NA | NA | NA | NA | 2.0611 | NA | NA | 2.3588
O00144 | Frizzled-9 | FZD9 | 1:1 | Block_1 | -0.50895 | 7.4898 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q2WGN9 | GRB2-associated-binding protein 4 | GAB4 | 1:1 | Block_1 | 0.31656 | 11.0487 | TRUE | NA | NA | NA | NA | 0.44174 | NA | NA | NA
O14976 | Cyclin-G-associated kinase | GAK | 1:1 | Block_1 | 6.5998 | 13.7453 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TET4 | Neutral alpha-glucosidase C | GANC | 1:1 | Block_1 | 1.4511 | 10.8477 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O95995 | Dynein regulatory complex subunit 4 | GAS8 | 1:1 | Block_1 | -0.61149 | 7.9547 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q92538 | Golgi-specific brefeldin A-resistance guanine nucleotide exchange factor 1 | GBF1 | 1:1 | Block_1 | 3.511 | 12.5181 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N9F7 | Lysophospholipase D GDPD1 | GDPD1 | 1:1 | Block_1 | 2.1848 | 12.7507 | FALSE | 23.1929 | NA | NA | NA | NA | NA | NA | NA
Q9UJY4 | ADP-ribosylation factor-binding protein GGA2 | GGA2 | 1:1 | Block_1 | 3.0665 | 11.4036 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y6H8 | Gap junction alpha-3 protein | GJA3 | 1:1 | Block_1 | 0.66928 | 9.904 | TRUE | NA | NA | NA | NA | 0.69302 | NA | NA | NA
Q5VSY0 | G kinase-anchoring protein 1 | GKAP1 | 1:1 | Block_1 | 0.001 | 12.6749 | TRUE | NA | 1.4241 | NA | NA | 0.27727 | NA | 0.099249 | 2.1221
Q53GS7 | mRNA export factor GLE1 | GLE1 | 1:1 | Block_1 | 0.74683 | 8.8443 | FALSE | NA | NA | NA | NA | 1.0022 | NA | NA | 0.81866
Q9NZD2 | Glycolipid transfer protein | GLTP | 1:1 | Block_1 | 5 | 14.8558 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UKD1 | Glucocorticoid modulatory element-binding protein 2 | GMEB2 | 1:1 | Block_1 | 2.4571 | 10.6079 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6NEF3_A6NEM1_A6NI86 | Putative golgin subfamily A member 6-like protein 3 | GOLGA6L4_GOLGA6L9_GOLGA6L10 | 1:1 | Block_1 | 2.5 | 12.4469 | TRUE | 37.7625 | NA | NA | NA | NA | NA | NA | NA
A6NN73 | Golgin subfamily A member 8C | GOLGA8CP | 1:1 | Block_1 | 6.796 | 15.9068 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O75487 | Glypican-4 | GPC4 | 1:1 | Block_1 | 0.1554 | 9.2514 | FALSE | 30.8914 | 24.2979 | NA | NA | NA | NA | NA | 0.5649
Q8IZ08 | G-protein coupled receptor 135 | GPR135 | 1:1 | Block_1 | 0.001 | 12.4394 | FALSE | NA | NA | NA | NA | NA | NA | NA | 0.20348
Q9Y4H4 | G-protein-signaling modulator 3 | GPSM3 | 1:1 | Block_1 | 2.109 | 8.1253 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P32298 | G protein-coupled receptor kinase 4 | GRK4 | 1:1 | Block_1 | 1 | 12.5565 | FALSE | NA | 2.5707 | NA | NA | 1.1511 | NA | NA | 1.2649
P57764 | Gasdermin-D | GSDMD | 1:1 | Block_1 | 0.64089 | 9.186 | FALSE | 2.8075 | 1.0817 | NA | 0.98327 | 3.4993 | 1.1722 | 3.8073 | 6.4039
Q6UXU4 | Germ cell-specific gene 1-like protein | GSG1L | 1:1 | Block_1 | 3.154 | 5.8887 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q969Y2 | tRNA modification GTPase GTPBP3, mitochondrial | GTPBP3 | 1:1 | Block_1 | 2.337 | 6.241 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H1H1 | Gametocyte-specific factor 1-like | GTSF1L | 1:1 | Block_1 | 2.2983 | 13.3468 | TRUE | 13.8091 | 22.476 | NA | NA | NA | NA | NA | NA
P02724 | Glycophorin-A | GYPA | 1:1 | Block_1 | 3.2405 | 12.0642 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UQL6 | Histone deacetylase 5 | HDAC5 | 1:1 | Block_1 | 1.2382 | 11.8214 | TRUE | NA | NA | NA | NA | 2.9324 | NA | 2.5507 | 5.3398
Q5TGJ6 | Hepatoma-derived growth factor-like protein 1 | HDGFL1 | 1:1 | Block_1 | 1.7768 | 7.4733 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P2P5 | E3 ubiquitin-protein ligase HECW2 | HECW2 | 1:1 | Block_1 | 4.1745 | 10.538 | TRUE | 16.8617 | NA | NA | NA | NA | NA | NA | NA
Q8NG08 | DNA helicase B | HELB | 1:1 | Block_1 | 0.28853 | 10.4343 | FALSE | 12.7128 | 9.306 | NA | NA | 0.53274 | NA | NA | 0.63535
Q9HCC6 | Transcription factor HES-4 | HES4 | 1:1 | Block_1 | 1.7028 | 8.9543 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
C9JL84 | HERV-H LTR-associating protein 1 | HHLA1 | 1:1 | Block_1 | 4.0501 | 13.0772 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H422 | Homeodomain-interacting protein kinase 3 | HIPK3 | 1:1 | Block_1 | 2.5463 | 9.9083 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NE63 | Homeodomain-interacting protein kinase 4 | HIPK4 | 1:1 | Block_1 | 2.102 | 10.1634 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P41235 | Hepatocyte nuclear factor 4-alpha | HNF4A | 1:1 | Block_1 | 1.6839 | 15.2047 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.5549
Q9UJC3 | Protein Hook homolog 1 | HOOK1 | 1:1 | Block_1 | 0.9013 | 12.0128 | TRUE | NA | 0.73102 | NA | NA | 1.0665 | NA | NA | 1.1761
Q9UPZ3 | BLOC-2 complex member HPS5 | HPS5 | 1:1 | Block_1 | 3.0041 | 10.2471 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q53GQ0 | Very-long-chain 3-oxoacyl-CoA reductase | HSD17B12 | 1:1 | Block_1 | 7.1819 | 13.8853 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P37059 | 17-beta-hydroxysteroid dehydrogenase type 2 | HSD17B2 | 1:1 | Block_1 | 3.1285 | 7.0944 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UBD0 | Heat shock transcription factor, X-linked | HSFX1_HSFX2 | 1:1 | Block_1 | 2.283 | 11.9286 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O43301 | Heat shock 70 kDa protein 12A | HSPA12A | 1:1 | Block_1 | 1.4923 | 11.064 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O95757 | Heat shock 70 kDa protein 4L | HSPA4L | 1:1 | Block_1 | -0.37907 | 10.3779 | TRUE | NA | NA | NA | NA | 0.097832 | NA | NA | 0.80689
Q96M11 | Centriolar and ciliogenesis-associated protein HYLS1 | HYLS1 | 1:1 | Block_1 | 1.0163 | 12.2552 | TRUE | NA | NA | NA | NA | 1.4497 | NA | NA | 2.1113
Q02363 | DNA-binding protein inhibitor ID-2 | ID2 | 1:1 | Block_1 | 1.0425 | 4.1772 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BTL4 | Immediate early response gene 2 protein | IER2 | 1:1 | Block_1 | 2.1672 | 7.4813 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P01563 | Interferon alpha-2 | IFNA2 | 1:1 | Block_1 | 7.2557 | 17.7962 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P01574 | Interferon beta | IFNB1 | 1:1 | Block_1 | 2 | 14.0724 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
K9M1U5 | Interferon lambda-4 | IFNL4 | 1:1 | Block_1 | 0.46846 | 10.6753 | FALSE | 0.79219 | 7.5902 | NA | NA | 0.98018 | NA | NA | 1.604
Q9H7X7 | Intraflagellar transport protein 22 homolog | IFT22 | 1:1 | Block_1 | 2.6018 | 10.9735 | TRUE | NA | NA | NA | NA | 2.6131 | NA | NA | NA
Q5DX21 | Immunoglobulin superfamily member 11 | IGSF11 | 1:1 | Block_1 | 2.0794 | 9.4264 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q14626 | Interleukin-11 receptor subunit alpha | IL11RA | 1:1 | Block_1 | 0.53069 | 9.7495 | FALSE | NA | 0.32029 | NA | NA | 1.1504 | NA | NA | 0.81399
Q8NFR9 | Interleukin-17 receptor E | IL17RE | 1:1 | Block_1 | 0.97729 | 9.6933 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NQS7 | Inner centromere protein | INCENP | 1:1 | Block_1 | 1.0498 | 10.5062 | TRUE | NA | 28.6232 | NA | NA | 1.3817 | NA | NA | 1.5624
Q9Y2H2 | Phosphatidylinositide phosphatase SAC2 | INPP5F | 1:1 | Block_1 | 3.5842 | 11.7406 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q2T9L4 | Inhibitory synaptic factor 1 | INSYN1 | 1:1 | Block_1 | -2.3597 | 7.802 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6NMK8 | Protein INSYN2B | INSYN2B | 1:1 | Block_1 | 1.859 | 13.418 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86XH1 | Dynein regulatory complex protein 11 | IQCA1 | 1:1 | Block_1 | 1.272 | 11.0519 | TRUE | NA | NA | NA | NA | 1.6816 | NA | NA | 1.6562
A6NCM1 | IQ and AAA domain-containing protein 1-like | IQCA1L | 1:1 | Block_1 | 2.459 | 14.3564 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NA54 | IQ and ubiquitin-like domain-containing protein | IQUB | 1:1 | Block_1 | 2.2504 | 13.1816 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5VVH5 | Interleukin-1 receptor-associated kinase 1-binding protein 1 | IRAK1BP1 | 1:1 | Block_1 | 1.2583 | 13.4201 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.2799
Q13349 | Integrin alpha-D | ITGAD | 1:1 | Block_1 | 0.98 | 8.451 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q13352 | Centromere protein R | ITGB3BP | 1:1 | Block_1 | 3.9691 | 13.0426 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P23677 | Inositol-trisphosphate 3-kinase A | ITPKA | 1:1 | Block_1 | 1.125 | 9.474 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P28290 | Protein ITPRID2 | ITPRID2 | 1:1 | Block_1 | 1.4886 | 11.892 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.6522
Q6IE81 | Protein Jade-1 | JADE1 | 1:1 | Block_1 | 3.5499 | 13.9061 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5VZ66 | Janus kinase and microtubule-interacting protein 3 | JAKMIP3 | 1:1 | Block_1 | 0.99721 | 13.4842 | TRUE | 6.641 | 3.1528 | NA | NA | NA | NA | NA | NA
Q86VZ6 | Juxtaposed with another zinc finger protein 1 | JAZF1 | 1:1 | Block_1 | -0.53913 | 7.0327 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P266 | Junctional cadherin 5-associated protein | JCAD | 1:1 | Block_1 | 2.1976 | 11.9454 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.625
Q6NYC1 | Bifunctional arginine demethylase and lysyl-hydroxylase JMJD6 | JMJD6 | 1:1 | Block_1 | 1.4881 | 9.8637 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O60229 | Kalirin | KALRN | 1:1 | Block_1 | 2.2149 | 11.2906 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5T7N3 | KN motif and ankyrin repeat domain-containing protein 4 | KANK4 | 1:1 | Block_1 | -0.05579 | 8.2954 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q92831 | Histone acetyltransferase KAT2B | KAT2B | 1:1 | Block_1 | 1.8099 | 12.2355 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y2W7 | Calsenilin | KCNIP3 | 1:1 | Block_1 | 2.8523 | 11.4706 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q14654 | ATP-sensitive inward rectifier potassium channel 11 | KCNJ11 | 1:1 | Block_1 | 4.0987 | 13.5108 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q7Z5Y7 | BTB/POZ domain-containing protein KCTD20 | KCTD20 | 1:1 | Block_1 | 1.9837 | 11.6223 | FALSE | NA | NA | NA | NA | 2.108 | NA | NA | 2.8718
O60341 | Lysine-specific histone demethylase 1A | KDM1A | 1:1 | Block_1 | 3.347 | 11.8169 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P29375 | Lysine-specific demethylase 5A | KDM5A | 1:1 | Block_1 | 1.6664 | 14.2352 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q92628 | Uncharacterized protein KIAA0232 | KIAA0232 | 1:1 | Block_1 | 1.6389 | 11.4193 | TRUE | NA | 9.9083 | NA | NA | NA | NA | NA | 2.4188
Q2KHM9 | Protein moonraker | KIAA0753 | 1:1 | Block_1 | 3.4372 | 13.1854 | TRUE | NA | NA | NA | NA | NA | NA | NA | 4.1643
Q5T5P2 | Sickle tail protein homolog | KIAA1217 | 1:1 | Block_1 | 1.428 | 11.2248 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5VZ46 | Uncharacterized protein KIAA1614 | KIAA1614 | 1:1 | Block_1 | 2.1201 | 14.1995 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N8K9 | Uncharacterized protein KIAA1958 | KIAA1958 | 1:1 | Block_1 | 1.2475 | 12.8315 | TRUE | NA | NA | NA | NA | 1.622 | NA | NA | 2.0787
Q5HYC2 | Uncharacterized protein KIAA2026 | KIAA2026 | 1:1 | Block_1 | 2 | 11.1166 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H1H9 | Kinesin-like protein KIF13A | KIF13A | 1:1 | Block_1 | 0.84835 | 10.8344 | TRUE | 9.7309 | 26.3142 | NA | NA | 1.3676 | NA | NA | 1.5666
Q9NQT8 | Kinesin-like protein KIF13B | KIF13B | 1:1 | Block_1 | 2 | 13.5578 | TRUE | NA | NA | NA | NA | 2.0818 | NA | NA | 2.3386
Q2KJY2 | Kinesin-like protein KIF26B | KIF26B | 1:1 | Block_1 | 1.5 | 12.3722 | TRUE | NA | NA | NA | NA | 1.6132 | NA | NA | NA
B7ZC32 | Kinesin-like protein KIF28P | KIF28P | 1:1 | Block_1 | 0.89729 | 11.1429 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N4N8 | Kinesin-like protein KIF2B | KIF2B | 1:1 | Block_1 | 1.4238 | 11.5146 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZMV9 | Kinesin-like protein KIF6 | KIF6 | 1:1 | Block_1 | 2.5 | 11.8396 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96AC6 | Kinesin-like protein KIFC2 | KIFC2 | 1:1 | Block_1 | 0.9065 | 10.5793 | TRUE | NA | 2.3836 | NA | NA | 1.2147 | NA | NA | 0.9751
O60870 | DNA/RNA-binding protein KIN17 | KIN | 1:1 | Block_1 | 2.1921 | 11.5689 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.794
Q8IZU9 | Kin of IRRE-like protein 3 | KIRREL3 | 1:1 | Block_1 | 1 | 9.6576 | TRUE | NA | 9.996 | NA | NA | 1.2159 | NA | NA | NA
Q9UIH9 | Krueppel-like factor 15 | KLF15 | 1:1 | Block_1 | 1.5 | 11.5885 | TRUE | 13.3852 | 16.3833 | NA | NA | NA | NA | NA | NA
O95600 | Krueppel-like factor 8 | KLF8 | 1:1 | Block_1 | 1.4651 | 8.9811 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NEZ4 | Histone-lysine N-methyltransferase 2C | KMT2C | 1:1 | Block_1 | 0.57477 | 11.2801 | TRUE | 11.9818 | 21.6123 | NA | NA | 0.84895 | NA | NA | 1.1636
Q76NI1 | Kinase non-catalytic C-lobe domain-containing protein 1 | KNDC1 | 1:1 | Block_1 | 1.57 | 11.4056 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y448 | Small kinetochore-associated protein | KNSTRN | 1:1 | Block_1 | 2.2799 | 13.2114 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A9QM74 | Importin subunit alpha-8 | KPNA7 | 1:1 | Block_1 | 2.41 | 10.3748 | TRUE | NA | 46.6076 | NA | NA | 2.4126 | NA | NA | NA
A5PL33 | Protein KRBA1 | KRBA1 | 1:1 | Block_1 | 1.6263 | 13.2647 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N9T8 | Protein KRI1 homolog | KRI1 | 1:1 | Block_1 | 0.993 | 12.2117 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.8492
O00522 | Krev interaction trapped protein 1 | KRIT1 | 1:1 | Block_1 | 0.82117 | 8.3286 | FALSE | 1.2348 | 13.5059 | NA | NA | 1.4282 | NA | NA | 0.8364
Q99456 | Keratin, type I cytoskeletal 12 | KRT12 | 1:1 | Block_1 | 2.5 | 11.3289 | FALSE | 6.8691 | 7.7996 | NA | NA | 3.8483 | NA | NA | 3.8143
P35900 | Keratin, type I cytoskeletal 20 | KRT20 | 1:1 | Block_1 | 1.3276 | 4.1108 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NA19 | Lethal(3)malignant brain tumor-like protein 4 | L3MBTL4 | 1:1 | Block_1 | 6.0118 | 13.3816 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6UWM7 | Lactase-like protein | LCTL | 1:1 | Block_1 | 2.7797 | 8.3738 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O75112 | LIM domain-binding protein 3 | LDB3 | 1:1 | Block_1 | 2.3178 | 11.5322 | FALSE | NA | NA | NA | NA | 2.7676 | NA | NA | 3.5768
P50458 | LIM/homeobox protein Lhx2 | LHX2 | 1:1 | Block_1 | 2.0294 | 10.3712 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UPQ0 | LIM and calponin homology domains-containing protein 1 | LIMCH1 | 1:1 | Block_1 | 1.6523 | 8.7682 | TRUE | NA | NA | NA | NA | 1.744 | NA | NA | 2.6281
Q5TKA1 | Protein lin-9 homolog | LIN9 | 1:1 | Block_1 | 4.6307 | 8.8867 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N7U9 | Putative uncharacterized protein encoded by LINC00469 | LINC00469 | 1:1 | Block_1 | 2.973 | 12.5941 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N535 | Putative uncharacterized protein encoded by LINC00471 | LINC00471 | 1:1 | Block_1 | 0.41201 | 11.0184 | TRUE | 18.9459 | 8.9292 | NA | NA | 1.2837 | NA | NA | 1.0954
Q8N9X3 | Putative uncharacterized protein encoded by LINC01356 | LINC01356 | 1:1 | Block_1 | 0.39149 | 13.35 | TRUE | NA | NA | NA | NA | 1.5112 | NA | NA | 1.2547
P61968 | LIM domain transcription factor LMO4 | LMO4 | 1:1 | Block_1 | 2.5 | 10.0598 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q1L5Z9 | LON peptidase N-terminal domain and RING finger protein 2 | LONRF2 | 1:1 | Block_1 | 2.2143 | 14.8125 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P244 | Leucine-rich repeat and fibronectin type III domain-containing protein 1 | LRFN1 | 1:1 | Block_1 | 1.1733 | 12.5366 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.5296
A6NDA9 | Leucine-rich repeat, immunoglobulin-like domain and transmembrane domain-containing protein 2 | LRIT2 | 1:1 | Block_1 | 0.50531 | 11.6141 | FALSE | NA | NA | NA | NA | 0.78301 | NA | NA | 1.1042
Q9BTT6 | Leucine-rich repeat-containing protein 1 | LRRC1 | 1:1 | Block_1 | 1.0664 | 12.3869 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9C0I9 | Leucine-rich repeat-containing protein 27 | LRRC27 | 1:1 | Block_1 | -0.22159 | 9.271 | TRUE | NA | 3.8961 | NA | NA | NA | NA | NA | NA
Q8N309 | Leucine-rich repeat-containing protein 43 | LRRC43 | 1:1 | Block_1 | 6.25 | 12.4899 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NAA5 | Leucine-rich repeat-containing protein 75A | LRRC75A | 1:1 | Block_1 | 1.2556 | 11.4723 | TRUE | NA | NA | NA | NA | 2.7837 | NA | 1.7139 | 3.3275
Q86UE6 | Leucine-rich repeat transmembrane neuronal protein 1 | LRRTM1 | 1:1 | Block_1 | 1.3968 | 9.6801 | TRUE | 7.8266 | 23.2143 | NA | NA | NA | NA | NA | NA
Q6Q4G3 | Aminopeptidase Q | LVRN | 1:1 | Block_1 | 0.3624 | 11.0458 | TRUE | NA | NA | NA | NA | 1.2575 | NA | NA | 1.7164
Q17RY6 | Lymphocyte antigen 6K | LY6K | 1:1 | Block_1 | 2.441 | 10.1824 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZN28 | Metastasis-associated in colon cancer protein 1 | MACC1 | 1:1 | Block_1 | 0.69496 | 8.9941 | TRUE | NA | NA | NA | NA | 1.0192 | NA | NA | 0.8214
A1Z1Q3 | ADP-ribose glycohydrolase MACROD2 | MACROD2 | 1:1 | Block_1 | 0.75442 | 9.7531 | FALSE | NA | NA | NA | NA | 1.065 | NA | NA | 1.3454
Q8WXG6 | MAP kinase-activating death domain protein | MADD | 1:1 | Block_1 | 1.58 | 9.6103 | TRUE | 29.7108 | 49.3613 | NA | NA | 1.8025 | NA | NA | 2.1491
P43355 | Melanoma-associated antigen 1 | MAGEA1 | 1:1 | Block_1 | 0.29227 | 11.1079 | FALSE | 15.1828 | 8.5028 | NA | NA | 1.2642 | NA | NA | 0.73324
O15480 | Melanoma-associated antigen B3 | MAGEB3 | 1:1 | Block_1 | 2.6892 | 10.9421 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O15481 | Melanoma-associated antigen B4 | MAGEB4 | 1:1 | Block_1 | 1.7721 | 13.0966 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96JG8 | Melanoma-associated antigen D4 | MAGED4 | 1:1 | Block_1 | 2.4267 | 11.675 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P20794 | Serine/threonine-protein kinase MAK | MAK | 1:1 | Block_1 | 0.034082 | 11.2642 | FALSE | 11.3868 | 8.4529 | NA | NA | 1.6593 | NA | NA | 1.4453
O43283 | Mitogen-activated protein kinase kinase kinase 13 | MAP3K13 | 1:1 | Block_1 | 2.4317 | 10.1959 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q99759 | Mitogen-activated protein kinase kinase kinase 3 | MAP3K3 | 1:1 | Block_1 | 2 | 11.4522 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.6276
Q9Y6R4 | Mitogen-activated protein kinase kinase kinase 4 | MAP3K4 | 1:1 | Block_1 | 6.1063 | 16.2149 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q14244 | Ensconsin | MAP7 | 1:1 | Block_1 | -0.36439 | 8.2615 | TRUE | 21.2654 | 7.275 | -0.26949 | 0.57444 | 1.1873 | 0.66722 | 1.4279 | 2.6634
P53778 | Mitogen-activated protein kinase 12 | MAPK12 | 1:1 | Block_1 | 0.001 | 7.7029 | FALSE | NA | NA | NA | NA | 0.015437 | NA | NA | NA
P31152 | Mitogen-activated protein kinase 4 | MAPK4 | 1:1 | Block_1 | 2.0375 | 11.5598 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q16644 | MAP kinase-activated protein kinase 3 | MAPKAPK3 | 1:1 | Block_1 | 2.0174 | 10.8696 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N8E1 | Putative uncharacterized protein encoded by MAPKAPK5-AS1 | MAPKAPK5-AS1 | 1:1 | Block_1 | 1.3172 | 12.133 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O60336 | Mitogen-activated protein kinase-binding protein 1 | MAPKBP1 | 1:1 | Block_1 | 0.79753 | 8.7879 | TRUE | NA | NA | NA | NA | 1.0924 | NA | NA | 1.0721
Q9Y2H9 | Microtubule-associated serine/threonine-protein kinase 1 | MAST1 | 1:1 | Block_1 | 1.6689 | 12.0969 | FALSE | NA | NA | NA | NA | 2.2544 | NA | NA | NA
A4D2B0 | Metallo-beta-lactamase domain-containing protein 1 | MBLAC1 | 1:1 | Block_1 | 3.4762 | 14.3675 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96RQ3 | Methylcrotonoyl-CoA carboxylase subunit alpha, mitochondrial | MCCC1 | 1:1 | Block_1 | 2.3444 | 11.2352 | FALSE | NA | NA | NA | NA | 2.8234 | NA | NA | NA
Q86YR7 | Probable guanine nucleotide exchange factor MCF2L2 | MCF2L2 | 1:1 | Block_1 | 3.5016 | 12.5223 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BTE3 | Mini-chromosome maintenance complex-binding protein | MCMBP | 1:1 | Block_1 | 0.63556 | 8.617 | TRUE | NA | NA | NA | NA | 0.8186 | NA | NA | 0.70366
Q14676 | Mediator of DNA damage checkpoint protein 1 | MDC1 | 1:1 | Block_1 | 1.5897 | 13.7652 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.3526
Q93074 | Mediator of RNA polymerase II transcription subunit 12 | MED12 | 1:1 | Block_1 | 3.076 | 14.3202 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O14770 | Homeobox protein Meis2 | MEIS2 | 1:1 | Block_1 | 1.9232 | 12.9005 | FALSE | 18.586 | 29.5708 | NA | NA | 2.0031 | NA | NA | 2.5743
P0C7V9 | Putative methyltransferase-like protein 15P1 | METTL15P1 | 1:1 | Block_1 | 2 | 10.9196 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86W50 | RNA N6-adenosine-methyltransferase METTL16 | METTL16 | 1:1 | Block_1 | 5 | 15.3983 | TRUE | NA | NA | NA | NA | NA | NA | NA | 5.4832
Q8N6Q8 | Probable methyltransferase-like protein 25 | METTL25 | 1:1 | Block_1 | 1.7891 | 12.8519 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.9528
Q5JXC2 | Migration and invasion-inhibitory protein | MIIP | 1:1 | Block_1 | 0.80778 | 8.5059 | FALSE | NA | NA | NA | NA | 0.82321 | NA | NA | 0.80779
Q4G0A6 | Probable ubiquitin carboxyl-terminal hydrolase MINDY-4 | MINDY4 | 1:1 | Block_1 | 2.2102 | 13.0702 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A8MYZ0 | Inactive ubiquitin carboxyl-terminal hydrolase MINDY-4B | MINDY4B | 1:1 | Block_1 | 1.9038 | 9.0127 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P58340 | Myeloid leukemia factor 1 | MLF1 | 1:1 | Block_1 | 1.5701 | 12.1841 | TRUE | 19.1876 | NA | NA | NA | NA | NA | NA | NA
Q9HAP2 | MLX-interacting protein | MLXIP | 1:1 | Block_1 | 1.6964 | 10.5993 | TRUE | 25.9127 | 11.915 | NA | NA | NA | NA | NA | 2.2716
Q495T6 | Membrane metallo-endopeptidase-like 1 | MMEL1 | 1:1 | Block_1 | 3.093 | 11.1403 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BWT6 | Meiotic nuclear division protein 1 homolog | MND1 | 1:1 | Block_1 | 2.0567 | 12.2389 | FALSE | 11.9553 | 16.1134 | NA | NA | NA | NA | NA | NA
Q99583 | Max-binding protein MNT | MNT | 1:1 | Block_1 | 0.75477 | 12.2139 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.7858
Q9Y6X9 | ATPase MORC2 | MORC2 | 1:1 | Block_1 | 2.8613 | 10.5693 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5VZ52 | MORN repeat-containing protein 5 | MORN5 | 1:1 | Block_1 | 3.7 | 15.3085 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UJG1 | Motile sperm domain-containing protein 1 | MOSPD1 | 1:1 | Block_1 | 3.7867 | 11.4573 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P29372 | DNA-3-methyladenine glycosylase | MPG | 1:1 | Block_1 | -0.68441 | 15.0417 | TRUE | 15.8082 | 4.8919 | NA | NA | 0.27379 | NA | -0.37182 | 1.9191
Q96G30 | Melanocortin-2 receptor accessory protein 2 | MRAP2 | 1:1 | Block_1 | 2.6863 | 11.2863 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N565 | Melanoregulin | MREG | 1:1 | Block_1 | 2.4056 | 16.0574 | FALSE | NA | NA | NA | NA | 2.4231 | NA | NA | NA
Q6IN84 | rRNA methyltransferase 1, mitochondrial | MRM1 | 1:1 | Block_1 | 2.5 | 12.9065 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6NES4 | Maestro heat-like repeat-containing protein family member 2A | MROH2A | 1:1 | Block_1 | 3.0279 | 12.5896 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P35548 | Homeobox protein MSX-2 | MSX2 | 1:1 | Block_1 | 0.0064976 | 10.9375 | FALSE | 1.5896 | 8.3443 | NA | 0.25034 | 1.4375 | NA | 0.16358 | 1.5072
Q15390 | Mitochondrial fission regulator 1 | MTFR1 | 1:1 | Block_1 | 1.7075 | 10.1352 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H4K7 | Mitochondrial ribosome-associated GTPase 2 | MTG2 | 1:1 | Block_1 | 2.056 | 14.8812 | TRUE | NA | NA | NA | NA | 2.5742 | NA | NA | NA
P46199 | Translation initiation factor IF-2, mitochondrial | MTIF2 | 1:1 | Block_1 | 2.164 | 12.4177 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q02505 | Mucin-3A | MUC3A | 1:1 | Block_1 | 2.0726 | 11.3161 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.209
Q8TAX7 | Mucin-7 | MUC7 | 1:1 | Block_1 | 1.8229 | 6.1293 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96NY9 | Crossover junction endonuclease MUS81 | MUS81 | 1:1 | Block_1 | -1.0176 | 10.0721 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TBZ2 | MYCBP-associated protein | MYCBPAP | 1:1 | Block_1 | 1.1893 | 13.5091 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UKX3 | Myosin-13 | MYH13 | 1:1 | Block_1 | 2.259 | 13.2303 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y2K3 | Myosin-15 | MYH15 | 1:1 | Block_1 | 2.7072 | 12.0946 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BUA6 | Myosin regulatory light chain 10 | MYL10 | 1:1 | Block_1 | 3.2878 | 10.3368 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9HD67 | Unconventional myosin-X | MYO10 | 1:1 | Block_1 | 3.5904 | 11.0514 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NQX4 | Unconventional myosin-Vc | MYO5C | 1:1 | Block_1 | 1.1414 | 11.2296 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P15172 | Myoblast determination protein 1 | MYOD1 | 1:1 | Block_1 | 1.3413 | 8.7645 | TRUE | NA | NA | NA | NA | 1.5875 | NA | NA | NA
Q9UL68 | Myelin transcription factor 1-like protein | MYT1L | 1:1 | Block_1 | 3.6271 | 12.6091 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H009 | Nascent polypeptide-associated complex subunit alpha-2 | NACA2 | 1:1 | Block_1 | 1.8884 | 12.5135 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H9S0 | Homeobox protein NANOG | NANOG | 1:1 | Block_1 | 4.1279 | 13.3196 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NFP9 | Neurobeachin | NBEA | 1:1 | Block_1 | 0.49117 | 9.2016 | TRUE | 31.8911 | 2.6772 | NA | 0.55427 | 1.6348 | NA | 0.94361 | 2.2753
O14513 | Nck-associated protein 5 | NCKAP5 | 1:1 | Block_1 | 4.7668 | 11.2671 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H4R4 | Putative nuclear receptor corepressor 1-like protein NCOR1P1 | NCOR1P1 | 1:1 | Block_1 | 3.27 | 9.7359 | TRUE | NA | 12.0968 | NA | NA | NA | NA | NA | 3.7109
Q9UHB4 | NADPH-dependent diflavin oxidoreductase 1 | NDOR1 | 1:1 | Block_1 | 0.81587 | 11.2635 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.3448
Q9Y375 | Complex I intermediate-associated protein 30, mitochondrial | NDUFAF1 | 1:1 | Block_1 | 4.7686 | 13.9883 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P032 | NADH dehydrogenase [ubiquinone] 1 alpha subcomplex assembly factor 4 | NDUFAF4 | 1:1 | Block_1 | 1 | 6.0753 | FALSE | NA | 0.72967 | NA | NA | 1.0841 | NA | NA | 1.0391
O75489 | NADH dehydrogenase [ubiquinone] iron-sulfur protein 3, mitochondrial | NDUFS3 | 1:1 | Block_1 | 1.9005 | 10.3256 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96P71 | N-terminal EF-hand calcium-binding protein 3 | NECAB3 | 1:1 | Block_1 | 2.0381 | 14.9086 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.2416
Q6P3R8 | Serine/threonine-protein kinase Nek5 | NEK5 | 1:1 | Block_1 | 0.59543 | 12.4712 | FALSE | 19.9939 | 26.9462 | NA | NA | 1.025 | NA | NA | 1.2848
Q7Z628 | Neuroepithelial cell-transforming gene 1 protein | NET1 | 1:1 | Block_1 | 1.87 | 9.8088 | TRUE | NA | NA | NA | NA | 2.9652 | NA | NA | NA
Q8NC67 | Neuropilin and tolloid-like protein 2 | NETO2 | 1:1 | Block_1 | 2.6748 | 11.3025 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q14934 | Nuclear factor of activated T-cells, cytoplasmic 4 | NFATC4 | 1:1 | Block_1 | 1.304 | 7.7202 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96IV0 | Peptide-N(4)-(N-acetyl-beta-glucosaminyl)asparagine amidase | NGLY1 | 1:1 | Block_1 | 1.0456 | 11.7433 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.5402
Q5HYW2 | NHS-like protein 2 | NHSL2 | 1:1 | Block_1 | 0.97421 | 10.4151 | TRUE | 12.5703 | 12.141 | NA | NA | 1.8328 | NA | NA | 2.8441
Q96TA1 | Protein Niban 2 | NIBAN2 | 1:1 | Block_1 | -1.0382 | 10.8139 | FALSE | 22.1001 | 22.683 | -0.93664 | 0.25428 | 1.2594 | -0.28871 | 0.46978 | 1.234
Q9NZG7 | Ninjurin-2 | NINJ2 | 1:1 | Block_1 | 2.3351 | 10.7396 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96S42 | Nodal homolog | NODAL | 1:1 | Block_1 | 4.3932 | 15.1645 | FALSE | NA | NA | NA | NA | 4.506 | NA | NA | 5.072
Q96MY1 | Nucleolar protein 4-like | NOL4L | 1:1 | Block_1 | 2.2831 | 11.7876 | TRUE | 3.3634 | NA | NA | NA | 2.862 | NA | NA | 3.0444
Q14207 | Protein NPAT | NPAT | 1:1 | Block_1 | 0.89222 | 5.3605 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q14994 | Nuclear receptor subfamily 1 group I member 3 | NR1I3 | 1:1 | Block_1 | 1.54 | 12.2327 | TRUE | NA | NA | NA | NA | 1.8988 | NA | NA | 1.8719
Q7Z6K4 | Notch-regulated ankyrin repeat-containing protein | NRARP | 1:1 | Block_1 | 1.5075 | 9.7242 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P01111 | GTPase NRas | NRAS | 1:1 | Block_1 | 2.0411 | 2.1679 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q16656 | Nuclear respiratory factor 1 | NRF1 | 1:1 | Block_1 | 3.4263 | 7.5358 | FALSE | NA | NA | NA | NA | NA | NA | NA | 4.4456
Q9NQ35 | Nuclear receptor-interacting protein 3 | NRIP3 | 1:1 | Block_1 | 2.8061 | 13.2174 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q969T7 | 7-methylguanosine phosphate-specific 5'-nucleotidase | NT5C3B | 1:1 | Block_1 | 1.25 | 6.2837 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H093 | NUAK family SNF1-like kinase 2 | NUAK2 | 1:1 | Block_1 | 1.36 | 12.1861 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.6534
Q9NZJ9 | Diphosphoinositol polyphosphate phosphohydrolase 2 | NUDT4 | 1:1 | Block_1 | 1.6462 | 10.2786 | FALSE | 5.4174 | 15.2697 | NA | NA | 1.9262 | NA | NA | 1.7915
Q9Y6R0 | Numb-like protein | NUMBL | 1:1 | Block_1 | 1.2277 | 10.9462 | TRUE | 4.268 | 12.3708 | NA | NA | 1.3745 | NA | NA | 1.6464
Q12769 | Nuclear pore complex protein Nup160 | NUP160 | 1:1 | Block_1 | 0.13545 | 9.839 | FALSE | 3.1996 | 14.6964 | NA | NA | 0.96041 | NA | NA | 0.58085
A6NNL0 | NUT family member 2B | NUTM2B | 1:1 | Block_1 | 4 | 13.7497 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O15381 | Nuclear valosin-containing protein-like | NVL | 1:1 | Block_1 | 1.25 | 9.1055 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H4D5 | Nuclear RNA export factor 3 | NXF3 | 1:1 | Block_1 | 1.9503 | 11.839 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZVC0 | Neuronal tyrosine-phosphorylated phosphoinositide-3-kinase adapter 1 | NYAP1 | 1:1 | Block_1 | 1.122 | 9.1281 | TRUE | 4.88 | 8.9424 | NA | NA | NA | NA | NA | NA
P11926 | Ornithine decarboxylase | ODC1 | 1:1 | Block_1 | 2.7096 | 11.1466 | TRUE | NA | NA | NA | NA | NA | NA | NA | 3.0198
Q9ULJ1 | Protein BCAP | ODF2L | 1:1 | Block_1 | -2.1698 | 8.4527 | TRUE | 14.8694 | 8.8233 | NA | NA | -1.9188 | NA | NA | -1.4798
Q9NRC9 | Otoraplin | OTOR | 1:1 | Block_1 | 1.9734 | 11.2267 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q7L8S5 | OTU domain-containing protein 6A | OTUD6A | 1:1 | Block_1 | 2 | 13.786 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q7Z309 | PABIR family member 2 | PABIR2 | 1:1 | Block_1 | 1.2427 | 12.3835 | TRUE | 8.6463 | 19.5471 | NA | NA | 2.2845 | NA | NA | 2.1557
Q9NP74 | Palmdelphin | PALMD | 1:1 | Block_1 | 1.3908 | 11.6672 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NPB6 | Partitioning defective 6 homolog alpha | PARD6A | 1:1 | Block_1 | 0.72728 | 7.6864 | FALSE | 15.2759 | 37.4364 | NA | NA | NA | NA | NA | NA
Q6UWI2 | Prostate androgen-regulated mucin-like protein 1 | PARM1 | 1:1 | Block_1 | 4.8196 | 8.156 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y6F1 | Protein mono-ADP-ribosyltransferase PARP3 | PARP3 | 1:1 | Block_1 | 1.6077 | 12.8594 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UKK3 | Protein mono-ADP-ribosyltransferase PARP4 | PARP4 | 1:1 | Block_1 | 1.1527 | 10.0908 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IV76 | Circadian clock protein PASD1 | PASD1 | 1:1 | Block_1 | 1.9847 | 13.8906 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96RG2 | PAS domain-containing serine/threonine-protein kinase | PASK | 1:1 | Block_1 | 1.12 | 11.4442 | TRUE | NA | 15.8219 | NA | NA | 1.786 | NA | NA | 1.4824
O43316 | Paired box protein Pax-4 | PAX4 | 1:1 | Block_1 | 1.8316 | 11.3661 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q15004 | PCNA-associated factor | PCLAF | 1:1 | Block_1 | 0.42462 | 9.2525 | TRUE | 6.7445 | 7.8776 | NA | NA | 0.91747 | NA | NA | 1.0802
Q15154 | Pericentriolar material 1 protein | PCM1 | 1:1 | Block_1 | 0.78082 | 9.9846 | TRUE | NA | NA | NA | NA | 1.1943 | NA | 0.78711 | 1.7873
Q9NV79 | Protein-L-isoaspartate O-methyltransferase domain-containing protein 2 | PCMTD2 | 1:1 | Block_1 | 2.2409 | 12.4728 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q08493 | cAMP-specific 3',5'-cyclic phosphodiesterase 4C | PDE4C | 1:1 | Block_1 | 0.54751 | 10.9175 | FALSE | 17.9337 | 25.8108 | NA | NA | 1.3169 | NA | NA | 1.0292
Q5EBL8 | PDZ domain-containing protein 11 | PDZD11 | 1:1 | Block_1 | 2.7724 | 8.7417 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UPQ7 | E3 ubiquitin-protein ligase PDZRN3 | PDZRN3 | 1:1 | Block_1 | -1.1796 | 9.0127 | TRUE | NA | NA | NA | NA | -0.99344 | NA | NA | -0.44549
Q96FA3 | E3 ubiquitin-protein ligase pellino homolog 1 | PELI1 | 1:1 | Block_1 | 1.169 | 6.077 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P40855 | Peroxisomal biogenesis factor 19 | PEX19 | 1:1 | Block_1 | 1 | 4.913 | FALSE | NA | 9.6894 | NA | NA | 1.1897 | NA | NA | 1.028
P16118 | 6-phosphofructo-2-kinase/fructose-2,6-bisphosphatase 1 | PFKFB1 | 1:1 | Block_1 | 7.118 | 13.6503 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q16877 | 6-phosphofructo-2-kinase/fructose-2,6-bisphosphatase 4 | PFKFB4 | 1:1 | Block_1 | 0.31302 | 7.3373 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6NSJ2 | Pleckstrin homology-like domain family B member 3 | PHLDB3 | 1:1 | Block_1 | 0.58946 | 10.762 | FALSE | 16.8816 | 36.6451 | NA | NA | NA | NA | NA | NA
O00329 | Phosphatidylinositol 4,5-bisphosphate 3-kinase catalytic subunit delta isoform | PIK3CD | 1:1 | Block_1 | 1.3849 | 10.7431 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q7Z3Z3 | Piwi-like protein 3 | PIWIL3 | 1:1 | Block_1 | 1.5385 | 14.4925 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P0L9 | Polycystic kidney disease 2-like 1 protein | PKD2L1 | 1:1 | Block_1 | 1.4938 | 12.115 | FALSE | NA | 12.8929 | NA | NA | NA | NA | NA | NA
Q86XP0 | Cytosolic phospholipase A2 delta | PLA2G4D | 1:1 | Block_1 | 1.3143 | 5.8061 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P53816 | Phospholipase A and acyltransferase 3 | PLAAT3 | 1:1 | Block_1 | 0.62109 | 9.706 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O95397 | Putative protein PLEKHA9 | PLEKHA8P1 | 1:1 | Block_1 | 2.5 | 14.8981 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A1L390 | Pleckstrin homology domain-containing family G member 3 | PLEKHG3 | 1:1 | Block_1 | 1.5575 | 11.2865 | FALSE | NA | NA | NA | NA | 1.5593 | NA | NA | 1.981
Q96PX9 | Pleckstrin homology domain-containing family G member 4B | PLEKHG4B | 1:1 | Block_1 | 2.5 | 11.52 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5SXH7 | Pleckstrin homology domain-containing family S member 1 | PLEKHS1 | 1:1 | Block_1 | 2.3396 | 12.5111 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q496M5 | Inactive serine/threonine-protein kinase PLK5 | PLK5 | 1:1 | Block_1 | 1.2112 | 7.563 | FALSE | NA | 16.3417 | NA | NA | NA | NA | NA | NA
P60201 | Myelin proteolipid protein | PLP1 | 1:1 | Block_1 | 1.2715 | 5.985 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TBY8 | Polyamine-modulated factor 1-binding protein 1 | PMFBP1 | 1:1 | Block_1 | 0.85891 | 11.3426 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q13401 | Putative postmeiotic segregation increased 2-like protein 3 | PMS2P3 | 1:1 | Block_1 | 1.2593 | 9.4123 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96PV4 | Paraneoplastic antigen-like protein 5 | PNMA5 | 1:1 | Block_1 | 4.466 | 11.8283 | FALSE | NA | NA | NA | NA | NA | NA | NA | 5.0654
Q8N8W4 | Omega-hydroxyceramide transacylase | PNPLA1 | 1:1 | Block_1 | 3.2346 | 12.4514 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q3SYA9 | Putative POM121-like protein 1 | POM121L1P | 1:1 | Block_1 | 3.2577 | 13.7622 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
H3BUK9 | POTE ankyrin domain family member B2 | POTEB2 | 1:1 | Block_1 | 1.8426 | 12.1695 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q7Z5V6 | Protein phosphatase 1 regulatory subunit 32 | PPP1R32 | 1:1 | Block_1 | 2.9965 | 14.2358 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q08209 | Protein phosphatase 3 catalytic subunit alpha | PPP3CA | 1:1 | Block_1 | 1.0416 | 7.4224 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96LZ3 | Calcineurin subunit B type 2 | PPP3R2 | 1:1 | Block_1 | 1.177 | 8.177 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6NEY8 | Putative prolyl-tRNA synthetase associated domain-containing protein 1 | PRORSD1P | 1:1 | Block_1 | 2.518 | 9.6 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q2NL68 | Proline and serine-rich protein 3 | PROSER3 | 1:1 | Block_1 | 1.7909 | 12.7216 | TRUE | NA | NA | NA | NA | 2.3014 | NA | 2.1188 | 3.7641
Q9UMS4 | Pre-mRNA-processing factor 19 | PRPF19 | 1:1 | Block_1 | 1.1168 | 8.183 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q96HE9 | Proline-rich protein 11 | PRR11 | 1:1 | Block_1 | -0.2977 | 9.4828 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N4B5 | Proline-rich protein 18 | PRR18 | 1:1 | Block_1 | 0.14801 | 13.1912 | TRUE | NA | 39.1071 | NA | NA | 0.89449 | NA | NA | 0.52086
P0DMB1 | Proline-rich protein 23D2 | PRR23D2 | 1:1 | Block_1 | 3.9492 | 13.9338 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96S07 | Proline-rich protein 25 | PRR25 | 1:1 | Block_1 | 1.2139 | 13.2579 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.2902
P79522 | Proline-rich protein 3 | PRR3 | 1:1 | Block_1 | 1.2756 | 9.452 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q7Z6L0 | Proline-rich transmembrane protein 2 | PRRT2 | 1:1 | Block_1 | -0.28216 | 11.5419 | TRUE | 10.3661 | 17.1552 | NA | 0.36041 | 0.93001 | 0.15522 | 0.84112 | 2.1519
P11801 | Serine/threonine-protein kinase H1 | PSKH1 | 1:1 | Block_1 | 1.7248 | 12.4088 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P25787 | Proteasome subunit alpha type-2 | PSMA2 | 1:1 | Block_1 | 1.98 | 11.6224 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
A5LHX3 | Proteasome subunit beta type-11 | PSMB11 | 1:1 | Block_1 | 1.17 | 11.5793 | TRUE | NA | 12.5959 | NA | NA | 2.5088 | NA | NA | 2.7894
Q9P2W1 | Homologous-pairing protein 2 homolog | PSMC3IP | 1:1 | Block_1 | 1.4898 | 11.871 | TRUE | 18.8825 | 11.4886 | NA | NA | 1.9941 | NA | NA | 1.5069
Q96PY0 | Putative uncharacterized protein PSMG3-AS1 | PSMG3-AS1 | 1:1 | Block_1 | 2.5 | 15.0544 | FALSE | 0.26488 | 39.8118 | NA | NA | 3.3638 | NA | NA | 4.5912
Q16825 | Tyrosine-protein phosphatase non-receptor type 21 | PTPN21 | 1:1 | Block_1 | 2.0077 | 13.0721 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NZH4 | Putative pituitary tumor-transforming gene 3 protein | PTTG3P | 1:1 | Block_1 | 3.0451 | 12.8543 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q15032 | R3H domain-containing protein 1 | R3HDM1 | 1:1 | Block_1 | 3.4839 | 12.6773 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A4D1S5 | Ras-related protein Rab-19 | RAB19 | 1:1 | Block_1 | 1.7999 | 13.0557 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P57729 | Ras-related protein Rab-38 | RAB38 | 1:1 | Block_1 | 0.77877 | 8.0171 | TRUE | 5.943 | 8.0724 | NA | NA | 1.4029 | NA | NA | 0.85084
P20337 | Ras-related protein Rab-3B | RAB3B | 1:1 | Block_1 | 3.7089 | 8.3093 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q92696 | Geranylgeranyl transferase type-2 subunit alpha | RABGGTA | 1:1 | Block_1 | 1.989 | 13.3768 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.8238
Q15311 | RalA-binding protein 1 | RALBP1 | 1:1 | Block_1 | 1.6323 | 8.6367 | FALSE | NA | NA | NA | NA | 1.644 | NA | NA | 3.1815
Q2PPJ7 | Ral GTPase-activating protein subunit alpha-2 | RALGAPA2 | 1:1 | Block_1 | 1.5667 | 11.0919 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.7745
Q86SE5 | RNA-binding Raly-like protein | RALYL | 1:1 | Block_1 | 2.8191 | 13.8232 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O60896 | Receptor activity-modifying protein 3 | RAMP3 | 1:1 | Block_1 | 1 | 7.7959 | TRUE | 8.5124 | 15.3351 | NA | NA | 1.5135 | NA | NA | 2.0053
Q86VV4 | Ran-binding protein 3-like | RANBP3L | 1:1 | Block_1 | 0.73058 | 11.1737 | FALSE | NA | NA | NA | NA | 1.2364 | NA | NA | 1.1558
O60518 | Ran-binding protein 6 | RANBP6 | 1:1 | Block_1 | 1.34 | 11.3442 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IZ41 | Ras and EF-hand domain-containing protein | RASEF | 1:1 | Block_1 | 2.747 | 14.1822 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O75901 | Ras association domain-containing protein 9 | RASSF9 | 1:1 | Block_1 | 5 | 16.7403 | TRUE | NA | NA | NA | NA | NA | NA | NA | 5.1117
Q99708 | DNA endonuclease RBBP8 | RBBP8 | 1:1 | Block_1 | 0.8425 | 9.6856 | FALSE | NA | 9.1204 | NA | NA | 1.0468 | NA | NA | 1.006
O75884 | Serine hydrolase RBBP9 | RBBP9 | 1:1 | Block_1 | 3.3857 | 8.0548 | TRUE | NA | NA | NA | NA | NA | NA | NA | 4.223
Q8N0V3 | Putative ribosome-binding factor A, mitochondrial | RBFA | 1:1 | Block_1 | 3.6677 | 11.0005 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5RL73 | RNA-binding protein 48 | RBM48 | 1:1 | Block_1 | 0.53103 | 8.984 | FALSE | NA | NA | NA | NA | 0.62562 | NA | NA | NA
O94762 | ATP-dependent DNA helicase Q5 | RECQL5 | 1:1 | Block_1 | 2.0227 | 12.9415 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.3202
Q96A58 | Ras-related and estrogen-regulated growth inhibitor | RERG | 1:1 | Block_1 | 2.1543 | 10.3253 | TRUE | 7.4499 | 2.7128 | NA | NA | NA | NA | NA | NA
Q6ZTI6 | Refilin-A | RFLNA | 1:1 | Block_1 | 3.7504 | 12.0241 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O14924 | Regulator of G-protein signaling 12 | RGS12 | 1:1 | Block_1 | 2.1539 | 12.133 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NE09 | Regulator of G-protein signaling 22 | RGS22 | 1:1 | Block_1 | 1.8112 | 12.6443 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O15539 | Regulator of G-protein signaling 5 | RGS5 | 1:1 | Block_1 | 4.2263 | 14.6791 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6MZT1 | Regulator of G-protein signaling 7-binding protein | RGS7BP | 1:1 | Block_1 | 0.23402 | 11.7068 | TRUE | NA | NA | NA | NA | 0.75072 | NA | NA | 1.4726
Q8TCX5 | Rhophilin-1 | RHPN1 | 1:1 | Block_1 | 2.0972 | 13.2526 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NVN3 | Synembryn-B | RIC8B | 1:1 | Block_1 | 5.6188 | 14.5342 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
A6NJZ7_Q9UFD9 | RIMS-binding protein 3A/C | RIMBP3C_RIMBP3 | 1:1 | Block_1 | 0.83458 | 11.3667 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q13671 | Ras and Rab interactor 1 | RIN1 | 1:1 | Block_1 | 2 | 8.8004 | TRUE | NA | NA | NA | NA | 2.2918 | NA | NA | NA
Q9Y572 | Receptor-interacting serine/threonine-protein kinase 3 | RIPK3 | 1:1 | Block_1 | 1.037 | 5.168 | FALSE | 5.9918 | 10.4479 | NA | NA | 2.4794 | NA | 1.9915 | 4.4315
Q8IUD6 | E3 ubiquitin-protein ligase RNF135 | RNF135 | 1:1 | Block_1 | 1.7426 | 11.6949 | TRUE | NA | NA | NA | NA | 1.8129 | NA | 2.131 | 3.5817
Q9BXT8 | RING finger protein 17 | RNF17 | 1:1 | Block_1 | 1.4874 | 10.6568 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H0X6 | RING finger protein 208 | RNF208 | 1:1 | Block_1 | 2.4087 | 12.4251 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A8MTL3 | RING finger protein 212B | RNF212B | 1:1 | Block_1 | 2.3722 | 13.0374 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H0A6 | RING finger protein 32 | RNF32 | 1:1 | Block_1 | 1.5758 | 12.0718 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q969K3 | E3 ubiquitin-protein ligase RNF34 | RNF34 | 1:1 | Block_1 | 3.441 | 9.6247 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H0F5 | E3 ubiquitin-protein ligase RNF38 | RNF38 | 1:1 | Block_1 | 2.6232 | 11.8217 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q01974 | Tyrosine-protein kinase transmembrane receptor ROR2 | ROR2 | 1:1 | Block_1 | 5.8468 | 9.7588 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P35398 | Nuclear receptor ROR-alpha | RORA | 1:1 | Block_1 | 3.8236 | 12.2089 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86UA6 | RPA-interacting protein | RPAIN | 1:1 | Block_1 | 2.7615 | 12.4818 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IXW5 | Putative RNA polymerase II subunit B1 CTD phosphatase RPAP2 | RPAP2 | 1:1 | Block_1 | 2.0468 | 10.4281 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P46777 | 60S ribosomal protein L5 | RPL5 | 1:1 | Block_1 | 0.42633 | 5.6816 | FALSE | 12.2481 | 28.502 | NA | NA | NA | NA | NA | 0.46887
P05386 | 60S acidic ribosomal protein P1 | RPLP1 | 1:1 | Block_1 | 0.82041 | 10.4727 | TRUE | NA | NA | NA | NA | 1.2246 | NA | NA | 1.1325
Q9NSQ0 | Putative ribosomal RNA-processing protein 7 homolog B | RRP7BP | 1:1 | Block_1 | 1.9179 | 14.0274 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.0442
Q9NZ71 | Regulator of telomere elongation helicase 1 | RTEL1 | 1:1 | Block_1 | 2.8427 | 15.0131 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NXZ1 | Sarcoma antigen 1 | SAGE1 | 1:1 | Block_1 | 1 | 4.3272 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P1V8 | Sterile alpha motif domain-containing protein 15 | SAMD15 | 1:1 | Block_1 | 2.8859 | 12.9274 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q99590 | Protein SCAF11 | SCAF11 | 1:1 | Block_1 | 1.7644 | 8.8884 | TRUE | NA | 12.6224 | NA | NA | NA | NA | NA | NA
O95171 | Sciellin | SCEL | 1:1 | Block_1 | 1.2344 | 8.7516 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P35499 | Sodium channel protein type 4 subunit alpha | SCN4A | 1:1 | Block_1 | 2.2173 | 10.0564 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
O75056 | Syndecan-3 | SDC3 | 1:1 | Block_1 | 0.36 | 7.4615 | TRUE | 17.9978 | 29.0625 | NA | NA | 0.98532 | NA | 0.79389 | 1.7781
Q93073 | Selenocysteine insertion sequence-binding protein 2-like | SECISBP2L | 1:1 | Block_1 | 3.0906 | 7.9838 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q14563 | Semaphorin-3A | SEMA3A | 1:1 | Block_1 | 3.5035 | 13.3544 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q13591 | Semaphorin-5A | SEMA5A | 1:1 | Block_1 | 2.1531 | 13.0008 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NFY4 | Semaphorin-6D | SEMA6D | 1:1 | Block_1 | 1.1683 | 9.8865 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.4097
Q14140 | SERTA domain-containing protein 2 | SERTAD2 | 1:1 | Block_1 | 3.1289 | 13.3006 | TRUE | 5.2332 | NA | NA | NA | NA | NA | NA | NA
Q8WTS6 | Histone-lysine N-methyltransferase SETD7 | SETD7 | 1:1 | Block_1 | 2.6352 | 12.6425 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q15393 | Splicing factor 3B subunit 3 | SF3B3 | 1:1 | Block_1 | 5.3703 | 14.9721 | FALSE | NA | NA | NA | NA | NA | NA | NA | 5.8795
O00141 | Serine/threonine-protein kinase Sgk1 | SGK1 | 1:1 | Block_1 | 1.62 | 11.245 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N5H7 | SH2 domain-containing protein 3C | SH2D3C | 1:1 | Block_1 | 1.4212 | 10.2257 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.4962
Q5VZ18 | SH2 domain-containing adapter protein E | SHE | 1:1 | Block_1 | 2.3351 | 11.5727 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y2K2 | Serine/threonine-protein kinase SIK3 | SIK3 | 1:1 | Block_1 | 2.6086 | 10.6817 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NTG7 | NAD-dependent protein deacetylase sirtuin-3, mitochondrial | SIRT3 | 1:1 | Block_1 | 1.5668 | 11.1983 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.0816
Q8N6T7 | NAD-dependent protein deacylase sirtuin-6 | SIRT6 | 1:1 | Block_1 | 0.31294 | 11.1162 | FALSE | 6.6477 | 8.1545 | NA | NA | 0.68457 | NA | 0.33992 | 1.2894
Q13621 | Solute carrier family 12 member 1 | SLC12A1 | 1:1 | Block_1 | 1 | 9.3326 | TRUE | NA | 21.3918 | NA | NA | 1.8087 | NA | NA | 1.5886
Q8IYB5 | Stromal membrane-associated protein 1 | SMAP1 | 1:1 | Block_1 | 0.001 | 11.8841 | FALSE | NA | NA | NA | NA | 0.41656 | NA | 0.26393 | 1.9034
Q8IY18 | Structural maintenance of chromosomes protein 5 | SMC5 | 1:1 | Block_1 | 1.15 | 13.8552 | TRUE | NA | 10.4186 | NA | NA | 1.3049 | NA | NA | 1.6738
Q96SB8 | Structural maintenance of chromosomes protein 6 | SMC6 | 1:1 | Block_1 | 2.0007 | 15.9424 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86US8 | Telomerase-binding protein EST1A | SMG6 | 1:1 | Block_1 | 3.3719 | 12.6604 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q92540 | Nonsense-mediated mRNA decay factor SMG7 | SMG7 | 1:1 | Block_1 | 2.4565 | 12.2916 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q2TAL5 | Smoothelin-like protein 2 | SMTNL2 | 1:1 | Block_1 | 1.1784 | 9.6748 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9HAU4 | E3 ubiquitin-protein ligase SMURF2 | SMURF2 | 1:1 | Block_1 | 2.5 | 11.6702 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6NMZ2 | Sentan | SNTN | 1:1 | Block_1 | 1.7071 | 12.2773 | FALSE | NA | NA | NA | NA | NA | NA | NA | 1.9496
P35712 | Transcription factor SOX-6 | SOX6 | 1:1 | Block_1 | 2.551 | 10.5973 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BT81 | Transcription factor SOX-7 | SOX7 | 1:1 | Block_1 | 3.5332 | 13.0967 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96KW9 | Sperm acrosome-associated protein 7 | SPACA7 | 1:1 | Block_1 | 2.081 | 9.102 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5MJ10 | Sperm protein associated with the nucleus on the X chromosome N2 | SPANXN2 | 1:1 | Block_1 | 1.7973 | 8.758 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q7Z5L4 | Spermatogenesis-associated protein 19, mitochondrial | SPATA19 | 1:1 | Block_1 | 3 | 8.2555 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86W54 | Spermatogenesis-associated protein 24 | SPATA24 | 1:1 | Block_1 | 1.4844 | 10.4671 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96N06 | Spermatogenesis-associated protein 33 | SPATA33 | 1:1 | Block_1 | 0.90324 | 12.4186 | TRUE | 11.0806 | 17.3619 | NA | NA | 2.3324 | NA | NA | 2.2522
Q86XZ4 | Spermatogenesis-associated serine-rich protein 2 | SPATS2 | 1:1 | Block_1 | 0.64446 | 7.909 | TRUE | 10.7758 | 9.4888 | NA | NA | 1.2678 | NA | NA | 1.0593
Q5M775 | Cytospin-B | SPECC1 | 1:1 | Block_1 | 5 | 16.6035 | FALSE | NA | NA | NA | NA | 6.1248 | NA | NA | 7.0852
Q9NZD8 | Maspardin | SPG21 | 1:1 | Block_1 | 2.5825 | 15.699 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9BXG8 | Spermatogenic leucine zipper protein 1 | SPZ1 | 1:1 | Block_1 | 0.96611 | 8.7036 | TRUE | NA | 6.5252 | NA | NA | NA | NA | NA | NA
P0DMP2 | SLIT-ROBO Rho GTPase-activating protein 2B | SRGAP2B | 1:1 | Block_1 | 2.5407 | 8.9892 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.5825
Q96JF0 | Beta-galactoside alpha-2,6-sialyltransferase 2 | ST6GAL2 | 1:1 | Block_1 | 2.0091 | 10.5593 | FALSE | NA | NA | NA | NA | 2.2921 | NA | NA | NA
P61647 | Alpha-2,8-sialyltransferase 8F | ST8SIA6 | 1:1 | Block_1 | -0.74056 | 9.3817 | TRUE | NA | NA | NA | NA | -0.51221 | NA | NA | -0.58555
Q99469 | SH3 and cysteine-rich domain-containing protein | STAC | 1:1 | Block_1 | 2.3037 | 12.3939 | FALSE | NA | NA | NA | NA | NA | NA | NA | 2.7788
Q6ZMT1 | SH3 and cysteine-rich domain-containing protein 2 | STAC2 | 1:1 | Block_1 | 3.2922 | 10.5418 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96MF2 | SH3 and cysteine-rich domain-containing protein 3 | STAC3 | 1:1 | Block_1 | 2.5 | 14.6966 | FALSE | 3.2586 | NA | NA | NA | 3.7744 | NA | NA | 4.6239
Q9UGK3 | Signal-transducing adaptor protein 2 | STAP2 | 1:1 | Block_1 | 3.32 | 11.5401 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P2P6 | StAR-related lipid transfer protein 9 | STARD9 | 1:1 | Block_1 | 1.725 | 12.1021 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P42224 | Signal transducer and activator of transcription 1-alpha/beta | STAT1 | 1:1 | Block_1 | 2.418 | 10.091 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UHE8 | Metalloreductase STEAP1 | STEAP1 | 1:1 | Block_1 | 3.264 | 12.1292 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H5V9 | STING ER exit protein | STEEP1 | 1:1 | Block_1 | 3.4102 | 10.6511 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NY57 | Serine/threonine-protein kinase 32B | STK32B | 1:1 | Block_1 | 1.7613 | 11.6262 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TDR2 | Serine/threonine-protein kinase 35 | STK35 | 1:1 | Block_1 | 1.3046 | 10.6387 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q93045 | Stathmin-2 | STMN2 | 1:1 | Block_1 | 6.1205 | 16.4611 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y6Q2 | Stonin-1 | STON1 | 1:1 | Block_1 | 1.9555 | 11.9542 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q7Z7C7 | Stimulated by retinoic acid gene 8 protein homolog | STRA8 | 1:1 | Block_1 | 0.63578 | 11.9857 | FALSE | NA | 10.0557 | NA | NA | 1.931 | NA | NA | 0.90657
Q86Y82 | Syntaxin-12 | STX12 | 1:1 | Block_1 | -0.86847 | 3.1854 | FALSE | 2.9194 | 6.7281 | NA | NA | 0.42572 | NA | -0.85748 | 1.2134
P32856 | Syntaxin-2 | STX2 | 1:1 | Block_1 | 2.4837 | 12.057 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZWJ1 | Syntaxin-binding protein 4 | STXBP4 | 1:1 | Block_1 | 1.1129 | 8.285 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UMX1 | Suppressor of fused homolog | SUFU | 1:1 | Block_1 | 2.9626 | 12.717 | FALSE | NA | NA | NA | NA | NA | NA | NA | 3.7178
O00338 | Sulfotransferase 1C2 | SULT1C2 | 1:1 | Block_1 | 0.001 | 11.3432 | TRUE | NA | NA | NA | NA | NA | NA | NA | 0.26145
P49888 | Sulfotransferase 1E1 | SULT1E1 | 1:1 | Block_1 | 2.0271 | 13.5152 | FALSE | NA | NA | NA | NA | 2.0318 | NA | NA | NA
Q92797 | Symplekin | SYMPK | 1:1 | Block_1 | 0.725 | 7.7211 | TRUE | NA | NA | NA | NA | 0.76504 | NA | NA | 0.73344
Q4VX76 | Synaptotagmin-like protein 3 | SYTL3 | 1:1 | Block_1 | 3.1123 | 12.5421 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TEA7 | TBC domain-containing protein kinase-like protein | TBCK | 1:1 | Block_1 | 1.3192 | 12.3989 | TRUE | NA | NA | NA | NA | 1.9067 | NA | NA | 1.9614
Q6SJ96 | TATA box-binding protein-like 2 | TBPL2 | 1:1 | Block_1 | 1.1003 | 7.184 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y458 | T-box transcription factor TBX22 | TBX22 | 1:1 | Block_1 | 1.5123 | 11.2593 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q7Z6L1 | Tectonin beta-propeller repeat-containing protein 1 | TECPR1 | 1:1 | Block_1 | 3.435 | 16.9446 | TRUE | NA | NA | NA | NA | NA | NA | NA | 4.5547
Q8WW24 | Tektin-4 | TEKT4 | 1:1 | Block_1 | 0.77438 | 12.273 | FALSE | 0.3153 | 26.5598 | NA | NA | 0.77877 | NA | NA | NA
Q9Y4R8 | Telomere length regulation protein TEL2 homolog | TELO2 | 1:1 | Block_1 | 1.8098 | 11.7109 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96IP4 | Terminal nucleotidyltransferase 5A | TENT5A | 1:1 | Block_1 | 0.9837 | 10.3091 | TRUE | NA | NA | NA | NA | 0.99273 | NA | NA | NA
Q96N21 | AP-4 complex accessory subunit Tepsin | TEPSIN | 1:1 | Block_1 | 2.2484 | 10.0571 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y4I5 | Tesmin | TESMIN | 1:1 | Block_1 | 0.43237 | 11.507 | FALSE | 10.2903 | 11.3068 | NA | NA | 1.1363 | NA | NA | 1.4283
Q96LM6 | Testis-expressed sequence 37 protein | TEX37 | 1:1 | Block_1 | 1.7512 | 13.6203 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q15583 | Homeobox protein TGIF1 | TGIF1 | 1:1 | Block_1 | 2.0344 | 11.8748 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P07101 | Tyrosine 3-monooxygenase | TH | 1:1 | Block_1 | 0.75308 | 10.1804 | FALSE | 17.3257 | 29.0503 | NA | NA | 0.94719 | NA | NA | 1.6932
Q7Z2Z1 | Treslin | TICRR | 1:1 | Block_1 | 0.44437 | 11.4403 | TRUE | NA | NA | NA | NA | 0.70441 | NA | NA | 1.0117
Q96MW7 | Tigger transposable element-derived protein 1 | TIGD1 | 1:1 | Block_1 | 0.3706 | 5.1432 | FALSE | NA | 29.5565 | NA | NA | 0.60933 | NA | NA | NA
Q7Z3E1 | Protein mono-ADP-ribosyltransferase TIPARP | TIPARP | 1:1 | Block_1 | 0.95299 | 9.638 | TRUE | NA | 4.8529 | NA | NA | NA | NA | NA | NA
Q04725 | Transducin-like enhancer protein 2 | TLE2 | 1:1 | Block_1 | 2.4481 | 7.623 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H808 | Transducin-like enhancer protein 6 | TLE6 | 1:1 | Block_1 | 3.2725 | 13.7716 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NWD8 | Transmembrane protein 248 | TMEM248 | 1:1 | Block_1 | 3.0586 | 13.3955 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NRS4 | Transmembrane protease serine 4 | TMPRSS4 | 1:1 | Block_1 | 5.7282 | 13.4092 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y320 | Thioredoxin-related transmembrane protein 2 | TMX2 | 1:1 | Block_1 | 0.47348 | 7.3372 | TRUE | NA | NA | NA | NA | 0.49739 | NA | NA | 0.71746
Q9C0C2 | 182 kDa tankyrase-1-binding protein | TNKS1BP1 | 1:1 | Block_1 | 2.0534 | 9.6082 | TRUE | NA | 12.1592 | NA | NA | NA | NA | NA | 2.6208
P63316 | Troponin C, slow skeletal and cardiac muscles | TNNC1 | 1:1 | Block_1 | 0.54618 | 9.5666 | TRUE | 12.3018 | 16.2285 | NA | NA | 1.795 | NA | 0.69809 | 2.6834
P45379 | Troponin T, cardiac muscle | TNNT2 | 1:1 | Block_1 | 1.3499 | 14.4343 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IZW8 | Tensin-4 | TNS4 | 1:1 | Block_1 | 0.84807 | 11.537 | TRUE | 6.1644 | 20.1777 | NA | NA | 1.5938 | NA | NA | 1.331
Q8N6T0 | Type 2 DNA topoisomerase 6 subunit B-like | TOP6BL | 1:1 | Block_1 | 4.3744 | 14.5127 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q5T0D9 | Tumor protein p63-regulated gene 1-like protein | TPRG1L | 1:1 | Block_1 | 2.0988 | 9.9779 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TDR0 | TRAF3-interacting protein 1 | TRAF3IP1 | 1:1 | Block_1 | 2.0447 | 15.2274 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.5734
Q9Y577 | E3 ubiquitin-protein ligase TRIM17 | TRIM17 | 1:1 | Block_1 | 0.89943 | 10.2844 | FALSE | 4.1926 | 6.3777 | NA | NA | 1.4334 | NA | NA | 0.93424
Q9BZY9 | E3 ubiquitin-protein ligase TRIM31 | TRIM31 | 1:1 | Block_1 | 0.72209 | 12.1275 | FALSE | 8.0735 | 5.5756 | NA | NA | 0.96984 | NA | NA | NA
Q9BU70 | tRNA (adenine(37)-N6)-methyltransferase | TRMO | 1:1 | Block_1 | 2.427 | 11.4309 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NUP7 | tRNA:m(4)X modification enzyme TRM13 homolog | TRMT13 | 1:1 | Block_1 | 1.2523 | 10.4346 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UBN4 | Short transient receptor potential channel 4 | TRPC4 | 1:1 | Block_1 | 1.5 | 8.4271 | FALSE | NA | 15.9846 | NA | NA | NA | NA | NA | NA
Q9UHF7 | Zinc finger transcription factor Trps1 | TRPS1 | 1:1 | Block_1 | 2.3486 | 9.6425 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q3SY00 | Testis-specific protein 10-interacting protein | TSGA10IP | 1:1 | Block_1 | 2.6693 | 13.2012 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q2TAA8 | Translin-associated factor X-interacting protein 1 | TSNAXIP1 | 1:1 | Block_1 | 1.8348 | 10.9893 | FALSE | 16.3986 | 12.6641 | NA | NA | NA | NA | NA | 2.2664
Q9UJK0 | 18S rRNA aminocarboxypropyltransferase | TSR3 | 1:1 | Block_1 | 1.9149 | 8.448 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A2A3L6 | Tetratricopeptide repeat protein 24 | TTC24 | 1:1 | Block_1 | 0.86537 | 10.3809 | FALSE | NA | 6.5744 | NA | NA | 1.5758 | NA | NA | 1.1531
Q6EMB2 | Tubulin polyglutamylase TTLL5 | TTLL5 | 1:1 | Block_1 | 0.65 | 10.7986 | FALSE | 17.2075 | 31.8569 | NA | NA | 1.2234 | NA | NA | 0.94513
Q8N841 | Tubulin polyglutamylase TTLL6 | TTLL6 | 1:1 | Block_1 | 0.97975 | 11.4174 | TRUE | 2.8004 | NA | NA | NA | NA | NA | NA | NA
Q6ZT98 | Tubulin polyglutamylase TTLL7 | TTLL7 | 1:1 | Block_1 | 3.5636 | 11.8482 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A6NNZ2 | Tubulin beta 8B | TUBB8B | 1:1 | Block_1 | 2.7932 | 12.0013 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96RT8 | Gamma-tubulin complex component 5 | TUBGCP5 | 1:1 | Block_1 | 2.0069 | 11.8765 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.6839
Q5TAX3 | Terminal uridylyltransferase 4 | TUT4 | 1:1 | Block_1 | 2.2693 | 11.0341 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
P62987 | Ubiquitin-60S ribosomal protein L40 | UBA52 | 1:1 | Block_1 | 0.46688 | 6.7116 | TRUE | NA | 10.5609 | NA | NA | 0.78032 | NA | NA | 0.91985
Q96LR5 | Ubiquitin-conjugating enzyme E2 E2 | UBE2E2 | 1:1 | Block_1 | 0.98325 | 13.4796 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.3325
Q5VVX9 | Ubiquitin-conjugating enzyme E2 U | UBE2U | 1:1 | Block_1 | 0.001 | 11.188 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q6ZU65 | Ubinuclein-2 | UBN2 | 1:1 | Block_1 | 0.42558 | 8.5291 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IWV7 | E3 ubiquitin-protein ligase UBR1 | UBR1 | 1:1 | Block_1 | 1.6318 | 7.8718 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N806 | Putative E3 ubiquitin-protein ligase UBR7 | UBR7 | 1:1 | Block_1 | -0.5 | 8.8923 | TRUE | 15.2489 | 6.9492 | NA | NA | 0.32698 | NA | NA | 0.51375
Q96LJ8 | UBX domain-containing protein 10 | UBXN10 | 1:1 | Block_1 | 1.4602 | 11.0972 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P68543 | UBX domain-containing protein 2A | UBXN2A | 1:1 | Block_1 | 1.1583 | 8.8005 | FALSE | 10.9432 | 8.7855 | NA | NA | NA | NA | NA | 1.4464
Q9HA47 | Uridine-cytidine kinase 1 | UCK1 | 1:1 | Block_1 | 1.5 | 8.4104 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N6Y0 | Harmonin-binding protein USHBP1 | USHBP1 | 1:1 | Block_1 | 0.41246 | 12.1278 | FALSE | NA | NA | NA | NA | 1.078 | NA | NA | 1.0589
P51784 | Ubiquitin carboxyl-terminal hydrolase 11 | USP11 | 1:1 | Block_1 | 1.1595 | 9.1093 | FALSE | NA | 9.8926 | NA | NA | NA | NA | NA | 1.5858
Q9P275 | Ubiquitin carboxyl-terminal hydrolase 36 | USP36 | 1:1 | Block_1 | 0.93725 | 11.8391 | TRUE | 15.1919 | 27.9114 | NA | NA | 1.8065 | NA | NA | 2.4149
Q9H9J4 | Ubiquitin carboxyl-terminal hydrolase 42 | USP42 | 1:1 | Block_1 | -0.26488 | 10.8135 | TRUE | NA | NA | NA | NA | NA | NA | NA | -0.25397
Q70EL4 | Ubiquitin carboxyl-terminal hydrolase 43 | USP43 | 1:1 | Block_1 | 2 | 11.5684 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q86UV5 | Ubiquitin carboxyl-terminal hydrolase 48 | USP48 | 1:1 | Block_1 | 1.208 | 9.9298 | FALSE | NA | 11.8393 | NA | NA | 1.7039 | NA | NA | 1.6906
P49754 | Vacuolar protein sorting-associated protein 41 homolog | VPS41 | 1:1 | Block_1 | 2.9023 | 12.9444 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8N398 | von Willebrand factor A domain-containing protein 5B2 | VWA5B2 | 1:1 | Block_1 | 1.3309 | 13.3834 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
A3KMH1 | von Willebrand factor A domain-containing protein 8 | VWA8 | 1:1 | Block_1 | 0.2572 | 14.0797 | FALSE | 9.9897 | 8.4614 | NA | NA | 0.50718 | NA | NA | 1.0443
Q7Z5K2 | Wings apart-like protein homolog | WAPL | 1:1 | Block_1 | 2 | 12.2523 | FALSE | NA | NA | NA | NA | 2.8749 | NA | NA | 5.4986
Q64LD2 | WD repeat-containing protein 25 | WDR25 | 1:1 | Block_1 | 1.9966 | 13.4644 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9H977 | WD repeat-containing protein 54 | WDR54 | 1:1 | Block_1 | 1.0198 | 11.3502 | TRUE | NA | NA | NA | NA | 1.2406 | NA | NA | 1.4434
B1ANS9 | WD repeat-containing protein 64 | WDR64 | 1:1 | Block_1 | -0.28178 | 10.8604 | TRUE | 13.1859 | 8.0661 | NA | NA | 0.30226 | NA | NA | 0.90097
Q9BQA1 | Methylosome protein 50 | WDR77 | 1:1 | Block_1 | 2.5 | 13.1352 | FALSE | NA | NA | NA | NA | 2.6835 | NA | NA | 4.0723
Q9P202 | Whirlin | WHRN | 1:1 | Block_1 | 3.62 | 13.2767 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96J92 | Serine/threonine-protein kinase WNK4 | WNK4 | 1:1 | Block_1 | 1.746 | 13.3824 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
O00755 | Protein Wnt-7a | WNT7A | 1:1 | Block_1 | 1.5 | 11.5395 | FALSE | 3.9848 | 2.7848 | NA | NA | 1.557 | NA | NA | NA
P19544 | Wilms tumor protein | WT1 | 1:1 | Block_1 | 1.859 | 9.9585 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UBH6 | Xenotropic and polytropic retrovirus receptor 1 | XPR1 | 1:1 | Block_1 | 1.2573 | 11.304 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IZH2 | 5'-3' exoribonuclease 1 | XRN1 | 1:1 | Block_1 | 3.3731 | 12.4077 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q6P2D8 | X-ray radiation resistance-associated protein 1 | XRRA1 | 1:1 | Block_1 | 2.1233 | 13.7444 | TRUE | NA | NA | NA | NA | 2.6008 | NA | NA | NA
Q9NRH1 | Protein YAE1 homolog | YAE1 | 1:1 | Block_1 | 1.3929 | 10.8092 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8IYB0 | Putative uncharacterized protein MGC39545 | YK038 | 1:1 | Block_1 | 1.9152 | 13.9585 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96M66 | Putative uncharacterized protein FLJ32790 | YP010 | 1:1 | Block_1 | 1.7853 | 13.0295 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q49AG3 | Zinc finger BED domain-containing protein 5 | ZBED5 | 1:1 | Block_1 | -0.063427 | 10.8707 | FALSE | NA | NA | NA | NA | 0.81171 | NA | NA | 1.0126
Q99592 | Zinc finger and BTB domain-containing protein 18 | ZBTB18 | 1:1 | Block_1 | 0.84798 | 9.1362 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9HC78 | Zinc finger and BTB domain-containing protein 20 | ZBTB20 | 1:1 | Block_1 | 2.2857 | 13.035 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8NCP5 | Zinc finger and BTB domain-containing protein 44 | ZBTB44 | 1:1 | Block_1 | 1.4789 | 9.9203 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q5TFG8 | Zinc finger C2HC domain-containing protein 1B | ZC2HC1B | 1:1 | Block_1 | 2.3829 | 13.1034 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WU90 | Zinc finger CCCH domain-containing protein 15 | ZC3H15 | 1:1 | Block_1 | 2.728 | 11.3347 | TRUE | 13.1704 | NA | NA | NA | NA | NA | NA | 2.9735
Q86WB0 | Zinc finger C3HC-type protein 1 | ZC3HC1 | 1:1 | Block_1 | -0.2987 | 11.5332 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8WYQ9 | Zinc finger CCHC domain-containing protein 14 | ZCCHC14 | 1:1 | Block_1 | -0.32213 | 14.5654 | FALSE | 6.7834 | 5.8394 | NA | 0.72326 | 2.078 | NA | 0.010997 | 2.1939
Q9H0M4 | Zinc finger CW-type PWWP domain protein 1 | ZCWPW1 | 1:1 | Block_1 | 2.3638 | 13.2919 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q504Y3 | Zinc finger CW-type PWWP domain protein 2 | ZCWPW2 | 1:1 | Block_1 | 1.5 | 13.0561 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Q86XD8 | AN1-type zinc finger protein 4 | ZFAND4 | 1:1 | Block_1 | -0.17021 | 13.7638 | TRUE | NA | NA | NA | NA | 0.92322 | NA | NA | 0.85214
Q9P243 | Zinc finger protein ZFAT | ZFAT | 1:1 | Block_1 | 1.0579 | 10.6927 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.1599
Q15911 | Zinc finger homeobox protein 3 | ZFHX3 | 1:1 | Block_1 | 1.9868 | 11.3538 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9NU63 | Zinc finger protein 57 homolog | ZFP57 | 1:1 | Block_1 | 1.1685 | 10.3845 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q8TF47 | Zinc finger protein 90 homolog | ZFP90 | 1:1 | Block_1 | 2.4211 | 11.4188 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9UPR6 | Zinc finger RNA-binding protein 2 | ZFR2 | 1:1 | Block_1 | 0.47375 | 10.0877 | TRUE | 16.413 | 22.8034 | NA | NA | 2.4619 | NA | 0.57859 | 2.5994
Q9HCC9 | Lateral signaling target protein 2 homolog | ZFYVE28 | 1:1 | Block_1 | 1.9508 | 9.895 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q96T25 | Zinc finger protein ZIC 5 | ZIC5 | 1:1 | Block_1 | 6.6457 | 16.0626 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
P52739 | Zinc finger protein 131 | ZNF131 | 1:1 | Block_1 | 1.9222 | 12.1556 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q0IIN9 | Putative uncharacterized protein ZNF252P-AS1 | ZNF252P-AS1 | 1:1 | Block_1 | 3.1253 | 14.4732 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9Y3S2 | Zinc finger protein 330 | ZNF330 | 1:1 | Block_1 | 1.5 | 8.9668 | FALSE | NA | NA | NA | NA | 2.7424 | NA | 2.1634 | 4.0634
Q9UL40 | Zinc finger protein 346 | ZNF346 | 1:1 | Block_1 | 1.1481 | 12.7593 | FALSE | NA | NA | NA | NA | 1.659 | NA | NA | NA
Q9GZX5 | Zinc finger protein 350 | ZNF350 | 1:1 | Block_1 | 0.98408 | 14.1669 | TRUE | NA | NA | NA | NA | NA | NA | NA | 2.2474
Q96N95 | Zinc finger protein 396 | ZNF396 | 1:1 | Block_1 | 2.6875 | 10.2762 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
A8K554 | Putative protein ZNF815 | ZNF815P | 1:1 | Block_1 | 3.4048 | 12.5794 | TRUE | NA | NA | NA | NA | NA | NA | NA | NA
Q9P2E3 | NFX1-type zinc finger-containing protein 1 | ZNFX1 | 1:1 | Block_1 | 1.6307 | 12.4448 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.6343
A7E2V4 | Zinc finger SWIM domain-containing protein 8 | ZSWIM8 | 1:1 | Block_1 | 1.0766 | 10.0296 | TRUE | NA | NA | NA | NA | NA | NA | NA | 1.1458
Q96AP4 | Zinc finger-containing ubiquitin peptidase 1 | ZUP1 | 1:1 | Block_1 | 1.2725 | 12.5817 | TRUE | 14.6597 | 8.8534 | NA | NA | NA | NA | NA | 1.5126
Q9H900 | Protein zwilch homolog | ZWILCH | 1:1 | Block_1 | 3.1196 | 12.6116 | FALSE | NA | NA | NA | NA | NA | NA | NA | NA
Olink® Explore HT Validation Data

## Biomarker Specificity
Olink® Explore HT Biomarker Specificity Information |  |  |  | Last updated October 4, 2023
Protein X |  | Protein Y |  | Sequence homology | Estimated signal contribution*
Uniprot | Gene name | Uniprot | Gene name
P05187 | ALPP | P10696 | ALPG | 96.4% | ~25%
P10696 | ALPG | P05187 | ALPP | 96.4% | ~50%
Q08708 | CD300C | Q9UGN4 | CD300A | 80.4% | ~12%
Q9UGN4 | CD300A | Q08708 | CD300C | 80.4% | ~25%
P41208 | CETN2 | Q12798 | CETN1 | 83.7% | 50-70%
P0DN86 | CGB3_CGB5_CGB8 | P01229 | LHB | 85.2% | 25-50%
P01241 | GH1 | P01242 | GH2 | 93.1% | ~12%
Q86SG3 | DAZ4 | Q9NQZ3 | DAZ1 | 100% | ~12%
Q9NQZ3 | DAZ1 | Q86SG3 | DAZ4 | 100% | ~100%
Q6ZTU2 | EP400P1 | Q96L91 | EP400 | 84.2% | ~12%
P31994 | FCGR2B | P12318 | FCGR2A | 92% | ~25%
P12318 | FCGR2A | P31994 | FCGR2B | 92% | ~50%
Q6NT46 | GAGE2A | O76087 | GAGE7 | 96.6% | ~17%
P62760 | VSNL1 | Q9UM19 | HPCAL4 | 89.5% | ~25%
O95239 | KIF4A | Q2VIQ3 | KIF4B | 93.8% | ~12%
Q2VIQ3 | KIF4B | O95239 | KIF4A | 93.8% | ~35%
P02533 | KRT14 | Q04695 | KRT17 | 82.6% | ~25%
Q8NFP7 | NUDT10 | Q9NZJ9 | NUDT4 | 90.2% | <25%
Q6P474 | PDXDC2P | Q6P996 | PDXDC1 | 97% | ~12%
A6NI47 | POTEM | B2RU33 | POTEC | 86.6% | ~18%
B2RU33 | POTEC | H3BUK9 | POTEB2 | 90.3% | ~25%
Q96PQ5 | PPP1R2P1 | Q6NXS1 | PPP1R2B | 87.8% | ~100%
P63098 | PPP3R1 | Q96LZ3 | PPP3R2 | 84.7% | ~100%
P05451 | REG1A | P48304 | REG1B | 86.7% | ~25%
Q96AT9 | RPE | Q2QD12 | RPEL1 | 96.1% | ~100%
Q9NY46 | SCN3A | Q99250 | SCN2A | 87% | ~25%
P12829 | MYL4 | P08590 | MYL3 | 82.4% | ~12%
*This table lists all biomarkers with observed cross-reactivity where the antibody probe for Protein X recognizes Protein Y, and where Protein Y gives a signal contribution of at least 10% as indicated in column F.